## Calibration (Start Here)
General Information | Overview Procedure for Creating a Calibration Curve | Overview Procedure for Testing Analytes
Background: This example document uses Major Mix calibrant on a Waters Cyclic Ion Mobility Spectrometer. Experiments were run using electrospray ionization in positive mode to create calibration power curves for ions between 311-1013 m/z. The separation times used for this experiment were 0.01, 2, 10, 15, 30, 32, 50, 75 and 100 ms and separation times exhibting wrap-around were removed. The accompanying example experiment data file can be found within the R Shiny application and in the supporting information. | Step 1. Note instrument settings for the separation cell so they can be duplicated during experimental analysis. The instrument settings used for this demonstration can be found in the following tab. Instrument settings in the cyclic control page were held constant between calibration curve creation and experimental analysis. | Step 1. Mirror separation cell settings from the calibration curve and collect by-pass (0.01 ms), single pass (2 ms) and all other desired separation time data.
Uploading: This document can be uploaded into the Multipass CCS Refiner R scripts or R Shiny web application as an example. The analyte names and CCS values can be copied into their respective fields in the application. The authors recommend saving a copy of this document and pasting user data directly to avoid error messages. Analyte Names: Sulfadimethoxide,Val-Tyr-Val,Terfenadine,Polyalanine (n7),Leucine Enkephalin,Polyalanine (n8),Resperine,Polyalanine (n9),Polyalanine (n10),Polyalanine (n11),Polyalanine(n12),Polyalanine (n13),Polyalanine (n14) CCS Values: 168.4,191.7,228.7,211.0,229.8,228.0,252.3,243.0,256.0,271.0,282.0,294.0,306.0 | Step 2. Collect a bypass time using a separation time of 0.01 ms. An acquistion time of at least 30 seconds is recommended for each separation time acquired. Step 3. Change the separation time to increase the number of passes starting with a single pass acquistion at 2 ms. Increasing the number of bins may be necessary for larger ions to ensure that data collection occurs over the total ion arrival time distribution. At least 5 additional separation times are recommended to increase the number of points used for linear fitting. | Step 2. Upload experimental data in .xlsx format. Again the authors recommend saving a copy of the accompanying experimental file and adding user data directly. Step 3. Comma separated analyte names are required but known CCS values are optional. All result files can be downloaded directly within the application in .pdf format.
Data Preparation: Tabs starting with the word Analyte will be evaluated by the code and other tabs will be ignored. The first row of each analyte tab is the separation time and the second row contains column labels. Separation times can be changed at the users discretion as long as a zero and one pass measurement are run first at 0.01 and 2 ms separation times. Time points where wrap-around occurs should be removed and can be replaced with similar separation times where wrap-around does not occur. | Step 4. Upload calibration data as a .xlsx file. Enter comma separated analyte names and corresponding CCS values. After a calibration curve has been run it can be saved to either calibration curve slot and used to calculate experimental CCS values. | The accompanying example experiment file investigates three lipids in SPLASH LIPIDOMIX and can be uploaded to the Process Data tab. Names and Known CCS values for the example data file are provided below. Analyte Names: d7-PC (15:0-18:1), d7-LysoPC (18:1), d7-PE (15:0-18:1), d7-LysoPE (18:1), d7-LysoPE (18:1) M+Na, d7-LysoPC (18:1) M+Na Comparison CCS values are from DT-IMS experiments sourced from the McClean Compendium. CCS Values: 285.4, 235.0, 273.5, 217.8, 224.3, 238.2

## Analyte Sulfadimethoxide
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.032 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.23 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.428 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.626 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
10.824 | 1272 | 12.804 | 0 | 20.789 | 25 | 25.805 | 23 | 30.821 | 27 | 40.787 | 26 | 42.767 | 0 | 60.784 | 9 | 85.797 | 0 | 110.81 | 21
11.022 | 69482 | 13.002 | 0 | 20.987 | 224 | 26.003 | 310 | 31.019 | 204 | 40.985 | 123 | 42.965 | 149 | 60.982 | 1139 | 85.995 | 331 | 111.008 | 170
11.22 | 726509 | 13.2 | 0 | 21.185 | 695 | 26.201 | 277 | 31.217 | 390 | 41.183 | 300 | 43.163 | 283 | 61.18 | 8418 | 86.193 | 1591 | 111.206 | 462
11.418 | 613009 | 13.398 | 0 | 21.383 | 512 | 26.399 | 321 | 31.415 | 307 | 41.381 | 167 | 43.361 | 246 | 61.378 | 24933 | 86.391 | 4969 | 111.404 | 1156
11.616 | 80158 | 13.596 | 15 | 21.581 | 363 | 26.597 | 264 | 31.613 | 391 | 41.579 | 247 | 43.559 | 332 | 61.576 | 64524 | 86.589 | 12955 | 111.602 | 3042
11.814 | 6405 | 13.794 | 8 | 21.779 | 447 | 26.795 | 241 | 31.811 | 268 | 41.777 | 314 | 43.757 | 365 | 61.774 | 114595 | 86.787 | 27735 | 111.8 | 7345
12.012 | 1353 | 13.992 | 87 | 21.977 | 527 | 26.993 | 386 | 32.009 | 290 | 41.975 | 304 | 43.955 | 467 | 61.972 | 150557 | 86.985 | 51494 | 111.998 | 14937
12.21 | 794 | 14.19 | 44 | 22.175 | 636 | 27.191 | 292 | 32.207 | 293 | 42.173 | 316 | 44.153 | 1105 | 62.17 | 152881 | 87.183 | 74480 | 112.196 | 26067
12.408 | 239 | 14.388 | 120 | 22.373 | 609 | 27.389 | 533 | 32.405 | 318 | 42.371 | 221 | 44.351 | 6122 | 62.368 | 113954 | 87.381 | 87951 | 112.394 | 41001
12.606 | 226 | 14.586 | 178 | 22.571 | 463 | 27.587 | 5542 | 32.603 | 368 | 42.569 | 218 | 44.549 | 31044 | 62.566 | 66356 | 87.579 | 89341 | 112.592 | 53406
12.804 | 94 | 14.784 | 155 | 22.769 | 376 | 27.785 | 54184 | 32.801 | 299 | 42.767 | 279 | 44.747 | 97912 | 62.764 | 27856 | 87.777 | 72986 | 112.79 | 63271
13.002 | 6 | 14.982 | 69 | 22.967 | 554 | 27.983 | 230902 | 32.999 | 539 | 42.965 | 279 | 44.945 | 192289 | 62.962 | 9789 | 87.975 | 49666 | 112.988 | 62350
13.2 | 5 | 15.18 | 58 | 23.165 | 753 | 28.181 | 440150 | 33.197 | 640 | 43.163 | 266 | 45.143 | 252551 | 63.16 | 2271 | 88.173 | 27819 | 113.186 | 53246
13.398 | 0 | 15.378 | 137 | 23.363 | 1156 | 28.379 | 385667 | 33.395 | 548 | 43.361 | 325 | 45.341 | 218595 | 63.358 | 615 | 88.371 | 12663 | 113.384 | 40059
13.596 | 6 | 15.576 | 152 | 23.561 | 875 | 28.577 | 165290 | 33.593 | 341 | 43.559 | 343 | 45.539 | 126806 | 63.556 | 278 | 88.569 | 5246 | 113.582 | 27097
13.794 | 0 | 15.773 | 132 | 23.759 | 586 | 28.775 | 35486 | 33.791 | 284 | 43.757 | 464 | 45.737 | 49416 | 63.754 | 178 | 88.767 | 1697 | 113.78 | 15009
13.992 | 4 | 15.971 | 235 | 23.957 | 411 | 28.973 | 4016 | 33.989 | 330 | 43.955 | 402 | 45.935 | 12312 | 63.952 | 240 | 88.965 | 470 | 113.978 | 7608
14.19 | 4 | 16.169 | 109 | 24.155 | 189 | 29.171 | 987 | 34.187 | 233 | 44.153 | 803 | 46.133 | 2549 | 64.15 | 221 | 89.163 | 157 | 114.176 | 3075
14.388 | 0 | 16.367 | 100 | 24.353 | 174 | 29.369 | 395 | 34.385 | 308 | 44.351 | 3760 | 46.331 | 635 | 64.348 | 239 | 89.361 | 179 | 114.374 | 1464
14.586 | 0 | 16.565 | 154 | 24.551 | 277 | 29.567 | 1003 | 34.583 | 358 | 44.549 | 19720 | 46.529 | 325 | 64.546 | 178 | 89.559 | 207 | 114.572 | 506
14.784 | 10 | 16.763 | 174 | 24.749 | 256 | 29.765 | 1363 | 34.781 | 354 | 44.747 | 74839 | 46.727 | 389 | 64.744 | 326 | 89.757 | 151 | 114.77 | 148
14.982 | 3 | 16.961 | 216 | 24.947 | 264 | 29.963 | 1332 | 34.979 | 476 | 44.945 | 172338 | 46.924 | 416 | 64.942 | 264 | 89.955 | 169 | 114.968 | 174
15.18 | 0 | 17.159 | 130 | 25.145 | 281 | 30.161 | 1099 | 35.177 | 704 | 45.143 | 258046 | 47.122 | 565 | 65.14 | 197 | 90.153 | 114 | 115.166 | 142
15.378 | 0 | 17.357 | 163 | 25.343 | 188 | 30.359 | 823 | 35.375 | 564 | 45.341 | 254973 | 47.32 | 404 | 65.338 | 272 | 90.351 | 136 | 115.364 | 170
15.576 | 0 | 17.555 | 91 | 25.541 | 273 | 30.557 | 524 | 35.573 | 501 | 45.539 | 167836 | 47.518 | 357 | 65.536 | 270 | 90.549 | 212 | 115.562 | 175
15.773 | 0 | 17.753 | 169 | 25.739 | 332 | 30.755 | 389 | 35.771 | 712 | 45.737 | 72600 | 47.716 | 451 | 65.734 | 214 | 90.747 | 181 | 115.76 | 140
15.971 | 0 | 17.951 | 128 | 25.937 | 452 | 30.953 | 308 | 35.969 | 2405 | 45.935 | 21666 | 47.914 | 526 | 65.932 | 147 | 90.945 | 146 | 115.958 | 145
16.169 | 0 | 18.149 | 158 | 26.135 | 387 | 31.151 | 233 | 36.167 | 18908 | 46.133 | 4230 | 48.112 | 779 | 66.13 | 293 | 91.143 | 195 | 116.156 | 193
16.367 | 0 | 18.347 | 149 | 26.333 | 292 | 31.349 | 272 | 36.365 | 88307 | 46.331 | 839 | 48.31 | 891 | 66.328 | 319 | 91.341 | 181 | 116.354 | 90
16.565 | 0 | 18.545 | 398 | 26.531 | 220 | 31.547 | 255 | 36.563 | 231210 | 46.529 | 365 | 48.508 | 948 | 66.526 | 346 | 91.539 | 128 | 116.552 | 99
16.763 | 0 | 18.743 | 628 | 26.729 | 221 | 31.745 | 270 | 36.761 | 354559 | 46.727 | 348 | 48.706 | 763 | 66.724 | 480 | 91.737 | 194 | 116.75 | 127
16.961 | 0 | 18.941 | 481 | 26.927 | 297 | 31.943 | 424 | 36.959 | 310357 | 46.924 | 402 | 48.904 | 564 | 66.922 | 499 | 91.935 | 149 | 116.948 | 108
17.159 | 0 | 19.139 | 470 | 27.125 | 287 | 32.141 | 216 | 37.157 | 164330 | 47.122 | 353 | 49.102 | 410 | 67.12 | 426 | 92.133 | 139 | 117.146 | 67
17.357 | 0 | 19.337 | 2935 | 27.323 | 345 | 32.339 | 169 | 37.355 | 52084 | 47.32 | 598 | 49.3 | 395 | 67.318 | 394 | 92.331 | 131 | 117.344 | 77
17.555 | 0 | 19.535 | 67164 | 27.521 | 388 | 32.537 | 275 | 37.553 | 10347 | 47.518 | 463 | 49.498 | 366 | 67.516 | 338 | 92.529 | 160 | 117.542 | 155
17.753 | 0 | 19.733 | 452780 | 27.719 | 4232 | 32.735 | 348 | 37.751 | 1718 | 47.716 | 485 | 49.696 | 385 | 67.714 | 275 | 92.727 | 99 | 117.74 | 93
17.951 | 0 | 19.931 | 773970 | 27.917 | 47416 | 32.933 | 276 | 37.949 | 602 | 47.914 | 774 | 49.894 | 302 | 67.912 | 259 | 92.925 | 170 | 117.938 | 95
18.149 | 0 | 20.129 | 395878 | 28.115 | 227140 | 33.131 | 544 | 38.147 | 346 | 48.112 | 840 | 50.092 | 247 | 68.11 | 245 | 93.123 | 99 | 118.136 | 129
18.347 | 0 | 20.327 | 63890 | 28.313 | 480596 | 33.329 | 656 | 38.345 | 251 | 48.31 | 923 | 50.29 | 198 | 68.308 | 171 | 93.321 | 132 | 118.334 | 104
18.545 | 0 | 20.525 | 5296 | 28.511 | 467695 | 33.527 | 383 | 38.543 | 287 | 48.508 | 963 | 50.488 | 162 | 68.506 | 279 | 93.519 | 93 | 118.532 | 66
18.743 | 0 | 20.723 | 2767 | 28.709 | 217230 | 33.725 | 218 | 38.741 | 309 | 48.706 | 742 | 50.686 | 113 | 68.704 | 239 | 93.717 | 95 | 118.73 | 70
18.941 | 0 | 20.921 | 2119 | 28.907 | 48945 | 33.923 | 144 | 38.939 | 795 | 48.904 | 516 | 50.884 | 228 | 68.902 | 161 | 93.915 | 114 | 118.928 | 126
19.139 | 0 | 21.119 | 1115 | 29.105 | 6326 | 34.121 | 134 | 39.137 | 905 | 49.102 | 479 | 51.082 | 135 | 69.1 | 166 | 94.113 | 243 | 119.126 | 96
19.337 | 0 | 21.317 | 533 | 29.303 | 930 | 34.319 | 152 | 39.335 | 1039 | 49.3 | 289 | 51.28 | 183 | 69.298 | 105 | 94.311 | 260 | 119.324 | 108
19.535 | 0 | 21.515 | 358 | 29.501 | 556 | 34.517 | 164 | 39.533 | 858 | 49.498 | 315 | 51.478 | 196 | 69.496 | 368 | 94.509 | 294 | 119.522 | 196
19.733 | 0 | 21.713 | 376 | 29.699 | 871 | 34.715 | 175 | 39.731 | 598 | 49.696 | 280 | 51.676 | 103 | 69.694 | 628 | 94.707 | 406 | 119.72 | 137
19.931 | 0 | 21.911 | 338 | 29.897 | 1363 | 34.913 | 181 | 39.929 | 559 | 49.894 | 254 | 51.874 | 151 | 69.892 | 193 | 94.905 | 335 | 119.918 | 73
20.129 | 0 | 22.109 | 377 | 30.095 | 1491 | 35.111 | 335 | 40.127 | 323 | 50.092 | 192 | 52.072 | 110 | 70.09 | 75 | 95.103 | 275 | 120.116 | 103
20.327 | 0 | 22.307 | 527 | 30.293 | 1011 | 35.309 | 437 | 40.325 | 320 | 50.29 | 154 | 52.27 | 98 | 70.288 | 52 | 95.301 | 227 | 120.314 | 34
20.525 | 0 | 22.505 | 592 | 30.491 | 752 | 35.507 | 409 | 40.523 | 202 | 50.488 | 164 | 52.468 | 98 | 70.486 | 87 | 95.499 | 197 | 120.512 | 59
20.723 | 0 | 22.703 | 563 | 30.689 | 451 | 35.705 | 501 | 40.721 | 136 | 50.686 | 134 | 52.666 | 109 | 70.684 | 171 | 95.697 | 138 | 120.71 | 64
20.921 | 0 | 22.901 | 371 | 30.887 | 350 | 35.903 | 399 | 40.919 | 122 | 50.884 | 151 | 52.864 | 92 | 70.882 | 134 | 95.895 | 59 | 120.908 | 58
21.119 | 0 | 23.099 | 510 | 31.085 | 150 | 36.101 | 349 | 41.117 | 78 | 51.082 | 74 | 53.062 | 53 | 71.08 | 178 | 96.093 | 64 | 121.106 | 22
21.317 | 0 | 23.297 | 674 | 31.283 | 167 | 36.299 | 322 | 41.315 | 65 | 51.28 | 22 | 53.26 | 13 | 71.278 | 129 | 96.291 | 41 | 121.304 | 24
21.515 | 0 | 23.495 | 1070 | 31.481 | 74 | 36.497 | 311 | 41.513 | 42 | 51.478 | 82 | 53.458 | 63 | 71.476 | 153 | 96.489 | 53 | 121.502 | 57
21.713 | 0 | 23.693 | 1004 | 31.679 | 120 | 36.695 | 186 | 41.711 | 51 | 51.676 | 53 | 53.656 | 35 | 71.674 | 153 | 96.687 | 32 | 121.7 | 58
21.911 | 0 | 23.891 | 775 | 31.877 | 37 | 36.893 | 175 | 41.909 | 34 | 51.874 | 79 | 53.854 | 79 | 71.872 | 141 | 96.885 | 58 | 121.898 | 44
22.109 | 0 | 24.089 | 482 | 32.075 | 40 | 37.091 | 213 | 42.107 | 48 | 52.072 | 57 | 54.052 | 32 | 72.07 | 94 | 97.083 | 36 | 122.096 | 52
22.307 | 0 | 24.287 | 220 | 32.273 | 60 | 37.289 | 155 | 42.305 | 16 | 52.27 | 99 | 54.25 | 61 | 72.268 | 64 | 97.281 | 79 | 122.294 | 34
22.505 | 0 | 24.485 | 160 | 32.471 | 45 | 37.487 | 122 | 42.503 | 31 | 52.468 | 60 | 54.448 | 26 | 72.466 | 85 | 97.479 | 25 | 122.492 | 30
22.703 | 0 | 24.683 | 141 | 32.669 | 55 | 37.685 | 70 | 42.701 | 10 | 52.666 | 42 | 54.646 | 35 | 72.664 | 104 | 97.677 | 93 | 122.69 | 56
22.901 | 0 | 24.881 | 117 | 32.867 | 19 | 37.883 | 92 | 42.899 | 6 | 52.864 | 47 | 54.844 | 10 | 72.862 | 52 | 97.875 | 31 | 122.888 | 29
23.099 | 0 | 25.079 | 214 | 33.065 | 56 | 38.081 | 72 | 43.097 | 18 | 53.062 | 50 | 55.042 | 24 | 73.06 | 137 | 98.073 | 33 | 123.086 | 66
23.297 | 0 | 25.277 | 175 | 33.263 | 40 | 38.279 | 107 | 43.295 | 41 | 53.26 | 29 | 55.24 | 38 | 73.258 | 80 | 98.271 | 21 | 123.284 | 9
23.495 | 0 | 25.475 | 137 | 33.461 | 31 | 38.477 | 42 | 43.493 | 9 | 53.458 | 79 | 55.438 | 5 | 73.456 | 65 | 98.469 | 12 | 123.482 | 9
23.693 | 0 | 25.673 | 96 | 33.659 | 18 | 38.675 | 116 | 43.691 | 9 | 53.656 | 30 | 55.636 | 34 | 73.654 | 82 | 98.667 | 7 | 123.68 | 23
23.891 | 0 | 25.871 | 124 | 33.857 | 29 | 38.873 | 132 | 43.889 | 8 | 53.854 | 58 | 55.834 | 23 | 73.852 | 5 | 98.865 | 2 | 123.878 | 11
24.089 | 0 | 26.069 | 87 | 34.055 | 22 | 39.071 | 111 | 44.087 | 8 | 54.052 | 19 | 56.032 | 8 | 74.05 | 18 | 99.063 | 0 | 124.076 | 26
24.287 | 0 | 26.267 | 68 | 34.253 | 17 | 39.269 | 63 | 44.285 | 9 | 54.25 | 22 | 56.23 | 26 | 74.248 | 25 | 99.261 | 32 | 124.274 | 10
24.485 | 0 | 26.465 | 134 | 34.451 | 7 | 39.467 | 70 | 44.483 | 22 | 54.448 | 41 | 56.428 | 19 | 74.446 | 14 | 99.459 | 9 | 124.472 | 0
24.683 | 0 | 26.663 | 69 | 34.649 | 28 | 39.665 | 44 | 44.681 | 6 | 54.646 | 46 | 56.626 | 46 | 74.644 | 5 | 99.657 | 8 | 124.67 | 0
24.881 | 0 | 26.861 | 74 | 34.847 | 10 | 39.863 | 44 | 44.879 | 15 | 54.844 | 23 | 56.824 | 18 | 74.842 | 10 | 99.855 | 0 | 124.868 | 0
25.079 | 0 | 27.059 | 98 | 35.045 | 0 | 40.061 | 29 | 45.077 | 23 | 55.042 | 32 | 57.022 | 16 | 75.04 | 30 | 100.053 | 4 | 125.066 | 13
25.277 | 0 | 27.257 | 102 | 35.243 | 41 | 40.259 | 52 | 45.275 | 61 | 55.24 | 37 | 57.22 | 20 | 75.238 | 17 | 100.251 | 0 | 125.264 | 0
25.475 | 0 | 27.455 | 51 | 35.441 | 13 | 40.457 | 34 | 45.473 | 10 | 55.438 | 20 | 57.418 | 26 | 75.436 | 31 | 100.449 | 6 | 125.462 | 9
25.673 | 0 | 27.653 | 38 | 35.639 | 7 | 40.655 | 63 | 45.671 | 26 | 55.636 | 5 | 57.616 | 28 | 75.634 | 0 | 100.647 | 0 | 125.66 | 7
25.871 | 0 | 27.851 | 47 | 35.837 | 30 | 40.853 | 40 | 45.869 | 39 | 55.834 | 26 | 57.814 | 9 | 75.832 | 0 | 100.845 | 16 | 125.858 | 16
26.069 | 0 | 28.049 | 43 | 36.035 | 4 | 41.051 | 38 | 46.067 | 8 | 56.032 | 47 | 58.012 | 33 | 76.03 | 19 | 101.043 | 0 | 126.056 | 0
26.267 | 0 | 28.247 | 59 | 36.233 | 15 | 41.249 | 27 | 46.265 | 31 | 56.23 | 8 | 58.21 | 0 | 76.228 | 5 | 101.241 | 7 | 126.254 | 14
26.465 | 0 | 28.445 | 71 | 36.431 | 45 | 41.447 | 7 | 46.463 | 22 | 56.428 | 2 | 58.408 | 20 | 76.426 | 0 | 101.439 | 15 | 126.452 | 0
26.663 | 0 | 28.643 | 37 | 36.629 | 64 | 41.645 | 18 | 46.661 | 0 | 56.626 | 7 | 58.606 | 17 | 76.624 | 30 | 101.637 | 0 | 126.65 | 6
26.861 | 0 | 28.841 | 26 | 36.827 | 68 | 41.843 | 5 | 46.859 | 13 | 56.824 | 16 | 58.804 | 32 | 76.822 | 16 | 101.835 | 0 | 126.848 | 0
27.059 | 0 | 29.039 | 30 | 37.025 | 81 | 42.041 | 13 | 47.056 | 18 | 57.022 | 5 | 59.002 | 47 | 77.02 | 29 | 102.033 | 0 | 127.046 | 0
27.257 | 0 | 29.237 | 46 | 37.223 | 91 | 42.239 | 0 | 47.254 | 19 | 57.22 | 12 | 59.2 | 40 | 77.218 | 9 | 102.231 | 3 | 127.244 | 0
27.455 | 0 | 29.435 | 23 | 37.421 | 112 | 42.437 | 0 | 47.452 | 4 | 57.418 | 5 | 59.398 | 0 | 77.416 | 18 | 102.429 | 6 | 127.442 | 0
27.653 | 0 | 29.633 | 56 | 37.619 | 66 | 42.635 | 0 | 47.65 | 12 | 57.616 | 16 | 59.596 | 5 | 77.614 | 14 | 102.627 | 0 | 127.64 | 10
27.851 | 0 | 29.831 | 37 | 37.817 | 63 | 42.833 | 7 | 47.848 | 37 | 57.814 | 16 | 59.794 | 0 | 77.812 | 13 | 102.825 | 6 | 127.838 | 0
28.049 | 0 | 30.029 | 37 | 38.015 | 36 | 43.031 | 6 | 48.046 | 17 | 58.012 | 0 | 59.992 | 6 | 78.01 | 15 | 103.023 | 6 | 128.036 | 5
28.247 | 0 | 30.227 | 50 | 38.213 | 35 | 43.229 | 6 | 48.244 | 32 | 58.21 | 8 | 60.19 | 0 | 78.208 | 11 | 103.221 | 0 | 128.234 | 7
28.445 | 0 | 30.425 | 60 | 38.411 | 14 | 43.427 | 0 | 48.442 | 5 | 58.408 | 6 | 60.388 | 23 | 78.405 | 6 | 103.419 | 0 | 128.432 | 0
28.643 | 0 | 30.623 | 39 | 38.609 | 8 | 43.625 | 0 | 48.64 | 21 | 58.606 | 5 | 60.586 | 0 | 78.603 | 0 | 103.617 | 7 | 128.63 | 0
28.841 | 0 | 30.821 | 19 | 38.807 | 22 | 43.823 | 6 | 48.838 | 6 | 58.804 | 0 | 60.784 | 11 | 78.801 | 12 | 103.815 | 9 | 128.828 | 0
29.039 | 0 | 31.019 | 47 | 39.005 | 9 | 44.021 | 0 | 49.036 | 18 | 59.002 | 0 | 60.982 | 4 | 78.999 | 8 | 104.013 | 32 | 129.026 | 0
29.237 | 0 | 31.217 | 58 | 39.203 | 28 | 44.219 | 0 | 49.234 | 35 | 59.2 | 6 | 61.18 | 0 | 79.197 | 0 | 104.211 | 16 | 129.224 | 0
29.435 | 0 | 31.415 | 25 | 39.401 | 19 | 44.417 | 4 | 49.432 | 14 | 59.398 | 4 | 61.378 | 8 | 79.395 | 0 | 104.409 | 5 | 129.422 | 17
29.633 | 0 | 31.613 | 67 | 39.599 | 2 | 44.615 | 4 | 49.63 | 18 | 59.596 | 11 | 61.576 | 9 | 79.593 | 0 | 104.607 | 0 | 129.62 | 7
29.831 | 0 | 31.811 | 34 | 39.797 | 13 | 44.813 | 5 | 49.828 | 24 | 59.794 | 15 | 61.774 | 43 | 79.791 | 7 | 104.805 | 11 | 129.818 | 16
30.029 | 0 | 32.009 | 69 | 39.995 | 17 | 45.011 | 0 | 50.026 | 0 | 59.992 | 0 | 61.972 | 16 | 79.989 | 5 | 105.003 | 0 | 130.016 | 11
30.227 | 0 | 32.207 | 34 | 40.193 | 4 | 45.209 | 4 | 50.224 | 19 | 60.19 | 0 | 62.17 | 46 | 80.187 | 0 | 105.201 | 0 | 130.214 | 10
30.425 | 0 | 32.405 | 67 | 40.391 | 0 | 45.407 | 0 | 50.422 | 21 | 60.388 | 10 | 62.368 | 61 | 80.385 | 7 | 105.399 | 0 | 130.412 | 0
30.623 | 0 | 32.603 | 24 | 40.589 | 4 | 45.605 | 0 | 50.62 | 7 | 60.586 | 10 | 62.566 | 39 | 80.583 | 9 | 105.597 | 25 | 130.61 | 16
30.821 | 0 | 32.801 | 17 | 40.787 | 0 | 45.803 | 0 | 50.818 | 5 | 60.784 | 15 | 62.764 | 23 | 80.781 | 2 | 105.795 | 7 | 130.808 | 7
31.019 | 0 | 32.999 | 31 | 40.985 | 9 | 46.001 | 0 | 51.016 | 21 | 60.982 | 10 | 62.962 | 30 | 80.979 | 0 | 105.993 | 9 | 131.006 | 10
31.217 | 0 | 33.197 | 53 | 41.183 | 0 | 46.199 | 0 | 51.214 | 23 | 61.18 | 10 | 63.16 | 44 | 81.177 | 3 | 106.191 | 0 | 131.204 | 40
31.415 | 0 | 33.395 | 40 | 41.381 | 13 | 46.397 | 0 | 51.412 | 51 | 61.378 | 13 | 63.358 | 20 | 81.375 | 0 | 106.389 | 0 | 131.402 | 20
31.613 | 0 | 33.593 | 64 | 41.579 | 8 | 46.595 | 0 | 51.61 | 24 | 61.576 | 13 | 63.556 | 16 | 81.573 | 5 | 106.587 | 0 | 131.6 | 0
31.811 | 0 | 33.791 | 67 | 41.777 | 11 | 46.793 | 0 | 51.808 | 29 | 61.774 | 16 | 63.754 | 8 | 81.771 | 0 | 106.785 | 0 | 131.798 | 3
32.009 | 0 | 33.989 | 33 | 41.975 | 12 | 46.99 | 0 | 52.006 | 53 | 61.972 | 26 | 63.952 | 5 | 81.969 | 0 | 106.983 | 5 | 131.996 | 10
32.207 | 0 | 34.187 | 47 | 42.173 | 9 | 47.188 | 6 | 52.204 | 38 | 62.17 | 39 | 64.15 | 0 | 82.167 | 0 | 107.181 | 0 | 132.194 | 0
32.405 | 0 | 34.385 | 29 | 42.371 | 3 | 47.386 | 0 | 52.402 | 7 | 62.368 | 75 | 64.348 | 5 | 82.365 | 4 | 107.379 | 0 | 132.392 | 10
32.603 | 0 | 34.583 | 12 | 42.569 | 4 | 47.584 | 0 | 52.6 | 6 | 62.566 | 67 | 64.546 | 0 | 82.563 | 5 | 107.577 | 0 | 132.59 | 5
32.801 | 0 | 34.781 | 0 | 42.767 | 0 | 47.782 | 0 | 52.798 | 13 | 62.764 | 73 | 64.744 | 0 | 82.761 | 0 | 107.775 | 0 | 132.788 | 9
32.999 | 0 | 34.979 | 34 | 42.965 | 5 | 47.98 | 0 | 52.996 | 13 | 62.962 | 61 | 64.942 | 0 | 82.959 | 0 | 107.973 | 0 | 132.986 | 8
33.197 | 0 | 35.177 | 14 | 43.163 | 11 | 48.178 | 5 | 53.194 | 5 | 63.16 | 61 | 65.14 | 7 | 83.157 | 13 | 108.171 | 0 | 133.184 | 4
33.395 | 0 | 35.375 | 39 | 43.361 | 0 | 48.376 | 0 | 53.392 | 8 | 63.358 | 32 | 65.338 | 0 | 83.355 | 0 | 108.369 | 0 | 133.382 | 6
33.593 | 0 | 35.573 | 36 | 43.559 | 0 | 48.574 | 0 | 53.59 | 0 | 63.556 | 24 | 65.536 | 0 | 83.553 | 0 | 108.567 | 0 | 133.58 | 0
33.791 | 0 | 35.771 | 7 | 43.757 | 0 | 48.772 | 0 | 53.788 | 0 | 63.754 | 10 | 65.734 | 0 | 83.751 | 13 | 108.765 | 0 | 133.778 | 0
33.989 | 0 | 35.969 | 31 | 43.955 | 6 | 48.97 | 0 | 53.986 | 0 | 63.952 | 3 | 65.932 | 0 | 83.949 | 6 | 108.963 | 7 | 133.976 | 0
34.187 | 0 | 36.167 | 8 | 44.153 | 0 | 49.168 | 9 | 54.184 | 0 | 64.15 | 9 | 66.13 | 14 | 84.147 | 0 | 109.161 | 0 | 134.174 | 0
34.385 | 0 | 36.365 | 22 | 44.351 | 12 | 49.366 | 0 | 54.382 | 0 | 64.348 | 11 | 66.328 | 0 | 84.345 | 0 | 109.359 | 0 | 134.372 | 0
34.583 | 0 | 36.563 | 20 | 44.549 | 0 | 49.564 | 0 | 54.58 | 4 | 64.546 | 0 | 66.526 | 0 | 84.543 | 3 | 109.556 | 0 | 134.57 | 0
34.781 | 0 | 36.761 | 75 | 44.747 | 0 | 49.762 | 0 | 54.778 | 0 | 64.744 | 0 | 66.724 | 0 | 84.741 | 10 | 109.754 | 0 | 134.768 | 0
34.979 | 0 | 36.959 | 62 | 44.945 | 5 | 49.96 | 0 | 54.976 | 0 | 64.942 | 0 | 66.922 | 4 | 84.939 | 0 | 109.952 | 0 | 134.966 | 4
35.177 | 0 | 37.157 | 133 | 45.143 | 0 | 50.158 | 0 | 55.174 | 0 | 65.14 | 0 | 67.12 | 9 | 85.137 | 0 | 110.15 | 4 | 135.164 | 0
35.375 | 0 | 37.355 | 104 | 45.341 | 0 | 50.356 | 13 | 55.372 | 3 | 65.338 | 12 | 67.318 | 0 | 85.335 | 0 | 110.348 | 0 | 135.362 | 0
35.573 | 0 | 37.553 | 85 | 45.539 | 12 | 50.554 | 0 | 55.57 | 0 | 65.536 | 7 | 67.516 | 13 | 85.533 | 0 | 110.546 | 9 | 135.56 | 0
35.771 | 0 | 37.751 | 80 | 45.737 | 11 | 50.752 | 0 | 55.768 | 0 | 65.734 | 0 | 67.714 | 0 | 85.731 | 5 | 110.744 | 3 | 135.758 | 0
35.969 | 0 | 37.949 | 37 | 45.935 | 7 | 50.95 | 0 | 55.966 | 0 | 65.932 | 14 | 67.912 | 0 | 85.929 | 0 | 110.942 | 0 | 135.956 | 0
36.167 | 0 | 38.147 | 58 | 46.133 | 4 | 51.148 | 13 | 56.164 | 0 | 66.13 | 4 | 68.11 | 0 | 86.127 | 0 | 111.14 | 11 | 136.154 | 0
36.365 | 0 | 38.345 | 41 | 46.331 | 0 | 51.346 | 15 | 56.362 | 3 | 66.328 | 0 | 68.308 | 0 | 86.325 | 4 | 111.338 | 0 | 136.352 | 0
36.563 | 0 | 38.543 | 25 | 46.529 | 0 | 51.544 | 47 | 56.56 | 0 | 66.526 | 0 | 68.506 | 0 | 86.523 | 0 | 111.536 | 0 | 136.55 | 0
36.761 | 0 | 38.741 | 6 | 46.727 | 0 | 51.742 | 9 | 56.758 | 0 | 66.724 | 0 | 68.704 | 0 | 86.721 | 0 | 111.734 | 0 | 136.748 | 0
36.959 | 0 | 38.939 | 5 | 46.924 | 0 | 51.94 | 4 | 56.956 | 0 | 66.922 | 5 | 68.902 | 4 | 86.919 | 0 | 111.932 | 0 | 136.946 | 0
37.157 | 0 | 39.137 | 0 | 47.122 | 0 | 52.138 | 19 | 57.154 | 0 | 67.12 | 4 | 69.1 | 22 | 87.117 | 0 | 112.13 | 7 | 137.144 | 0
37.355 | 0 | 39.335 | 6 | 47.32 | 0 | 52.336 | 27 | 57.352 | 0 | 67.318 | 0 | 69.298 | 6 | 87.315 | 0 | 112.328 | 3 | 137.342 | 0
37.553 | 0 | 39.533 | 10 | 47.518 | 12 | 52.534 | 8 | 57.55 | 0 | 67.516 | 0 | 69.496 | 0 | 87.513 | 0 | 112.526 | 0 | 137.54 | 0
37.751 | 0 | 39.731 | 14 | 47.716 | 0 | 52.732 | 26 | 57.748 | 0 | 67.714 | 0 | 69.694 | 3 | 87.711 | 0 | 112.724 | 0 | 137.738 | 0
37.949 | 0 | 39.929 | 5 | 47.914 | 0 | 52.93 | 41 | 57.946 | 0 | 67.912 | 0 | 69.892 | 0 | 87.909 | 0 | 112.922 | 0 | 137.936 | 0
38.147 | 0 | 40.127 | 4 | 48.112 | 0 | 53.128 | 10 | 58.144 | 0 | 68.11 | 0 | 70.09 | 12 | 88.107 | 0 | 113.12 | 0 | 138.134 | 0
38.345 | 0 | 40.325 | 18 | 48.31 | 8 | 53.326 | 0 | 58.342 | 0 | 68.308 | 0 | 70.288 | 3 | 88.305 | 0 | 113.318 | 0 | 138.332 | 0
38.543 | 0 | 40.523 | 9 | 48.508 | 0 | 53.524 | 20 | 58.54 | 0 | 68.506 | 11 | 70.486 | 2 | 88.503 | 0 | 113.516 | 0 | 138.53 | 0
38.741 | 0 | 40.721 | 2 | 48.706 | 13 | 53.722 | 16 | 58.738 | 0 | 68.704 | 6 | 70.684 | 4 | 88.701 | 0 | 113.714 | 0 | 138.728 | 0
38.939 | 0 | 40.919 | 0 | 48.904 | 7 | 53.92 | 0 | 58.936 | 0 | 68.902 | 20 | 70.882 | 16 | 88.899 | 0 | 113.912 | 0 | 138.926 | 0
39.137 | 0 | 41.117 | 18 | 49.102 | 0 | 54.118 | 0 | 59.134 | 0 | 69.1 | 0 | 71.08 | 0 | 89.097 | 0 | 114.11 | 0 | 139.124 | 0
39.335 | 0 | 41.315 | 0 | 49.3 | 0 | 54.316 | 10 | 59.332 | 0 | 69.298 | 0 | 71.278 | 8 | 89.295 | 0 | 114.308 | 5 | 139.322 | 0
39.533 | 0 | 41.513 | 10 | 49.498 | 0 | 54.514 | 0 | 59.53 | 0 | 69.496 | 0 | 71.476 | 0 | 89.493 | 8 | 114.506 | 0 | 139.52 | 0
39.731 | 0 | 41.711 | 5 | 49.696 | 8 | 54.712 | 0 | 59.728 | 0 | 69.694 | 4 | 71.674 | 6 | 89.691 | 12 | 114.704 | 0 | 139.718 | 0
39.929 | 0 | 41.909 | 48 | 49.894 | 0 | 54.91 | 0 | 59.926 | 0 | 69.892 | 0 | 71.872 | 0 | 89.889 | 0 | 114.902 | 0 | 139.916 | 0
40.127 | 0 | 42.107 | 5 | 50.092 | 0 | 55.108 | 0 | 60.124 | 0 | 70.09 | 0 | 72.07 | 0 | 90.087 | 0 | 115.1 | 0 | 140.114 | 0
40.325 | 0 | 42.305 | 9 | 50.29 | 6 | 55.306 | 15 | 60.322 | 0 | 70.288 | 6 | 72.268 | 0 | 90.285 | 0 | 115.298 | 0 | 140.312 | 0
40.523 | 0 | 42.503 | 13 | 50.488 | 0 | 55.504 | 0 | 60.52 | 0 | 70.486 | 0 | 72.466 | 5 | 90.483 | 0 | 115.496 | 0 | 140.51 | 8
40.721 | 0 | 42.701 | 0 | 50.686 | 15 | 55.702 | 0 | 60.718 | 0 | 70.684 | 0 | 72.664 | 0 | 90.681 | 14 | 115.694 | 0 | 140.708 | 0
40.919 | 0 | 42.899 | 0 | 50.884 | 9 | 55.9 | 0 | 60.916 | 0 | 70.882 | 0 | 72.862 | 0 | 90.879 | 0 | 115.892 | 0 | 140.906 | 5
41.117 | 0 | 43.097 | 7 | 51.082 | 16 | 56.098 | 0 | 61.114 | 0 | 71.08 | 25 | 73.06 | 0 | 91.077 | 0 | 116.09 | 0 | 141.103 | 0
41.315 | 0 | 43.295 | 5 | 51.28 | 10 | 56.296 | 0 | 61.312 | 0 | 71.278 | 0 | 73.258 | 5 | 91.275 | 12 | 116.288 | 0 | 141.301 | 0
41.513 | 0 | 43.493 | 18 | 51.478 | 11 | 56.494 | 0 | 61.51 | 0 | 71.476 | 0 | 73.456 | 0 | 91.473 | 8 | 116.486 | 0 | 141.499 | 0
41.711 | 0 | 43.691 | 11 | 51.676 | 21 | 56.692 | 0 | 61.708 | 0 | 71.674 | 0 | 73.654 | 0 | 91.671 | 11 | 116.684 | 0 | 141.697 | 6
41.909 | 0 | 43.889 | 8 | 51.874 | 12 | 56.89 | 0 | 61.906 | 0 | 71.872 | 0 | 73.852 | 4 | 91.869 | 7 | 116.882 | 0 | 141.895 | 0
42.107 | 0 | 44.087 | 0 | 52.072 | 31 | 57.088 | 0 | 62.104 | 0 | 72.07 | 0 | 74.05 | 11 | 92.067 | 19 | 117.08 | 0 | 142.093 | 0
42.305 | 0 | 44.285 | 0 | 52.27 | 13 | 57.286 | 0 | 62.302 | 0 | 72.268 | 0 | 74.248 | 0 | 92.265 | 19 | 117.278 | 0 | 142.291 | 5
42.503 | 0 | 44.483 | 0 | 52.468 | 5 | 57.484 | 0 | 62.5 | 0 | 72.466 | 0 | 74.446 | 0 | 92.463 | 10 | 117.476 | 0 | 142.489 | 5
42.701 | 0 | 44.681 | 0 | 52.666 | 4 | 57.682 | 0 | 62.698 | 0 | 72.664 | 0 | 74.644 | 4 | 92.661 | 9 | 117.674 | 0 | 142.687 | 0
42.899 | 0 | 44.879 | 0 | 52.864 | 15 | 57.88 | 0 | 62.896 | 0 | 72.862 | 0 | 74.842 | 0 | 92.859 | 19 | 117.872 | 0 | 142.885 | 0
43.097 | 0 | 45.077 | 0 | 53.062 | 15 | 58.078 | 0 | 63.094 | 0 | 73.06 | 0 | 75.04 | 0 | 93.057 | 10 | 118.07 | 0 | 143.083 | 0
43.295 | 0 | 45.275 | 0 | 53.26 | 5 | 58.276 | 0 | 63.292 | 0 | 73.258 | 0 | 75.238 | 0 | 93.255 | 7 | 118.268 | 0 | 143.281 | 0
43.493 | 0 | 45.473 | 36 | 53.458 | 12 | 58.474 | 0 | 63.49 | 0 | 73.456 | 0 | 75.436 | 0 | 93.453 | 0 | 118.466 | 0 | 143.479 | 0
43.691 | 0 | 45.671 | 0 | 53.656 | 0 | 58.672 | 0 | 63.688 | 0 | 73.654 | 0 | 75.634 | 0 | 93.651 | 0 | 118.664 | 0 | 143.677 | 0
43.889 | 0 | 45.869 | 0 | 53.854 | 6 | 58.87 | 0 | 63.886 | 0 | 73.852 | 0 | 75.832 | 0 | 93.849 | 4 | 118.862 | 0 | 143.875 | 0
44.087 | 0 | 46.067 | 7 | 54.052 | 4 | 59.068 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 0 | 94.047 | 4 | 119.06 | 0 | 144.073 | 0
44.285 | 0 | 46.265 | 0 | 54.25 | 0 | 59.266 | 0 | 64.282 | 0 | 74.248 | 0 | 76.228 | 0 | 94.245 | 0 | 119.258 | 0 | 144.271 | 0
44.483 | 0 | 46.463 | 0 | 54.448 | 0 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 0 | 94.443 | 4 | 119.456 | 0 | 144.469 | 0
44.681 | 0 | 46.661 | 10 | 54.646 | 4 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 0 | 119.654 | 0 | 144.667 | 0
44.879 | 0 | 46.859 | 0 | 54.844 | 0 | 59.86 | 0 | 64.876 | 0 | 74.842 | 0 | 76.822 | 0 | 94.839 | 6 | 119.852 | 0 | 144.865 | 0
45.077 | 0 | 47.056 | 0 | 55.042 | 4 | 60.058 | 0 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 0 | 120.05 | 0 | 145.063 | 0
45.275 | 0 | 47.254 | 0 | 55.24 | 9 | 60.256 | 0 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 6 | 120.248 | 0 | 145.261 | 0
45.473 | 0 | 47.452 | 5 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 0 | 145.459 | 0
45.671 | 0 | 47.65 | 5 | 55.636 | 0 | 60.652 | 0 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 0 | 145.657 | 0
45.869 | 0 | 47.848 | 0 | 55.834 | 0 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 0 | 120.842 | 0 | 145.855 | 0
46.067 | 0 | 48.046 | 0 | 56.032 | 0 | 61.048 | 0 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 0 | 121.04 | 0 | 146.053 | 0
46.265 | 0 | 48.244 | 0 | 56.23 | 0 | 61.246 | 0 | 66.262 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 0 | 146.251 | 0
46.463 | 0 | 48.442 | 0 | 56.428 | 5 | 61.444 | 0 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 0 | 146.449 | 0
46.661 | 0 | 48.64 | 0 | 56.626 | 0 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 0 | 146.647 | 0
46.859 | 0 | 48.838 | 0 | 56.824 | 0 | 61.84 | 0 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 0 | 146.845 | 0
47.056 | 0 | 49.036 | 0 | 57.022 | 0 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 0 | 122.03 | 0 | 147.043 | 0
47.254 | 0 | 49.234 | 0 | 57.22 | 0 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 0 | 147.241 | 0
47.452 | 0 | 49.432 | 0 | 57.418 | 0 | 62.434 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 0 | 147.439 | 0
47.65 | 0 | 49.63 | 0 | 57.616 | 0 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 0 | 147.637 | 0
47.848 | 0 | 49.828 | 0 | 57.814 | 0 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 0 | 122.822 | 0 | 147.835 | 0
48.046 | 0 | 50.026 | 16 | 58.012 | 0 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 0 | 148.033 | 0
48.244 | 0 | 50.224 | 0 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 0 | 148.231 | 0
48.442 | 0 | 50.422 | 0 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 0 | 148.429 | 0
48.64 | 0 | 50.62 | 0 | 58.606 | 0 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 4 | 123.614 | 0 | 148.627 | 0
48.838 | 0 | 50.818 | 0 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0 | 148.825 | 0
49.036 | 0 | 51.016 | 6 | 59.002 | 0 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 0 | 124.01 | 0 | 149.023 | 0
49.234 | 0 | 51.214 | 4 | 59.2 | 0 | 64.216 | 0 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 0 | 124.208 | 0 | 149.221 | 0
49.432 | 0 | 51.412 | 25 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 0 | 149.419 | 0

## Analyte Val-Tyr-Val
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 85 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 7974 | 12.804 | 0 | 20.789 | 0 | 25.805 | 8 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 39156 | 13.002 | 0 | 20.987 | 127 | 26.003 | 47 | 31.019 | 563 | 40.985 | 98 | 42.965 | 160 | 60.982 | 35 | 85.995 | 82 | 111.008 | 201
11.418 | 451079 | 13.2 | 0 | 21.185 | 637 | 26.201 | 245 | 31.217 | 710 | 41.183 | 334 | 43.163 | 284 | 61.18 | 188 | 86.193 | 226 | 111.206 | 247
11.616 | 731121 | 13.398 | 0 | 21.383 | 536 | 26.399 | 569 | 31.415 | 1087 | 41.381 | 366 | 43.361 | 527 | 61.378 | 477 | 86.391 | 465 | 111.404 | 370
11.814 | 273295 | 13.596 | 0 | 21.581 | 441 | 26.597 | 690 | 31.613 | 1265 | 41.579 | 590 | 43.559 | 576 | 61.576 | 517 | 86.589 | 408 | 111.602 | 442
12.012 | 31806 | 13.794 | 0 | 21.779 | 432 | 26.795 | 354 | 31.811 | 1458 | 41.777 | 344 | 43.757 | 660 | 61.774 | 753 | 86.787 | 627 | 111.8 | 447
12.21 | 10767 | 13.992 | 0 | 21.977 | 713 | 26.993 | 458 | 32.009 | 1684 | 41.975 | 726 | 43.955 | 768 | 61.972 | 744 | 86.985 | 452 | 111.998 | 381
12.408 | 7276 | 14.19 | 0 | 22.175 | 869 | 27.191 | 660 | 32.207 | 1888 | 42.173 | 435 | 44.153 | 583 | 62.17 | 541 | 87.183 | 472 | 112.196 | 351
12.606 | 1477 | 14.388 | 0 | 22.373 | 763 | 27.389 | 721 | 32.405 | 1846 | 42.371 | 328 | 44.351 | 630 | 62.368 | 510 | 87.381 | 351 | 112.394 | 294
12.804 | 130 | 14.586 | 0 | 22.571 | 920 | 27.587 | 857 | 32.603 | 1236 | 42.569 | 512 | 44.549 | 609 | 62.566 | 568 | 87.579 | 442 | 112.592 | 428
13.002 | 49 | 14.784 | 4 | 22.769 | 7720 | 27.785 | 686 | 32.801 | 874 | 42.767 | 501 | 44.747 | 574 | 62.764 | 558 | 87.777 | 420 | 112.79 | 357
13.2 | 53 | 14.982 | 12 | 22.967 | 67308 | 27.983 | 607 | 32.999 | 861 | 42.965 | 533 | 44.945 | 638 | 62.962 | 509 | 87.975 | 597 | 112.988 | 399
13.398 | 0 | 15.18 | 5 | 23.165 | 255317 | 28.181 | 418 | 33.197 | 741 | 43.163 | 387 | 45.143 | 612 | 63.16 | 401 | 88.173 | 384 | 113.186 | 344
13.596 | 0 | 15.378 | 5 | 23.363 | 387761 | 28.379 | 654 | 33.395 | 786 | 43.361 | 639 | 45.341 | 619 | 63.358 | 525 | 88.371 | 388 | 113.384 | 347
13.794 | 0 | 15.576 | 3 | 23.561 | 258426 | 28.577 | 627 | 33.593 | 709 | 43.559 | 555 | 45.539 | 691 | 63.556 | 632 | 88.569 | 453 | 113.582 | 390
13.992 | 0 | 15.773 | 11 | 23.759 | 80532 | 28.775 | 500 | 33.791 | 823 | 43.757 | 651 | 45.737 | 638 | 63.754 | 610 | 88.767 | 385 | 113.78 | 309
14.19 | 0 | 15.971 | 27 | 23.957 | 11488 | 28.973 | 602 | 33.989 | 900 | 43.955 | 697 | 45.935 | 635 | 63.952 | 682 | 88.965 | 467 | 113.978 | 304
14.388 | 0 | 16.169 | 27 | 24.155 | 1608 | 29.171 | 775 | 34.187 | 895 | 44.153 | 622 | 46.133 | 638 | 64.15 | 683 | 89.163 | 299 | 114.176 | 343
14.586 | 0 | 16.367 | 22 | 24.353 | 1140 | 29.369 | 1034 | 34.385 | 1155 | 44.351 | 499 | 46.331 | 2157 | 64.348 | 606 | 89.361 | 451 | 114.374 | 393
14.784 | 0 | 16.565 | 67 | 24.551 | 1283 | 29.567 | 782 | 34.583 | 4948 | 44.549 | 806 | 46.529 | 9325 | 64.546 | 645 | 89.559 | 357 | 114.572 | 443
14.982 | 0 | 16.763 | 69 | 24.749 | 1332 | 29.765 | 857 | 34.781 | 26362 | 44.747 | 593 | 46.727 | 32344 | 64.744 | 651 | 89.757 | 371 | 114.77 | 330
15.18 | 0 | 16.961 | 58 | 24.947 | 1092 | 29.963 | 750 | 34.979 | 94969 | 44.945 | 757 | 46.924 | 77839 | 64.942 | 688 | 89.955 | 470 | 114.968 | 556
15.378 | 0 | 17.159 | 59 | 25.145 | 729 | 30.161 | 806 | 35.177 | 195904 | 45.143 | 523 | 47.122 | 128126 | 65.14 | 641 | 90.153 | 392 | 115.166 | 457
15.576 | 0 | 17.357 | 61 | 25.343 | 643 | 30.359 | 810 | 35.375 | 238731 | 45.341 | 625 | 47.32 | 151362 | 65.338 | 739 | 90.351 | 553 | 115.364 | 660
15.773 | 0 | 17.555 | 366 | 25.541 | 483 | 30.557 | 879 | 35.573 | 187575 | 45.539 | 566 | 47.518 | 125238 | 65.536 | 706 | 90.549 | 524 | 115.562 | 391
15.971 | 0 | 17.753 | 516 | 25.739 | 456 | 30.755 | 756 | 35.771 | 90962 | 45.737 | 611 | 47.716 | 76140 | 65.734 | 825 | 90.747 | 369 | 115.76 | 492
16.169 | 0 | 17.951 | 1432 | 25.937 | 598 | 30.953 | 983 | 35.969 | 28433 | 45.935 | 454 | 47.914 | 32926 | 65.932 | 604 | 90.945 | 391 | 115.958 | 332
16.367 | 0 | 18.149 | 634 | 26.135 | 592 | 31.151 | 1219 | 36.167 | 6606 | 46.133 | 733 | 48.112 | 10811 | 66.13 | 576 | 91.143 | 423 | 116.156 | 320
16.565 | 0 | 18.347 | 247 | 26.333 | 718 | 31.349 | 1255 | 36.365 | 2514 | 46.331 | 1451 | 48.31 | 2669 | 66.328 | 617 | 91.341 | 501 | 116.354 | 360
16.763 | 0 | 18.545 | 170 | 26.531 | 670 | 31.547 | 1063 | 36.563 | 1267 | 46.529 | 6127 | 48.508 | 955 | 66.526 | 621 | 91.539 | 405 | 116.552 | 301
16.961 | 0 | 18.743 | 283 | 26.729 | 476 | 31.745 | 1133 | 36.761 | 1182 | 46.727 | 24901 | 48.706 | 622 | 66.724 | 456 | 91.737 | 443 | 116.75 | 389
17.159 | 0 | 18.941 | 174 | 26.927 | 416 | 31.943 | 1515 | 36.959 | 1378 | 46.924 | 66774 | 48.904 | 413 | 66.922 | 576 | 91.935 | 499 | 116.948 | 339
17.357 | 0 | 19.139 | 195 | 27.125 | 536 | 32.141 | 1804 | 37.157 | 1302 | 47.122 | 124337 | 49.102 | 501 | 67.12 | 640 | 92.133 | 476 | 117.146 | 269
17.555 | 0 | 19.337 | 183 | 27.323 | 551 | 32.339 | 2032 | 37.355 | 1476 | 47.32 | 158377 | 49.3 | 428 | 67.318 | 459 | 92.331 | 498 | 117.344 | 447
17.753 | 0 | 19.535 | 159 | 27.521 | 792 | 32.537 | 1494 | 37.553 | 1552 | 47.518 | 147283 | 49.498 | 463 | 67.516 | 443 | 92.529 | 481 | 117.542 | 652
17.951 | 0 | 19.733 | 299 | 27.719 | 817 | 32.735 | 1535 | 37.751 | 1134 | 47.716 | 98643 | 49.696 | 425 | 67.714 | 431 | 92.727 | 525 | 117.74 | 1403
18.149 | 0 | 19.931 | 345 | 27.917 | 956 | 32.933 | 979 | 37.949 | 1183 | 47.914 | 47739 | 49.894 | 351 | 67.912 | 501 | 92.925 | 447 | 117.938 | 2592
18.347 | 0 | 20.129 | 363 | 28.115 | 711 | 33.131 | 736 | 38.147 | 1183 | 48.112 | 17007 | 50.092 | 418 | 68.11 | 450 | 93.123 | 432 | 118.136 | 5369
18.545 | 0 | 20.327 | 399 | 28.313 | 553 | 33.329 | 929 | 38.345 | 1263 | 48.31 | 4804 | 50.29 | 449 | 68.308 | 369 | 93.321 | 416 | 118.334 | 9694
18.743 | 0 | 20.525 | 661 | 28.511 | 586 | 33.527 | 737 | 38.543 | 1210 | 48.508 | 1290 | 50.488 | 494 | 68.506 | 342 | 93.519 | 522 | 118.532 | 15559
18.941 | 0 | 20.723 | 424 | 28.709 | 587 | 33.725 | 683 | 38.741 | 1025 | 48.706 | 562 | 50.686 | 458 | 68.704 | 408 | 93.717 | 682 | 118.73 | 22726
19.139 | 0 | 20.921 | 287 | 28.907 | 744 | 33.923 | 716 | 38.939 | 820 | 48.904 | 517 | 50.884 | 432 | 68.902 | 481 | 93.915 | 1462 | 118.928 | 27015
19.337 | 0 | 21.119 | 328 | 29.105 | 761 | 34.121 | 644 | 39.137 | 778 | 49.102 | 522 | 51.082 | 478 | 69.1 | 329 | 94.113 | 3403 | 119.126 | 29473
19.535 | 0 | 21.317 | 515 | 29.303 | 680 | 34.319 | 833 | 39.335 | 502 | 49.3 | 573 | 51.28 | 468 | 69.298 | 405 | 94.311 | 7320 | 119.324 | 28891
19.733 | 0 | 21.515 | 422 | 29.501 | 819 | 34.517 | 1140 | 39.533 | 505 | 49.498 | 482 | 51.478 | 568 | 69.496 | 355 | 94.509 | 14563 | 119.522 | 25226
19.931 | 0 | 21.713 | 383 | 29.699 | 1227 | 34.715 | 4686 | 39.731 | 603 | 49.696 | 382 | 51.676 | 439 | 69.694 | 434 | 94.707 | 25172 | 119.72 | 19015
20.129 | 0 | 21.911 | 423 | 29.897 | 1015 | 34.913 | 25908 | 39.929 | 519 | 49.894 | 468 | 51.874 | 360 | 69.892 | 600 | 94.905 | 36238 | 119.918 | 13382
20.327 | 0 | 22.109 | 568 | 30.095 | 700 | 35.111 | 97198 | 40.127 | 512 | 50.092 | 296 | 52.072 | 512 | 70.09 | 1086 | 95.103 | 44158 | 120.116 | 8629
20.525 | 0 | 22.307 | 724 | 30.293 | 682 | 35.309 | 209371 | 40.325 | 587 | 50.29 | 304 | 52.27 | 530 | 70.288 | 2761 | 95.301 | 46787 | 120.314 | 4470
20.723 | 0 | 22.505 | 766 | 30.491 | 703 | 35.507 | 265272 | 40.523 | 474 | 50.488 | 403 | 52.468 | 676 | 70.486 | 7960 | 95.499 | 42448 | 120.512 | 1942
20.921 | 0 | 22.703 | 822 | 30.689 | 678 | 35.705 | 207601 | 40.721 | 417 | 50.686 | 533 | 52.666 | 846 | 70.684 | 19904 | 95.697 | 32405 | 120.71 | 1057
21.119 | 0 | 22.901 | 2426 | 30.887 | 587 | 35.903 | 105265 | 40.919 | 757 | 50.884 | 471 | 52.864 | 1024 | 70.882 | 40736 | 95.895 | 21516 | 120.908 | 687
21.317 | 0 | 23.099 | 26660 | 31.085 | 661 | 36.101 | 33518 | 41.117 | 768 | 51.082 | 614 | 53.062 | 1051 | 71.08 | 64464 | 96.093 | 12720 | 121.106 | 439
21.515 | 0 | 23.297 | 171483 | 31.283 | 507 | 36.299 | 8042 | 41.315 | 582 | 51.28 | 478 | 53.26 | 914 | 71.278 | 79394 | 96.291 | 6241 | 121.304 | 275
21.713 | 0 | 23.495 | 400399 | 31.481 | 502 | 36.497 | 2412 | 41.513 | 518 | 51.478 | 427 | 53.458 | 999 | 71.476 | 80577 | 96.489 | 2584 | 121.502 | 345
21.911 | 0 | 23.693 | 402414 | 31.679 | 524 | 36.695 | 1447 | 41.711 | 373 | 51.676 | 409 | 53.656 | 851 | 71.674 | 68820 | 96.687 | 1062 | 121.7 | 385
22.109 | 0 | 23.891 | 180836 | 31.877 | 720 | 36.893 | 1129 | 41.909 | 319 | 51.874 | 289 | 53.854 | 503 | 71.872 | 45909 | 96.885 | 633 | 121.898 | 257
22.307 | 0 | 24.089 | 37078 | 32.075 | 1135 | 37.091 | 1245 | 42.107 | 417 | 52.072 | 418 | 54.052 | 387 | 72.07 | 25774 | 97.083 | 371 | 122.096 | 212
22.505 | 0 | 24.287 | 3767 | 32.273 | 1523 | 37.289 | 1146 | 42.305 | 168 | 52.27 | 511 | 54.25 | 325 | 72.268 | 10860 | 97.281 | 278 | 122.294 | 245
22.703 | 0 | 24.485 | 749 | 32.471 | 1964 | 37.487 | 1199 | 42.503 | 262 | 52.468 | 628 | 54.448 | 227 | 72.466 | 4479 | 97.479 | 341 | 122.492 | 194
22.901 | 0 | 24.683 | 804 | 32.669 | 1878 | 37.685 | 1133 | 42.701 | 250 | 52.666 | 667 | 54.646 | 232 | 72.664 | 1640 | 97.677 | 339 | 122.69 | 172
23.099 | 0 | 24.881 | 1055 | 32.867 | 1547 | 37.883 | 1210 | 42.899 | 301 | 52.864 | 848 | 54.844 | 313 | 72.862 | 878 | 97.875 | 297 | 122.888 | 176
23.297 | 0 | 25.079 | 937 | 33.065 | 778 | 38.081 | 1240 | 43.097 | 199 | 53.062 | 1024 | 55.042 | 237 | 73.06 | 609 | 98.073 | 242 | 123.086 | 165
23.495 | 0 | 25.277 | 864 | 33.263 | 464 | 38.279 | 1084 | 43.295 | 237 | 53.26 | 1048 | 55.24 | 322 | 73.258 | 520 | 98.271 | 191 | 123.284 | 208
23.693 | 0 | 25.475 | 565 | 33.461 | 383 | 38.477 | 1065 | 43.493 | 238 | 53.458 | 900 | 55.438 | 394 | 73.456 | 640 | 98.469 | 179 | 123.482 | 167
23.891 | 0 | 25.673 | 439 | 33.659 | 538 | 38.675 | 895 | 43.691 | 395 | 53.656 | 695 | 55.636 | 371 | 73.654 | 716 | 98.667 | 151 | 123.68 | 199
24.089 | 0 | 25.871 | 426 | 33.857 | 290 | 38.873 | 965 | 43.889 | 444 | 53.854 | 739 | 55.834 | 374 | 73.852 | 556 | 98.865 | 166 | 123.878 | 183
24.287 | 0 | 26.069 | 663 | 34.055 | 432 | 39.071 | 842 | 44.087 | 479 | 54.052 | 428 | 56.032 | 316 | 74.05 | 663 | 99.063 | 220 | 124.076 | 176
24.485 | 0 | 26.267 | 683 | 34.253 | 326 | 39.269 | 635 | 44.285 | 307 | 54.25 | 366 | 56.23 | 419 | 74.248 | 452 | 99.261 | 231 | 124.274 | 146
24.683 | 0 | 26.465 | 770 | 34.451 | 377 | 39.467 | 591 | 44.483 | 342 | 54.448 | 372 | 56.428 | 262 | 74.446 | 284 | 99.459 | 103 | 124.472 | 186
24.881 | 0 | 26.663 | 663 | 34.649 | 468 | 39.665 | 566 | 44.681 | 274 | 54.646 | 327 | 56.626 | 243 | 74.644 | 350 | 99.657 | 162 | 124.67 | 195
25.079 | 0 | 26.861 | 589 | 34.847 | 388 | 39.863 | 673 | 44.879 | 229 | 54.844 | 225 | 56.824 | 353 | 74.842 | 275 | 99.855 | 134 | 124.868 | 122
25.277 | 0 | 27.059 | 414 | 35.045 | 597 | 40.061 | 395 | 45.077 | 175 | 55.042 | 226 | 57.022 | 290 | 75.04 | 301 | 100.053 | 265 | 125.066 | 81
25.475 | 0 | 27.257 | 425 | 35.243 | 606 | 40.259 | 374 | 45.275 | 212 | 55.24 | 280 | 57.22 | 238 | 75.238 | 256 | 100.251 | 149 | 125.264 | 204
25.673 | 0 | 27.455 | 595 | 35.441 | 630 | 40.457 | 409 | 45.473 | 220 | 55.438 | 282 | 57.418 | 242 | 75.436 | 177 | 100.449 | 163 | 125.462 | 154
25.871 | 0 | 27.653 | 829 | 35.639 | 528 | 40.655 | 523 | 45.671 | 259 | 55.636 | 391 | 57.616 | 297 | 75.634 | 165 | 100.647 | 170 | 125.66 | 209
26.069 | 0 | 27.851 | 995 | 35.837 | 760 | 40.853 | 482 | 45.869 | 150 | 55.834 | 280 | 57.814 | 236 | 75.832 | 131 | 100.845 | 228 | 125.858 | 173
26.267 | 0 | 28.049 | 1076 | 36.035 | 752 | 41.051 | 394 | 46.067 | 232 | 56.032 | 306 | 58.012 | 267 | 76.03 | 172 | 101.043 | 213 | 126.056 | 218
26.465 | 0 | 28.247 | 873 | 36.233 | 958 | 41.249 | 448 | 46.265 | 340 | 56.23 | 290 | 58.21 | 286 | 76.228 | 239 | 101.241 | 207 | 126.254 | 80
26.663 | 0 | 28.445 | 666 | 36.431 | 984 | 41.447 | 195 | 46.463 | 213 | 56.428 | 263 | 58.408 | 317 | 76.426 | 229 | 101.439 | 150 | 126.452 | 119
26.861 | 0 | 28.643 | 501 | 36.629 | 1080 | 41.645 | 279 | 46.661 | 257 | 56.626 | 213 | 58.606 | 329 | 76.624 | 187 | 101.637 | 200 | 126.65 | 127
27.059 | 0 | 28.841 | 521 | 36.827 | 1250 | 41.843 | 194 | 46.859 | 283 | 56.824 | 190 | 58.804 | 274 | 76.822 | 219 | 101.835 | 160 | 126.848 | 147
27.257 | 0 | 29.039 | 603 | 37.025 | 1099 | 42.041 | 166 | 47.056 | 226 | 57.022 | 110 | 59.002 | 309 | 77.02 | 168 | 102.033 | 174 | 127.046 | 144
27.455 | 0 | 29.237 | 578 | 37.223 | 1175 | 42.239 | 111 | 47.254 | 142 | 57.22 | 151 | 59.2 | 233 | 77.218 | 189 | 102.231 | 184 | 127.244 | 89
27.653 | 0 | 29.435 | 573 | 37.421 | 1148 | 42.437 | 137 | 47.452 | 279 | 57.418 | 223 | 59.398 | 285 | 77.416 | 177 | 102.429 | 147 | 127.442 | 123
27.851 | 0 | 29.633 | 537 | 37.619 | 1047 | 42.635 | 99 | 47.65 | 366 | 57.616 | 184 | 59.596 | 276 | 77.614 | 110 | 102.627 | 128 | 127.64 | 160
28.049 | 0 | 29.831 | 651 | 37.817 | 1056 | 42.833 | 151 | 47.848 | 175 | 57.814 | 146 | 59.794 | 164 | 77.812 | 139 | 102.825 | 168 | 127.838 | 121
28.247 | 0 | 30.029 | 747 | 38.015 | 972 | 43.031 | 87 | 48.046 | 261 | 58.012 | 163 | 59.992 | 294 | 78.01 | 130 | 103.023 | 111 | 128.036 | 135
28.445 | 0 | 30.227 | 579 | 38.213 | 919 | 43.229 | 114 | 48.244 | 270 | 58.21 | 198 | 60.19 | 341 | 78.208 | 146 | 103.221 | 198 | 128.234 | 120
28.643 | 0 | 30.425 | 572 | 38.411 | 909 | 43.427 | 106 | 48.442 | 245 | 58.408 | 287 | 60.388 | 292 | 78.405 | 212 | 103.419 | 199 | 128.432 | 182
28.841 | 0 | 30.623 | 656 | 38.609 | 683 | 43.625 | 123 | 48.64 | 256 | 58.606 | 203 | 60.586 | 342 | 78.603 | 84 | 103.617 | 142 | 128.63 | 114
29.039 | 0 | 30.821 | 637 | 38.807 | 712 | 43.823 | 109 | 48.838 | 221 | 58.804 | 224 | 60.784 | 327 | 78.801 | 157 | 103.815 | 155 | 128.828 | 111
29.237 | 0 | 31.019 | 591 | 39.005 | 648 | 44.021 | 86 | 49.036 | 226 | 59.002 | 243 | 60.982 | 381 | 78.999 | 146 | 104.013 | 99 | 129.026 | 134
29.435 | 0 | 31.217 | 754 | 39.203 | 730 | 44.219 | 116 | 49.234 | 210 | 59.2 | 252 | 61.18 | 514 | 79.197 | 124 | 104.211 | 147 | 129.224 | 99
29.633 | 0 | 31.415 | 655 | 39.401 | 550 | 44.417 | 88 | 49.432 | 174 | 59.398 | 220 | 61.378 | 364 | 79.395 | 209 | 104.409 | 83 | 129.422 | 81
29.831 | 0 | 31.613 | 669 | 39.599 | 397 | 44.615 | 63 | 49.63 | 257 | 59.596 | 233 | 61.576 | 325 | 79.593 | 65 | 104.607 | 130 | 129.62 | 232
30.029 | 0 | 31.811 | 625 | 39.797 | 393 | 44.813 | 70 | 49.828 | 224 | 59.794 | 189 | 61.774 | 318 | 79.791 | 51 | 104.805 | 101 | 129.818 | 114
30.227 | 0 | 32.009 | 875 | 39.995 | 369 | 45.011 | 121 | 50.026 | 246 | 59.992 | 144 | 61.972 | 409 | 79.989 | 112 | 105.003 | 183 | 130.016 | 112
30.425 | 0 | 32.207 | 1209 | 40.193 | 227 | 45.209 | 61 | 50.224 | 197 | 60.19 | 161 | 62.17 | 291 | 80.187 | 121 | 105.201 | 125 | 130.214 | 130
30.623 | 0 | 32.405 | 1845 | 40.391 | 273 | 45.407 | 70 | 50.422 | 211 | 60.388 | 305 | 62.368 | 294 | 80.385 | 149 | 105.399 | 46 | 130.412 | 120
30.821 | 0 | 32.603 | 2367 | 40.589 | 228 | 45.605 | 54 | 50.62 | 222 | 60.586 | 309 | 62.566 | 309 | 80.583 | 98 | 105.597 | 89 | 130.61 | 94
31.019 | 0 | 32.801 | 2028 | 40.787 | 306 | 45.803 | 53 | 50.818 | 224 | 60.784 | 382 | 62.764 | 294 | 80.781 | 113 | 105.795 | 75 | 130.808 | 117
31.217 | 0 | 32.999 | 1802 | 40.985 | 279 | 46.001 | 97 | 51.016 | 113 | 60.982 | 346 | 62.962 | 342 | 80.979 | 103 | 105.993 | 124 | 131.006 | 57
31.415 | 0 | 33.197 | 1171 | 41.183 | 124 | 46.199 | 44 | 51.214 | 147 | 61.18 | 361 | 63.16 | 331 | 81.177 | 122 | 106.191 | 61 | 131.204 | 107
31.613 | 0 | 33.395 | 563 | 41.381 | 162 | 46.397 | 17 | 51.412 | 85 | 61.378 | 364 | 63.358 | 419 | 81.375 | 133 | 106.389 | 89 | 131.402 | 63
31.811 | 0 | 33.593 | 563 | 41.579 | 178 | 46.595 | 35 | 51.61 | 91 | 61.576 | 469 | 63.556 | 305 | 81.573 | 115 | 106.587 | 50 | 131.6 | 37
32.009 | 0 | 33.791 | 561 | 41.777 | 87 | 46.793 | 34 | 51.808 | 42 | 61.774 | 403 | 63.754 | 372 | 81.771 | 123 | 106.785 | 34 | 131.798 | 60
32.207 | 0 | 33.989 | 482 | 41.975 | 167 | 46.99 | 30 | 52.006 | 49 | 61.972 | 442 | 63.952 | 295 | 81.969 | 137 | 106.983 | 101 | 131.996 | 54
32.405 | 0 | 34.187 | 441 | 42.173 | 99 | 47.188 | 54 | 52.204 | 34 | 62.17 | 342 | 64.15 | 306 | 82.167 | 129 | 107.181 | 67 | 132.194 | 16
32.603 | 0 | 34.385 | 467 | 42.371 | 125 | 47.386 | 55 | 52.402 | 32 | 62.368 | 366 | 64.348 | 281 | 82.365 | 140 | 107.379 | 76 | 132.392 | 61
32.801 | 0 | 34.583 | 390 | 42.569 | 91 | 47.584 | 31 | 52.6 | 14 | 62.566 | 350 | 64.546 | 200 | 82.563 | 147 | 107.577 | 93 | 132.59 | 51
32.999 | 0 | 34.781 | 594 | 42.767 | 105 | 47.782 | 23 | 52.798 | 9 | 62.764 | 367 | 64.744 | 371 | 82.761 | 124 | 107.775 | 80 | 132.788 | 36
33.197 | 0 | 34.979 | 628 | 42.965 | 92 | 47.98 | 26 | 52.996 | 6 | 62.962 | 435 | 64.942 | 281 | 82.959 | 119 | 107.973 | 44 | 132.986 | 67
33.395 | 0 | 35.177 | 660 | 43.163 | 93 | 48.178 | 34 | 53.194 | 13 | 63.16 | 377 | 65.14 | 264 | 83.157 | 139 | 108.171 | 52 | 133.184 | 81
33.593 | 0 | 35.375 | 601 | 43.361 | 58 | 48.376 | 24 | 53.392 | 8 | 63.358 | 336 | 65.338 | 289 | 83.355 | 104 | 108.369 | 60 | 133.382 | 37
33.791 | 0 | 35.573 | 505 | 43.559 | 126 | 48.574 | 38 | 53.59 | 0 | 63.556 | 412 | 65.536 | 340 | 83.553 | 101 | 108.567 | 92 | 133.58 | 111
33.989 | 0 | 35.771 | 779 | 43.757 | 100 | 48.772 | 25 | 53.788 | 4 | 63.754 | 394 | 65.734 | 234 | 83.751 | 93 | 108.765 | 78 | 133.778 | 70
34.187 | 0 | 35.969 | 676 | 43.955 | 86 | 48.97 | 7 | 53.986 | 0 | 63.952 | 409 | 65.932 | 190 | 83.949 | 72 | 108.963 | 127 | 133.976 | 124
34.385 | 0 | 36.167 | 857 | 44.153 | 101 | 49.168 | 15 | 54.184 | 5 | 64.15 | 304 | 66.13 | 179 | 84.147 | 164 | 109.161 | 140 | 134.174 | 39
34.583 | 0 | 36.365 | 953 | 44.351 | 72 | 49.366 | 9 | 54.382 | 0 | 64.348 | 338 | 66.328 | 330 | 84.345 | 95 | 109.359 | 112 | 134.372 | 90
34.781 | 0 | 36.563 | 1054 | 44.549 | 82 | 49.564 | 8 | 54.58 | 0 | 64.546 | 308 | 66.526 | 169 | 84.543 | 128 | 109.556 | 80 | 134.57 | 126
34.979 | 0 | 36.761 | 1325 | 44.747 | 64 | 49.762 | 18 | 54.778 | 16 | 64.744 | 312 | 66.724 | 175 | 84.741 | 175 | 109.754 | 77 | 134.768 | 103
35.177 | 0 | 36.959 | 1309 | 44.945 | 48 | 49.96 | 2 | 54.976 | 5 | 64.942 | 265 | 66.922 | 214 | 84.939 | 153 | 109.952 | 140 | 134.966 | 80
35.375 | 0 | 37.157 | 1294 | 45.143 | 118 | 50.158 | 4 | 55.174 | 0 | 65.14 | 299 | 67.12 | 137 | 85.137 | 130 | 110.15 | 121 | 135.164 | 105
35.573 | 0 | 37.355 | 1181 | 45.341 | 53 | 50.356 | 8 | 55.372 | 6 | 65.338 | 290 | 67.318 | 167 | 85.335 | 223 | 110.348 | 103 | 135.362 | 43
35.771 | 0 | 37.553 | 1270 | 45.539 | 56 | 50.554 | 0 | 55.57 | 5 | 65.536 | 308 | 67.516 | 129 | 85.533 | 161 | 110.546 | 99 | 135.56 | 75
35.969 | 0 | 37.751 | 1121 | 45.737 | 30 | 50.752 | 6 | 55.768 | 5 | 65.734 | 244 | 67.714 | 160 | 85.731 | 290 | 110.744 | 99 | 135.758 | 96
36.167 | 0 | 37.949 | 1045 | 45.935 | 80 | 50.95 | 4 | 55.966 | 0 | 65.932 | 258 | 67.912 | 149 | 85.929 | 167 | 110.942 | 120 | 135.956 | 90
36.365 | 0 | 38.147 | 1101 | 46.133 | 160 | 51.148 | 3 | 56.164 | 0 | 66.13 | 253 | 68.11 | 81 | 86.127 | 209 | 111.14 | 142 | 136.154 | 92
36.563 | 0 | 38.345 | 1154 | 46.331 | 46 | 51.346 | 0 | 56.362 | 0 | 66.328 | 190 | 68.308 | 122 | 86.325 | 140 | 111.338 | 113 | 136.352 | 47
36.761 | 0 | 38.543 | 1132 | 46.529 | 41 | 51.544 | 13 | 56.56 | 8 | 66.526 | 218 | 68.506 | 81 | 86.523 | 99 | 111.536 | 65 | 136.55 | 37
36.959 | 0 | 38.741 | 878 | 46.727 | 74 | 51.742 | 0 | 56.758 | 4 | 66.724 | 155 | 68.704 | 85 | 86.721 | 63 | 111.734 | 35 | 136.748 | 27
37.157 | 0 | 38.939 | 727 | 46.924 | 15 | 51.94 | 0 | 56.956 | 0 | 66.922 | 216 | 68.902 | 90 | 86.919 | 67 | 111.932 | 28 | 136.946 | 39
37.355 | 0 | 39.137 | 791 | 47.122 | 37 | 52.138 | 0 | 57.154 | 0 | 67.12 | 167 | 69.1 | 118 | 87.117 | 30 | 112.13 | 17 | 137.144 | 20
37.553 | 0 | 39.335 | 728 | 47.32 | 71 | 52.336 | 5 | 57.352 | 0 | 67.318 | 149 | 69.298 | 140 | 87.315 | 12 | 112.328 | 18 | 137.342 | 13
37.751 | 0 | 39.533 | 611 | 47.518 | 41 | 52.534 | 10 | 57.55 | 0 | 67.516 | 160 | 69.496 | 111 | 87.513 | 0 | 112.526 | 26 | 137.54 | 4
37.949 | 0 | 39.731 | 606 | 47.716 | 45 | 52.732 | 0 | 57.748 | 0 | 67.714 | 139 | 69.694 | 65 | 87.711 | 6 | 112.724 | 5 | 137.738 | 22
38.147 | 0 | 39.929 | 501 | 47.914 | 37 | 52.93 | 13 | 57.946 | 0 | 67.912 | 155 | 69.892 | 73 | 87.909 | 0 | 112.922 | 8 | 137.936 | 5
38.345 | 0 | 40.127 | 421 | 48.112 | 9 | 53.128 | 11 | 58.144 | 0 | 68.11 | 117 | 70.09 | 137 | 88.107 | 8 | 113.12 | 2 | 138.134 | 0
38.543 | 0 | 40.325 | 318 | 48.31 | 9 | 53.326 | 9 | 58.342 | 0 | 68.308 | 104 | 70.288 | 66 | 88.305 | 0 | 113.318 | 3 | 138.332 | 12
38.741 | 0 | 40.523 | 300 | 48.508 | 22 | 53.524 | 0 | 58.54 | 0 | 68.506 | 147 | 70.486 | 126 | 88.503 | 4 | 113.516 | 19 | 138.53 | 18
38.939 | 0 | 40.721 | 319 | 48.706 | 16 | 53.722 | 5 | 58.738 | 0 | 68.704 | 116 | 70.684 | 106 | 88.701 | 0 | 113.714 | 0 | 138.728 | 2
39.137 | 0 | 40.919 | 214 | 48.904 | 29 | 53.92 | 0 | 58.936 | 0 | 68.902 | 65 | 70.882 | 93 | 88.899 | 7 | 113.912 | 0 | 138.926 | 6
39.335 | 0 | 41.117 | 299 | 49.102 | 35 | 54.118 | 0 | 59.134 | 0 | 69.1 | 98 | 71.08 | 56 | 89.097 | 0 | 114.11 | 3 | 139.124 | 17
39.533 | 0 | 41.315 | 180 | 49.3 | 3 | 54.316 | 5 | 59.332 | 0 | 69.298 | 81 | 71.278 | 47 | 89.295 | 0 | 114.308 | 24 | 139.322 | 8
39.731 | 0 | 41.513 | 219 | 49.498 | 11 | 54.514 | 7 | 59.53 | 11 | 69.496 | 140 | 71.476 | 90 | 89.493 | 9 | 114.506 | 0 | 139.52 | 8
39.929 | 0 | 41.711 | 188 | 49.696 | 0 | 54.712 | 0 | 59.728 | 0 | 69.694 | 95 | 71.674 | 41 | 89.691 | 5 | 114.704 | 10 | 139.718 | 0
40.127 | 0 | 41.909 | 166 | 49.894 | 30 | 54.91 | 0 | 59.926 | 0 | 69.892 | 96 | 71.872 | 49 | 89.889 | 0 | 114.902 | 22 | 139.916 | 3
40.325 | 0 | 42.107 | 169 | 50.092 | 0 | 55.108 | 0 | 60.124 | 0 | 70.09 | 104 | 72.07 | 56 | 90.087 | 0 | 115.1 | 8 | 140.114 | 0
40.523 | 0 | 42.305 | 102 | 50.29 | 13 | 55.306 | 5 | 60.322 | 0 | 70.288 | 86 | 72.268 | 61 | 90.285 | 5 | 115.298 | 5 | 140.312 | 11
40.721 | 0 | 42.503 | 108 | 50.488 | 0 | 55.504 | 0 | 60.52 | 0 | 70.486 | 89 | 72.466 | 63 | 90.483 | 0 | 115.496 | 17 | 140.51 | 4
40.919 | 0 | 42.701 | 136 | 50.686 | 0 | 55.702 | 0 | 60.718 | 0 | 70.684 | 82 | 72.664 | 58 | 90.681 | 7 | 115.694 | 6 | 140.708 | 0
41.117 | 0 | 42.899 | 101 | 50.884 | 5 | 55.9 | 0 | 60.916 | 0 | 70.882 | 51 | 72.862 | 19 | 90.879 | 0 | 115.892 | 0 | 140.906 | 0
41.315 | 0 | 43.097 | 105 | 51.082 | 0 | 56.098 | 0 | 61.114 | 0 | 71.08 | 62 | 73.06 | 22 | 91.077 | 0 | 116.09 | 3 | 141.103 | 3
41.513 | 0 | 43.295 | 95 | 51.28 | 5 | 56.296 | 0 | 61.312 | 0 | 71.278 | 54 | 73.258 | 14 | 91.275 | 7 | 116.288 | 8 | 141.301 | 0
41.711 | 0 | 43.493 | 56 | 51.478 | 0 | 56.494 | 0 | 61.51 | 0 | 71.476 | 70 | 73.456 | 70 | 91.473 | 0 | 116.486 | 22 | 141.499 | 7
41.909 | 0 | 43.691 | 63 | 51.676 | 12 | 56.692 | 5 | 61.708 | 0 | 71.674 | 18 | 73.654 | 14 | 91.671 | 0 | 116.684 | 0 | 141.697 | 11
42.107 | 0 | 43.889 | 130 | 51.874 | 3 | 56.89 | 0 | 61.906 | 0 | 71.872 | 12 | 73.852 | 27 | 91.869 | 0 | 116.882 | 5 | 141.895 | 15
42.305 | 0 | 44.087 | 90 | 52.072 | 3 | 57.088 | 0 | 62.104 | 0 | 72.07 | 0 | 74.05 | 25 | 92.067 | 0 | 117.08 | 0 | 142.093 | 16
42.503 | 0 | 44.285 | 96 | 52.27 | 0 | 57.286 | 0 | 62.302 | 0 | 72.268 | 10 | 74.248 | 31 | 92.265 | 0 | 117.278 | 3 | 142.291 | 7
42.701 | 0 | 44.483 | 80 | 52.468 | 7 | 57.484 | 0 | 62.5 | 0 | 72.466 | 0 | 74.446 | 40 | 92.463 | 12 | 117.476 | 0 | 142.489 | 10
42.899 | 0 | 44.681 | 123 | 52.666 | 5 | 57.682 | 0 | 62.698 | 0 | 72.664 | 5 | 74.644 | 28 | 92.661 | 0 | 117.674 | 0 | 142.687 | 0
43.097 | 0 | 44.879 | 104 | 52.864 | 0 | 57.88 | 0 | 62.896 | 0 | 72.862 | 0 | 74.842 | 40 | 92.859 | 6 | 117.872 | 9 | 142.885 | 3
43.295 | 0 | 45.077 | 95 | 53.062 | 21 | 58.078 | 11 | 63.094 | 0 | 73.06 | 0 | 75.04 | 15 | 93.057 | 0 | 118.07 | 0 | 143.083 | 0
43.493 | 0 | 45.275 | 81 | 53.26 | 6 | 58.276 | 0 | 63.292 | 0 | 73.258 | 0 | 75.238 | 24 | 93.255 | 0 | 118.268 | 0 | 143.281 | 0
43.691 | 0 | 45.473 | 115 | 53.458 | 0 | 58.474 | 10 | 63.49 | 0 | 73.456 | 0 | 75.436 | 19 | 93.453 | 5 | 118.466 | 4 | 143.479 | 11
43.889 | 0 | 45.671 | 129 | 53.656 | 8 | 58.672 | 0 | 63.688 | 0 | 73.654 | 0 | 75.634 | 22 | 93.651 | 0 | 118.664 | 0 | 143.677 | 29
44.087 | 0 | 45.869 | 91 | 53.854 | 0 | 58.87 | 0 | 63.886 | 0 | 73.852 | 0 | 75.832 | 21 | 93.849 | 0 | 118.862 | 0 | 143.875 | 4
44.285 | 0 | 46.067 | 67 | 54.052 | 9 | 59.068 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 8 | 94.047 | 3 | 119.06 | 0 | 144.073 | 4
44.483 | 0 | 46.265 | 42 | 54.25 | 8 | 59.266 | 0 | 64.282 | 0 | 74.248 | 0 | 76.228 | 0 | 94.245 | 0 | 119.258 | 0 | 144.271 | 5
44.681 | 0 | 46.463 | 100 | 54.448 | 0 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 5 | 94.443 | 0 | 119.456 | 10 | 144.469 | 2
44.879 | 0 | 46.661 | 57 | 54.646 | 0 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 7 | 119.654 | 7 | 144.667 | 14
45.077 | 0 | 46.859 | 32 | 54.844 | 0 | 59.86 | 0 | 64.876 | 0 | 74.842 | 0 | 76.822 | 4 | 94.839 | 5 | 119.852 | 0 | 144.865 | 0
45.275 | 0 | 47.056 | 45 | 55.042 | 4 | 60.058 | 0 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 5 | 120.05 | 0 | 145.063 | 0
45.473 | 0 | 47.254 | 43 | 55.24 | 0 | 60.256 | 0 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 0 | 120.248 | 0 | 145.261 | 0
45.671 | 0 | 47.452 | 44 | 55.438 | 3 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 8 | 145.459 | 0
45.869 | 0 | 47.65 | 32 | 55.636 | 0 | 60.652 | 0 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 4 | 145.657 | 0
46.067 | 0 | 47.848 | 5 | 55.834 | 0 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 11 | 120.842 | 7 | 145.855 | 5
46.265 | 0 | 48.046 | 44 | 56.032 | 6 | 61.048 | 0 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 0 | 121.04 | 0 | 146.053 | 0
46.463 | 0 | 48.244 | 49 | 56.23 | 0 | 61.246 | 0 | 66.262 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 0 | 146.251 | 0
46.661 | 0 | 48.442 | 35 | 56.428 | 0 | 61.444 | 0 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 0 | 146.449 | 0
46.859 | 0 | 48.64 | 37 | 56.626 | 0 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 0 | 146.647 | 0
47.056 | 0 | 48.838 | 20 | 56.824 | 0 | 61.84 | 0 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 13 | 146.845 | 0
47.254 | 0 | 49.036 | 11 | 57.022 | 0 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 9 | 122.03 | 0 | 147.043 | 0
47.452 | 0 | 49.234 | 6 | 57.22 | 0 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 8 | 122.228 | 3 | 147.241 | 0
47.65 | 0 | 49.432 | 15 | 57.418 | 0 | 62.434 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 17 | 57.616 | 0 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 0 | 147.637 | 0
48.046 | 0 | 49.828 | 5 | 57.814 | 0 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 8 | 122.822 | 0 | 147.835 | 0
48.244 | 0 | 50.026 | 19 | 58.012 | 0 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 0 | 148.033 | 0
48.442 | 0 | 50.224 | 35 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 13 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 0 | 148.429 | 3
48.838 | 0 | 50.62 | 15 | 58.606 | 0 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 6 | 123.614 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 33 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 15 | 59.002 | 4 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 7 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 0 | 59.2 | 0 | 64.216 | 0 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 0 | 124.208 | 11 | 149.221 | 0
|  | 51.412 | 0 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 0 | 149.419 | 0

## Analyte Terfenadine
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 4 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 400 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 3306 | 13.002 | 0 | 20.987 | 31 | 26.003 | 0 | 31.019 | 11 | 40.985 | 0 | 42.965 | 4 | 60.982 | 6 | 85.995 | 43 | 111.008 | 10
11.418 | 2062 | 13.2 | 0 | 21.185 | 154 | 26.201 | 55 | 31.217 | 58 | 41.183 | 89 | 43.163 | 48 | 61.18 | 40 | 86.193 | 79 | 111.206 | 95
11.616 | 16231 | 13.398 | 3 | 21.383 | 134 | 26.399 | 383 | 31.415 | 114 | 41.381 | 166 | 43.361 | 99 | 61.378 | 190 | 86.391 | 706 | 111.404 | 79
11.814 | 285125 | 13.596 | 5 | 21.581 | 116 | 26.597 | 375 | 31.613 | 349 | 41.579 | 218 | 43.559 | 215 | 61.576 | 375 | 86.589 | 9698 | 111.602 | 178
12.012 | 841145 | 13.794 | 27 | 21.779 | 139 | 26.795 | 335 | 31.811 | 291 | 41.777 | 123 | 43.757 | 237 | 61.774 | 359 | 86.787 | 15345 | 111.8 | 205
12.21 | 720768 | 13.992 | 10 | 21.977 | 123 | 26.993 | 274 | 32.009 | 214 | 41.975 | 163 | 43.955 | 259 | 61.972 | 320 | 86.985 | 12206 | 111.998 | 143
12.408 | 151015 | 14.19 | 0 | 22.175 | 130 | 27.191 | 317 | 32.207 | 153 | 42.173 | 314 | 44.153 | 220 | 62.17 | 395 | 87.183 | 10662 | 112.196 | 209
12.606 | 7254 | 14.388 | 8 | 22.373 | 96 | 27.389 | 464 | 32.405 | 155 | 42.371 | 217 | 44.351 | 213 | 62.368 | 439 | 87.381 | 9291 | 112.394 | 86
12.804 | 192 | 14.586 | 18 | 22.571 | 131 | 27.587 | 381 | 32.603 | 91 | 42.569 | 179 | 44.549 | 388 | 62.566 | 325 | 87.579 | 6869 | 112.592 | 100
13.002 | 329 | 14.784 | 0 | 22.769 | 212 | 27.785 | 608 | 32.801 | 43 | 42.767 | 149 | 44.747 | 384 | 62.764 | 427 | 87.777 | 4714 | 112.79 | 153
13.2 | 201 | 14.982 | 0 | 22.967 | 101 | 27.983 | 795 | 32.999 | 98 | 42.965 | 89 | 44.945 | 579 | 62.962 | 377 | 87.975 | 2653 | 112.988 | 104
13.398 | 5 | 15.18 | 0 | 23.165 | 159 | 28.181 | 974 | 33.197 | 80 | 43.163 | 206 | 45.143 | 493 | 63.16 | 323 | 88.173 | 1346 | 113.186 | 85
13.596 | 0 | 15.378 | 0 | 23.363 | 149 | 28.379 | 1261 | 33.395 | 71 | 43.361 | 222 | 45.341 | 501 | 63.358 | 283 | 88.371 | 895 | 113.384 | 137
13.794 | 0 | 15.576 | 0 | 23.561 | 110 | 28.577 | 2746 | 33.593 | 79 | 43.559 | 256 | 45.539 | 524 | 63.556 | 287 | 88.569 | 453 | 113.582 | 108
13.992 | 0 | 15.773 | 0 | 23.759 | 115 | 28.775 | 15000 | 33.791 | 87 | 43.757 | 235 | 45.737 | 512 | 63.754 | 401 | 88.767 | 162 | 113.78 | 99
14.19 | 0 | 15.971 | 0 | 23.957 | 157 | 28.973 | 81315 | 33.989 | 87 | 43.955 | 270 | 45.935 | 669 | 63.952 | 1068 | 88.965 | 156 | 113.978 | 216
14.388 | 0 | 16.169 | 0 | 24.155 | 180 | 29.171 | 267828 | 34.187 | 105 | 44.153 | 220 | 46.133 | 891 | 64.15 | 3192 | 89.163 | 84 | 114.176 | 116
14.586 | 0 | 16.367 | 0 | 24.353 | 167 | 29.369 | 539480 | 34.385 | 115 | 44.351 | 162 | 46.331 | 3059 | 64.348 | 10639 | 89.361 | 109 | 114.374 | 151
14.784 | 0 | 16.565 | 16 | 24.551 | 339 | 29.567 | 657418 | 34.583 | 68 | 44.549 | 425 | 46.529 | 11986 | 64.546 | 27423 | 89.559 | 52 | 114.572 | 170
14.982 | 0 | 16.763 | 7 | 24.749 | 445 | 29.765 | 539075 | 34.781 | 45 | 44.747 | 407 | 46.727 | 41429 | 64.744 | 60408 | 89.757 | 134 | 114.77 | 154
15.18 | 0 | 16.961 | 16 | 24.947 | 401 | 29.963 | 302715 | 34.979 | 107 | 44.945 | 581 | 46.924 | 105344 | 64.942 | 110497 | 89.955 | 100 | 114.968 | 144
15.378 | 0 | 17.159 | 25 | 25.145 | 612 | 30.161 | 138559 | 35.177 | 57 | 45.143 | 653 | 47.122 | 204821 | 65.14 | 163275 | 90.153 | 101 | 115.166 | 87
15.576 | 0 | 17.357 | 11 | 25.343 | 485 | 30.359 | 79400 | 35.375 | 59 | 45.341 | 574 | 47.32 | 305703 | 65.338 | 202192 | 90.351 | 117 | 115.364 | 85
15.773 | 0 | 17.555 | 22 | 25.541 | 361 | 30.557 | 70055 | 35.573 | 124 | 45.539 | 583 | 47.518 | 355427 | 65.536 | 210905 | 90.549 | 143 | 115.562 | 110
15.971 | 0 | 17.753 | 30 | 25.739 | 379 | 30.755 | 55826 | 35.771 | 77 | 45.737 | 596 | 47.716 | 321307 | 65.734 | 184346 | 90.747 | 125 | 115.76 | 115
16.169 | 0 | 17.951 | 34 | 25.937 | 217 | 30.953 | 30591 | 35.969 | 56 | 45.935 | 701 | 47.914 | 228549 | 65.932 | 136907 | 90.945 | 104 | 115.958 | 102
16.367 | 0 | 18.149 | 17 | 26.135 | 425 | 31.151 | 10833 | 36.167 | 171 | 46.133 | 828 | 48.112 | 129103 | 66.13 | 84902 | 91.143 | 98 | 116.156 | 104
16.565 | 0 | 18.347 | 39 | 26.333 | 445 | 31.349 | 2479 | 36.365 | 274 | 46.331 | 2070 | 48.31 | 60531 | 66.328 | 45368 | 91.341 | 81 | 116.354 | 160
16.763 | 0 | 18.545 | 24 | 26.531 | 385 | 31.547 | 498 | 36.563 | 235 | 46.529 | 8844 | 48.508 | 23892 | 66.526 | 20847 | 91.539 | 108 | 116.552 | 44
16.961 | 0 | 18.743 | 16 | 26.729 | 330 | 31.745 | 141 | 36.761 | 402 | 46.727 | 32236 | 48.706 | 12674 | 66.724 | 8084 | 91.737 | 147 | 116.75 | 92
17.159 | 0 | 18.941 | 22 | 26.927 | 277 | 31.943 | 159 | 36.959 | 253 | 46.924 | 90041 | 48.904 | 15711 | 66.922 | 3111 | 91.935 | 164 | 116.948 | 88
17.357 | 0 | 19.139 | 67 | 27.125 | 292 | 32.141 | 92 | 37.157 | 355 | 47.122 | 190809 | 49.102 | 24383 | 67.12 | 1750 | 92.133 | 168 | 117.146 | 176
17.555 | 0 | 19.337 | 26 | 27.323 | 367 | 32.339 | 107 | 37.355 | 290 | 47.32 | 303362 | 49.3 | 32967 | 67.318 | 2808 | 92.331 | 169 | 117.344 | 498
17.753 | 0 | 19.535 | 168 | 27.521 | 410 | 32.537 | 113 | 37.553 | 406 | 47.518 | 380584 | 49.498 | 34893 | 67.516 | 5339 | 92.529 | 141 | 117.542 | 1061
17.951 | 0 | 19.733 | 450 | 27.719 | 587 | 32.735 | 94 | 37.751 | 329 | 47.716 | 365919 | 49.696 | 30428 | 67.714 | 9335 | 92.727 | 107 | 117.74 | 2284
18.149 | 0 | 19.931 | 740 | 27.917 | 749 | 32.933 | 81 | 37.949 | 271 | 47.914 | 279376 | 49.894 | 20294 | 67.912 | 14655 | 92.925 | 86 | 117.938 | 4774
18.347 | 0 | 20.129 | 568 | 28.115 | 840 | 33.131 | 62 | 38.147 | 332 | 48.112 | 166553 | 50.092 | 10900 | 68.11 | 17904 | 93.123 | 152 | 118.136 | 9940
18.545 | 0 | 20.327 | 424 | 28.313 | 1042 | 33.329 | 55 | 38.345 | 357 | 48.31 | 82066 | 50.29 | 4940 | 68.308 | 20745 | 93.321 | 130 | 118.334 | 17469
18.743 | 0 | 20.525 | 266 | 28.511 | 1524 | 33.527 | 62 | 38.543 | 437 | 48.508 | 32607 | 50.488 | 1700 | 68.506 | 18984 | 93.519 | 83 | 118.532 | 28333
18.941 | 0 | 20.723 | 150 | 28.709 | 3120 | 33.725 | 96 | 38.741 | 344 | 48.706 | 14984 | 50.686 | 565 | 68.704 | 16029 | 93.717 | 90 | 118.73 | 40515
19.139 | 0 | 20.921 | 79 | 28.907 | 16578 | 33.923 | 58 | 38.939 | 203 | 48.904 | 15042 | 50.884 | 275 | 68.902 | 11398 | 93.915 | 118 | 118.928 | 54163
19.337 | 0 | 21.119 | 68 | 29.105 | 88134 | 34.121 | 69 | 39.137 | 213 | 49.102 | 23582 | 51.082 | 129 | 69.1 | 6445 | 94.113 | 133 | 119.126 | 65988
19.535 | 0 | 21.317 | 78 | 29.303 | 293775 | 34.319 | 63 | 39.335 | 218 | 49.3 | 31234 | 51.28 | 93 | 69.298 | 3458 | 94.311 | 92 | 119.324 | 73089
19.733 | 0 | 21.515 | 81 | 29.501 | 588844 | 34.517 | 78 | 39.533 | 246 | 49.498 | 38093 | 51.478 | 196 | 69.496 | 1601 | 94.509 | 109 | 119.522 | 71798
19.931 | 0 | 21.713 | 73 | 29.699 | 730366 | 34.715 | 62 | 39.731 | 216 | 49.696 | 34686 | 51.676 | 258 | 69.694 | 697 | 94.707 | 127 | 119.72 | 65798
20.129 | 0 | 21.911 | 77 | 29.897 | 605256 | 34.913 | 56 | 39.929 | 153 | 49.894 | 24851 | 51.874 | 232 | 69.892 | 323 | 94.905 | 118 | 119.918 | 55432
20.327 | 0 | 22.109 | 115 | 30.095 | 339292 | 35.111 | 68 | 40.127 | 225 | 50.092 | 14402 | 52.072 | 267 | 70.09 | 112 | 95.103 | 175 | 120.116 | 43414
20.525 | 0 | 22.307 | 118 | 30.293 | 155270 | 35.309 | 117 | 40.325 | 181 | 50.29 | 6952 | 52.27 | 202 | 70.288 | 105 | 95.301 | 139 | 120.314 | 29705
20.723 | 0 | 22.505 | 53 | 30.491 | 88762 | 35.507 | 60 | 40.523 | 186 | 50.488 | 2469 | 52.468 | 202 | 70.486 | 93 | 95.499 | 163 | 120.512 | 19549
20.921 | 0 | 22.703 | 123 | 30.689 | 79848 | 35.705 | 47 | 40.721 | 187 | 50.686 | 818 | 52.666 | 177 | 70.684 | 71 | 95.697 | 151 | 120.71 | 11338
21.119 | 0 | 22.901 | 116 | 30.887 | 62205 | 35.903 | 108 | 40.919 | 301 | 50.884 | 249 | 52.864 | 211 | 70.882 | 65 | 95.895 | 144 | 120.908 | 5948
21.317 | 0 | 23.099 | 133 | 31.085 | 33795 | 36.101 | 121 | 41.117 | 268 | 51.082 | 171 | 53.062 | 184 | 71.08 | 103 | 96.093 | 165 | 121.106 | 2935
21.515 | 0 | 23.297 | 100 | 31.283 | 12013 | 36.299 | 71 | 41.315 | 219 | 51.28 | 112 | 53.26 | 152 | 71.278 | 82 | 96.291 | 98 | 121.304 | 1359
21.713 | 0 | 23.495 | 167 | 31.481 | 2835 | 36.497 | 41 | 41.513 | 191 | 51.478 | 73 | 53.458 | 166 | 71.476 | 45 | 96.489 | 101 | 121.502 | 518
21.911 | 0 | 23.693 | 194 | 31.679 | 523 | 36.695 | 79 | 41.711 | 141 | 51.676 | 162 | 53.656 | 152 | 71.674 | 68 | 96.687 | 128 | 121.7 | 216
22.109 | 0 | 23.891 | 104 | 31.877 | 86 | 36.893 | 100 | 41.909 | 111 | 51.874 | 219 | 53.854 | 163 | 71.872 | 60 | 96.885 | 69 | 121.898 | 76
22.307 | 0 | 24.089 | 99 | 32.075 | 87 | 37.091 | 42 | 42.107 | 119 | 52.072 | 296 | 54.052 | 126 | 72.07 | 85 | 97.083 | 119 | 122.096 | 66
22.505 | 0 | 24.287 | 123 | 32.273 | 38 | 37.289 | 82 | 42.305 | 130 | 52.27 | 269 | 54.25 | 115 | 72.268 | 69 | 97.281 | 95 | 122.294 | 63
22.703 | 0 | 24.485 | 167 | 32.471 | 98 | 37.487 | 58 | 42.503 | 111 | 52.468 | 141 | 54.448 | 60 | 72.466 | 80 | 97.479 | 63 | 122.492 | 34
22.901 | 0 | 24.683 | 98 | 32.669 | 73 | 37.685 | 97 | 42.701 | 78 | 52.666 | 186 | 54.646 | 57 | 72.664 | 70 | 97.677 | 132 | 122.69 | 53
23.099 | 0 | 24.881 | 224 | 32.867 | 41 | 37.883 | 69 | 42.899 | 155 | 52.864 | 124 | 54.844 | 61 | 72.862 | 48 | 97.875 | 107 | 122.888 | 139
23.297 | 0 | 25.079 | 663 | 33.065 | 35 | 38.081 | 150 | 43.097 | 139 | 53.062 | 125 | 55.042 | 89 | 73.06 | 60 | 98.073 | 102 | 123.086 | 221
23.495 | 0 | 25.277 | 611 | 33.263 | 30 | 38.279 | 111 | 43.295 | 133 | 53.26 | 130 | 55.24 | 149 | 73.258 | 46 | 98.271 | 132 | 123.284 | 324
23.693 | 0 | 25.475 | 399 | 33.461 | 20 | 38.477 | 301 | 43.493 | 174 | 53.458 | 178 | 55.438 | 83 | 73.456 | 101 | 98.469 | 85 | 123.482 | 732
23.891 | 0 | 25.673 | 402 | 33.659 | 36 | 38.675 | 330 | 43.691 | 234 | 53.656 | 135 | 55.636 | 84 | 73.654 | 66 | 98.667 | 52 | 123.68 | 1021
24.089 | 0 | 25.871 | 358 | 33.857 | 5 | 38.873 | 311 | 43.889 | 228 | 53.854 | 92 | 55.834 | 91 | 73.852 | 82 | 98.865 | 51 | 123.878 | 1494
24.287 | 0 | 26.069 | 390 | 34.055 | 3 | 39.071 | 203 | 44.087 | 172 | 54.052 | 137 | 56.032 | 180 | 74.05 | 43 | 99.063 | 54 | 124.076 | 2497
24.485 | 0 | 26.267 | 425 | 34.253 | 0 | 39.269 | 296 | 44.285 | 205 | 54.25 | 158 | 56.23 | 133 | 74.248 | 46 | 99.261 | 91 | 124.274 | 3933
24.683 | 0 | 26.465 | 469 | 34.451 | 17 | 39.467 | 180 | 44.483 | 203 | 54.448 | 184 | 56.428 | 102 | 74.446 | 18 | 99.459 | 133 | 124.472 | 5030
24.881 | 0 | 26.663 | 404 | 34.649 | 4 | 39.665 | 215 | 44.681 | 229 | 54.646 | 77 | 56.626 | 78 | 74.644 | 42 | 99.657 | 398 | 124.67 | 6085
25.079 | 0 | 26.861 | 334 | 34.847 | 5 | 39.863 | 217 | 44.879 | 258 | 54.844 | 55 | 56.824 | 76 | 74.842 | 26 | 99.855 | 1085 | 124.868 | 6929
25.277 | 0 | 27.059 | 227 | 35.045 | 62 | 40.061 | 228 | 45.077 | 376 | 55.042 | 66 | 57.022 | 101 | 75.04 | 27 | 100.053 | 2515 | 125.066 | 7183
25.475 | 0 | 27.257 | 234 | 35.243 | 0 | 40.259 | 138 | 45.275 | 472 | 55.24 | 103 | 57.22 | 104 | 75.238 | 62 | 100.251 | 6462 | 125.264 | 6835
25.673 | 0 | 27.455 | 288 | 35.441 | 30 | 40.457 | 162 | 45.473 | 411 | 55.438 | 89 | 57.418 | 37 | 75.436 | 47 | 100.449 | 13240 | 125.462 | 5762
25.871 | 0 | 27.653 | 397 | 35.639 | 23 | 40.655 | 99 | 45.671 | 506 | 55.636 | 94 | 57.616 | 75 | 75.634 | 59 | 100.647 | 24852 | 125.66 | 4895
26.069 | 0 | 27.851 | 457 | 35.837 | 8 | 40.853 | 185 | 45.869 | 461 | 55.834 | 136 | 57.814 | 96 | 75.832 | 34 | 100.845 | 40290 | 125.858 | 3491
26.267 | 0 | 28.049 | 551 | 36.035 | 27 | 41.051 | 181 | 46.067 | 541 | 56.032 | 204 | 58.012 | 74 | 76.03 | 103 | 101.043 | 62883 | 126.056 | 2519
26.465 | 0 | 28.247 | 684 | 36.233 | 29 | 41.249 | 196 | 46.265 | 539 | 56.23 | 87 | 58.21 | 81 | 76.228 | 55 | 101.241 | 85132 | 126.254 | 1626
26.663 | 0 | 28.445 | 876 | 36.431 | 43 | 41.447 | 136 | 46.463 | 1065 | 56.428 | 81 | 58.408 | 82 | 76.426 | 37 | 101.439 | 100559 | 126.452 | 887
26.861 | 0 | 28.643 | 909 | 36.629 | 31 | 41.645 | 49 | 46.661 | 3594 | 56.626 | 19 | 58.606 | 84 | 76.624 | 85 | 101.637 | 109039 | 126.65 | 440
27.059 | 0 | 28.841 | 1849 | 36.827 | 29 | 41.843 | 9 | 46.859 | 15518 | 56.824 | 17 | 58.804 | 43 | 76.822 | 70 | 101.835 | 105574 | 126.848 | 307
27.257 | 0 | 29.039 | 7665 | 37.025 | 0 | 42.041 | 3 | 47.056 | 54365 | 57.022 | 13 | 59.002 | 76 | 77.02 | 81 | 102.033 | 93150 | 127.046 | 124
27.455 | 0 | 29.237 | 49739 | 37.223 | 0 | 42.239 | 6 | 47.254 | 138148 | 57.22 | 3 | 59.2 | 109 | 77.218 | 69 | 102.231 | 73352 | 127.244 | 75
27.653 | 0 | 29.435 | 204368 | 37.421 | 4 | 42.437 | 12 | 47.452 | 272269 | 57.418 | 9 | 59.398 | 65 | 77.416 | 69 | 102.429 | 51534 | 127.442 | 44
27.851 | 0 | 29.633 | 505783 | 37.619 | 3 | 42.635 | 0 | 47.65 | 404110 | 57.616 | 0 | 59.596 | 59 | 77.614 | 68 | 102.627 | 32416 | 127.64 | 26
28.049 | 0 | 29.831 | 772579 | 37.817 | 0 | 42.833 | 10 | 47.848 | 470154 | 57.814 | 6 | 59.794 | 29 | 77.812 | 53 | 102.825 | 18513 | 127.838 | 62
28.247 | 0 | 30.029 | 757931 | 38.015 | 2 | 43.031 | 29 | 48.046 | 422339 | 58.012 | 4 | 59.992 | 14 | 78.01 | 30 | 103.023 | 9891 | 128.036 | 34
28.445 | 0 | 30.227 | 499161 | 38.213 | 0 | 43.229 | 16 | 48.244 | 299189 | 58.21 | 0 | 60.19 | 23 | 78.208 | 85 | 103.221 | 4405 | 128.234 | 42
28.643 | 0 | 30.425 | 243146 | 38.411 | 0 | 43.427 | 3 | 48.442 | 172471 | 58.408 | 20 | 60.388 | 48 | 78.405 | 41 | 103.419 | 2133 | 128.432 | 6
28.841 | 0 | 30.623 | 116946 | 38.609 | 3 | 43.625 | 3 | 48.64 | 79178 | 58.606 | 0 | 60.586 | 35 | 78.603 | 21 | 103.617 | 629 | 128.63 | 0
29.039 | 0 | 30.821 | 91192 | 38.807 | 2 | 43.823 | 8 | 48.838 | 32048 | 58.804 | 3 | 60.784 | 37 | 78.801 | 15 | 103.815 | 253 | 128.828 | 16
29.237 | 0 | 31.019 | 80094 | 39.005 | 14 | 44.021 | 7 | 49.036 | 16723 | 59.002 | 0 | 60.982 | 45 | 78.999 | 20 | 104.013 | 64 | 129.026 | 11
29.435 | 0 | 31.217 | 50971 | 39.203 | 9 | 44.219 | 4 | 49.234 | 20157 | 59.2 | 6 | 61.18 | 49 | 79.197 | 7 | 104.211 | 117 | 129.224 | 0
29.633 | 0 | 31.415 | 22081 | 39.401 | 0 | 44.417 | 4 | 49.432 | 31422 | 59.398 | 4 | 61.378 | 32 | 79.395 | 9 | 104.409 | 123 | 129.422 | 4
29.831 | 0 | 31.613 | 6412 | 39.599 | 0 | 44.615 | 6 | 49.63 | 41930 | 59.596 | 21 | 61.576 | 41 | 79.593 | 8 | 104.607 | 224 | 129.62 | 6
30.029 | 0 | 31.811 | 1334 | 39.797 | 0 | 44.813 | 11 | 49.828 | 46333 | 59.794 | 0 | 61.774 | 45 | 79.791 | 0 | 104.805 | 479 | 129.818 | 2
30.227 | 0 | 32.009 | 211 | 39.995 | 0 | 45.011 | 0 | 50.026 | 40357 | 59.992 | 5 | 61.972 | 52 | 79.989 | 4 | 105.003 | 760 | 130.016 | 6
30.425 | 0 | 32.207 | 124 | 40.193 | 0 | 45.209 | 0 | 50.224 | 27164 | 60.19 | 20 | 62.17 | 76 | 80.187 | 12 | 105.201 | 1197 | 130.214 | 0
30.623 | 0 | 32.405 | 56 | 40.391 | 2 | 45.407 | 4 | 50.422 | 15686 | 60.388 | 11 | 62.368 | 82 | 80.385 | 4 | 105.399 | 1739 | 130.412 | 9
30.821 | 0 | 32.603 | 84 | 40.589 | 17 | 45.605 | 17 | 50.62 | 6706 | 60.586 | 38 | 62.566 | 56 | 80.583 | 9 | 105.597 | 1734 | 130.61 | 4
31.019 | 0 | 32.801 | 63 | 40.787 | 29 | 45.803 | 7 | 50.818 | 2078 | 60.784 | 47 | 62.764 | 53 | 80.781 | 0 | 105.795 | 1434 | 130.808 | 16
31.217 | 0 | 32.999 | 40 | 40.985 | 15 | 46.001 | 6 | 51.016 | 726 | 60.982 | 95 | 62.962 | 20 | 80.979 | 9 | 105.993 | 595 | 131.006 | 5
31.415 | 0 | 33.197 | 20 | 41.183 | 11 | 46.199 | 8 | 51.214 | 146 | 61.18 | 52 | 63.16 | 35 | 81.177 | 21 | 106.191 | 273 | 131.204 | 16
31.613 | 0 | 33.395 | 17 | 41.381 | 31 | 46.397 | 10 | 51.412 | 64 | 61.378 | 128 | 63.358 | 14 | 81.375 | 8 | 106.389 | 47 | 131.402 | 0
31.811 | 0 | 33.593 | 33 | 41.579 | 14 | 46.595 | 4 | 51.61 | 24 | 61.576 | 84 | 63.556 | 35 | 81.573 | 5 | 106.587 | 4 | 131.6 | 4
32.009 | 0 | 33.791 | 44 | 41.777 | 15 | 46.793 | 12 | 51.808 | 14 | 61.774 | 36 | 63.754 | 0 | 81.771 | 27 | 106.785 | 0 | 131.798 | 7
32.207 | 0 | 33.989 | 2 | 41.975 | 0 | 46.99 | 0 | 52.006 | 13 | 61.972 | 34 | 63.952 | 0 | 81.969 | 0 | 106.983 | 0 | 131.996 | 0
32.405 | 0 | 34.187 | 15 | 42.173 | 19 | 47.188 | 18 | 52.204 | 29 | 62.17 | 90 | 64.15 | 0 | 82.167 | 0 | 107.181 | 0 | 132.194 | 5
32.603 | 0 | 34.385 | 11 | 42.371 | 13 | 47.386 | 21 | 52.402 | 0 | 62.368 | 54 | 64.348 | 0 | 82.365 | 7 | 107.379 | 0 | 132.392 | 8
32.801 | 0 | 34.583 | 5 | 42.569 | 0 | 47.584 | 7 | 52.6 | 30 | 62.566 | 31 | 64.546 | 0 | 82.563 | 27 | 107.577 | 3 | 132.59 | 0
32.999 | 0 | 34.781 | 11 | 42.767 | 17 | 47.782 | 10 | 52.798 | 12 | 62.764 | 27 | 64.744 | 8 | 82.761 | 7 | 107.775 | 10 | 132.788 | 0
33.197 | 0 | 34.979 | 0 | 42.965 | 16 | 47.98 | 19 | 52.996 | 0 | 62.962 | 48 | 64.942 | 0 | 82.959 | 20 | 107.973 | 5 | 132.986 | 7
33.395 | 0 | 35.177 | 8 | 43.163 | 17 | 48.178 | 23 | 53.194 | 0 | 63.16 | 31 | 65.14 | 0 | 83.157 | 12 | 108.171 | 0 | 133.184 | 0
33.593 | 0 | 35.375 | 6 | 43.361 | 5 | 48.376 | 13 | 53.392 | 0 | 63.358 | 32 | 65.338 | 13 | 83.355 | 11 | 108.369 | 0 | 133.382 | 0
33.791 | 0 | 35.573 | 18 | 43.559 | 17 | 48.574 | 28 | 53.59 | 0 | 63.556 | 15 | 65.536 | 0 | 83.553 | 0 | 108.567 | 0 | 133.58 | 0
33.989 | 0 | 35.771 | 3 | 43.757 | 7 | 48.772 | 9 | 53.788 | 0 | 63.754 | 17 | 65.734 | 0 | 83.751 | 0 | 108.765 | 17 | 133.778 | 6
34.187 | 0 | 35.969 | 9 | 43.955 | 3 | 48.97 | 24 | 53.986 | 0 | 63.952 | 10 | 65.932 | 6 | 83.949 | 19 | 108.963 | 7 | 133.976 | 0
34.385 | 0 | 36.167 | 41 | 44.153 | 10 | 49.168 | 21 | 54.184 | 9 | 64.15 | 3 | 66.13 | 0 | 84.147 | 0 | 109.161 | 8 | 134.174 | 4
34.583 | 0 | 36.365 | 39 | 44.351 | 6 | 49.366 | 19 | 54.382 | 6 | 64.348 | 5 | 66.328 | 0 | 84.345 | 0 | 109.359 | 0 | 134.372 | 6
34.781 | 0 | 36.563 | 21 | 44.549 | 24 | 49.564 | 5 | 54.58 | 0 | 64.546 | 0 | 66.526 | 0 | 84.543 | 13 | 109.556 | 13 | 134.57 | 0
34.979 | 0 | 36.761 | 23 | 44.747 | 5 | 49.762 | 8 | 54.778 | 0 | 64.744 | 8 | 66.724 | 0 | 84.741 | 15 | 109.754 | 0 | 134.768 | 4
35.177 | 0 | 36.959 | 21 | 44.945 | 0 | 49.96 | 6 | 54.976 | 0 | 64.942 | 0 | 66.922 | 5 | 84.939 | 0 | 109.952 | 24 | 134.966 | 0
35.375 | 0 | 37.157 | 18 | 45.143 | 11 | 50.158 | 11 | 55.174 | 0 | 65.14 | 0 | 67.12 | 0 | 85.137 | 18 | 110.15 | 14 | 135.164 | 0
35.573 | 0 | 37.355 | 10 | 45.341 | 0 | 50.356 | 0 | 55.372 | 4 | 65.338 | 0 | 67.318 | 20 | 85.335 | 0 | 110.348 | 29 | 135.362 | 0
35.771 | 0 | 37.553 | 18 | 45.539 | 7 | 50.554 | 9 | 55.57 | 16 | 65.536 | 0 | 67.516 | 2 | 85.533 | 0 | 110.546 | 22 | 135.56 | 3
35.969 | 0 | 37.751 | 0 | 45.737 | 4 | 50.752 | 0 | 55.768 | 0 | 65.734 | 0 | 67.714 | 0 | 85.731 | 0 | 110.744 | 18 | 135.758 | 0
36.167 | 0 | 37.949 | 0 | 45.935 | 20 | 50.95 | 0 | 55.966 | 0 | 65.932 | 0 | 67.912 | 0 | 85.929 | 4 | 110.942 | 19 | 135.956 | 4
36.365 | 0 | 38.147 | 12 | 46.133 | 3 | 51.148 | 7 | 56.164 | 6 | 66.13 | 4 | 68.11 | 0 | 86.127 | 0 | 111.14 | 15 | 136.154 | 0
36.563 | 0 | 38.345 | 15 | 46.331 | 7 | 51.346 | 15 | 56.362 | 0 | 66.328 | 0 | 68.308 | 6 | 86.325 | 0 | 111.338 | 79 | 136.352 | 0
36.761 | 0 | 38.543 | 6 | 46.529 | 6 | 51.544 | 0 | 56.56 | 0 | 66.526 | 0 | 68.506 | 0 | 86.523 | 0 | 111.536 | 15 | 136.55 | 0
36.959 | 0 | 38.741 | 0 | 46.727 | 23 | 51.742 | 18 | 56.758 | 0 | 66.724 | 0 | 68.704 | 0 | 86.721 | 0 | 111.734 | 11 | 136.748 | 0
37.157 | 0 | 38.939 | 20 | 46.924 | 11 | 51.94 | 9 | 56.956 | 6 | 66.922 | 0 | 68.902 | 0 | 86.919 | 9 | 111.932 | 35 | 136.946 | 0
37.355 | 0 | 39.137 | 4 | 47.122 | 48 | 52.138 | 0 | 57.154 | 8 | 67.12 | 0 | 69.1 | 8 | 87.117 | 0 | 112.13 | 6 | 137.144 | 5
37.553 | 0 | 39.335 | 4 | 47.32 | 19 | 52.336 | 9 | 57.352 | 0 | 67.318 | 0 | 69.298 | 8 | 87.315 | 13 | 112.328 | 44 | 137.342 | 0
37.751 | 0 | 39.533 | 0 | 47.518 | 27 | 52.534 | 0 | 57.55 | 22 | 67.516 | 0 | 69.496 | 0 | 87.513 | 0 | 112.526 | 13 | 137.54 | 0
37.949 | 0 | 39.731 | 8 | 47.716 | 19 | 52.732 | 10 | 57.748 | 0 | 67.714 | 0 | 69.694 | 8 | 87.711 | 5 | 112.724 | 0 | 137.738 | 0
38.147 | 0 | 39.929 | 0 | 47.914 | 22 | 52.93 | 0 | 57.946 | 4 | 67.912 | 4 | 69.892 | 0 | 87.909 | 9 | 112.922 | 23 | 137.936 | 10
38.345 | 0 | 40.127 | 3 | 48.112 | 21 | 53.128 | 0 | 58.144 | 0 | 68.11 | 0 | 70.09 | 0 | 88.107 | 0 | 113.12 | 28 | 138.134 | 4
38.543 | 0 | 40.325 | 0 | 48.31 | 0 | 53.326 | 0 | 58.342 | 11 | 68.308 | 0 | 70.288 | 0 | 88.305 | 8 | 113.318 | 34 | 138.332 | 0
38.741 | 0 | 40.523 | 11 | 48.508 | 13 | 53.524 | 0 | 58.54 | 8 | 68.506 | 2 | 70.486 | 0 | 88.503 | 0 | 113.516 | 16 | 138.53 | 0
38.939 | 0 | 40.721 | 31 | 48.706 | 12 | 53.722 | 12 | 58.738 | 0 | 68.704 | 4 | 70.684 | 0 | 88.701 | 0 | 113.714 | 14 | 138.728 | 0
39.137 | 0 | 40.919 | 0 | 48.904 | 8 | 53.92 | 0 | 58.936 | 0 | 68.902 | 0 | 70.882 | 15 | 88.899 | 0 | 113.912 | 55 | 138.926 | 0
39.335 | 0 | 41.117 | 18 | 49.102 | 12 | 54.118 | 0 | 59.134 | 0 | 69.1 | 0 | 71.08 | 0 | 89.097 | 0 | 114.11 | 12 | 139.124 | 0
39.533 | 0 | 41.315 | 26 | 49.3 | 5 | 54.316 | 9 | 59.332 | 0 | 69.298 | 2 | 71.278 | 0 | 89.295 | 0 | 114.308 | 10 | 139.322 | 0
39.731 | 0 | 41.513 | 36 | 49.498 | 16 | 54.514 | 7 | 59.53 | 9 | 69.496 | 9 | 71.476 | 3 | 89.493 | 0 | 114.506 | 0 | 139.52 | 0
39.929 | 0 | 41.711 | 6 | 49.696 | 4 | 54.712 | 24 | 59.728 | 11 | 69.694 | 3 | 71.674 | 7 | 89.691 | 5 | 114.704 | 0 | 139.718 | 0
40.127 | 0 | 41.909 | 7 | 49.894 | 0 | 54.91 | 0 | 59.926 | 11 | 69.892 | 2 | 71.872 | 4 | 89.889 | 12 | 114.902 | 7 | 139.916 | 0
40.325 | 0 | 42.107 | 0 | 50.092 | 8 | 55.108 | 12 | 60.124 | 0 | 70.09 | 5 | 72.07 | 0 | 90.087 | 0 | 115.1 | 0 | 140.114 | 0
40.523 | 0 | 42.305 | 26 | 50.29 | 5 | 55.306 | 0 | 60.322 | 29 | 70.288 | 12 | 72.268 | 0 | 90.285 | 3 | 115.298 | 0 | 140.312 | 0
40.721 | 0 | 42.503 | 40 | 50.488 | 20 | 55.504 | 0 | 60.52 | 19 | 70.486 | 8 | 72.466 | 4 | 90.483 | 0 | 115.496 | 0 | 140.51 | 0
40.919 | 0 | 42.701 | 0 | 50.686 | 0 | 55.702 | 0 | 60.718 | 12 | 70.684 | 11 | 72.664 | 0 | 90.681 | 12 | 115.694 | 4 | 140.708 | 0
41.117 | 0 | 42.899 | 0 | 50.884 | 10 | 55.9 | 3 | 60.916 | 15 | 70.882 | 14 | 72.862 | 0 | 90.879 | 0 | 115.892 | 0 | 140.906 | 0
41.315 | 0 | 43.097 | 14 | 51.082 | 13 | 56.098 | 3 | 61.114 | 45 | 71.08 | 2 | 73.06 | 0 | 91.077 | 0 | 116.09 | 9 | 141.103 | 0
41.513 | 0 | 43.295 | 16 | 51.28 | 6 | 56.296 | 0 | 61.312 | 77 | 71.278 | 0 | 73.258 | 4 | 91.275 | 5 | 116.288 | 3 | 141.301 | 0
41.711 | 0 | 43.493 | 5 | 51.478 | 9 | 56.494 | 4 | 61.51 | 78 | 71.476 | 0 | 73.456 | 3 | 91.473 | 2 | 116.486 | 0 | 141.499 | 13
41.909 | 0 | 43.691 | 5 | 51.676 | 0 | 56.692 | 0 | 61.708 | 39 | 71.674 | 0 | 73.654 | 0 | 91.671 | 0 | 116.684 | 6 | 141.697 | 0
42.107 | 0 | 43.889 | 25 | 51.874 | 0 | 56.89 | 0 | 61.906 | 53 | 71.872 | 0 | 73.852 | 6 | 91.869 | 0 | 116.882 | 0 | 141.895 | 7
42.305 | 0 | 44.087 | 27 | 52.072 | 12 | 57.088 | 0 | 62.104 | 75 | 72.07 | 6 | 74.05 | 0 | 92.067 | 0 | 117.08 | 0 | 142.093 | 0
42.503 | 0 | 44.285 | 8 | 52.27 | 5 | 57.286 | 0 | 62.302 | 63 | 72.268 | 0 | 74.248 | 0 | 92.265 | 2 | 117.278 | 0 | 142.291 | 0
42.701 | 0 | 44.483 | 5 | 52.468 | 0 | 57.484 | 0 | 62.5 | 84 | 72.466 | 0 | 74.446 | 9 | 92.463 | 0 | 117.476 | 0 | 142.489 | 0
42.899 | 0 | 44.681 | 7 | 52.666 | 3 | 57.682 | 0 | 62.698 | 75 | 72.664 | 0 | 74.644 | 0 | 92.661 | 0 | 117.674 | 0 | 142.687 | 0
43.097 | 0 | 44.879 | 4 | 52.864 | 3 | 57.88 | 0 | 62.896 | 35 | 72.862 | 0 | 74.842 | 4 | 92.859 | 0 | 117.872 | 6 | 142.885 | 0
43.295 | 0 | 45.077 | 11 | 53.062 | 5 | 58.078 | 0 | 63.094 | 38 | 73.06 | 0 | 75.04 | 0 | 93.057 | 0 | 118.07 | 6 | 143.083 | 0
43.493 | 0 | 45.275 | 0 | 53.26 | 3 | 58.276 | 11 | 63.292 | 27 | 73.258 | 0 | 75.238 | 0 | 93.255 | 0 | 118.268 | 0 | 143.281 | 0
43.691 | 0 | 45.473 | 0 | 53.458 | 0 | 58.474 | 13 | 63.49 | 10 | 73.456 | 0 | 75.436 | 7 | 93.453 | 0 | 118.466 | 10 | 143.479 | 0
43.889 | 0 | 45.671 | 7 | 53.656 | 0 | 58.672 | 0 | 63.688 | 15 | 73.654 | 0 | 75.634 | 0 | 93.651 | 0 | 118.664 | 0 | 143.677 | 0
44.087 | 0 | 45.869 | 27 | 53.854 | 0 | 58.87 | 0 | 63.886 | 30 | 73.852 | 0 | 75.832 | 4 | 93.849 | 4 | 118.862 | 0 | 143.875 | 0
44.285 | 0 | 46.067 | 0 | 54.052 | 2 | 59.068 | 0 | 64.084 | 4 | 74.05 | 0 | 76.03 | 0 | 94.047 | 0 | 119.06 | 6 | 144.073 | 17
44.483 | 0 | 46.265 | 19 | 54.25 | 0 | 59.266 | 0 | 64.282 | 0 | 74.248 | 0 | 76.228 | 8 | 94.245 | 0 | 119.258 | 4 | 144.271 | 0
44.681 | 0 | 46.463 | 16 | 54.448 | 0 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 0 | 94.443 | 0 | 119.456 | 5 | 144.469 | 0
44.879 | 0 | 46.661 | 4 | 54.646 | 0 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 0 | 119.654 | 0 | 144.667 | 0
45.077 | 0 | 46.859 | 10 | 54.844 | 9 | 59.86 | 7 | 64.876 | 0 | 74.842 | 0 | 76.822 | 0 | 94.839 | 0 | 119.852 | 8 | 144.865 | 0
45.275 | 0 | 47.056 | 0 | 55.042 | 0 | 60.058 | 4 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 0 | 120.05 | 6 | 145.063 | 0
45.473 | 0 | 47.254 | 4 | 55.24 | 0 | 60.256 | 3 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 0 | 120.248 | 5 | 145.261 | 0
45.671 | 0 | 47.452 | 30 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 0 | 145.459 | 0
45.869 | 0 | 47.65 | 19 | 55.636 | 4 | 60.652 | 27 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 0 | 145.657 | 0
46.067 | 0 | 47.848 | 62 | 55.834 | 0 | 60.85 | 20 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 0 | 120.842 | 0 | 145.855 | 0
46.265 | 0 | 48.046 | 44 | 56.032 | 0 | 61.048 | 38 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 15 | 121.04 | 10 | 146.053 | 0
46.463 | 0 | 48.244 | 19 | 56.23 | 0 | 61.246 | 32 | 66.262 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 0 | 146.251 | 0
46.661 | 0 | 48.442 | 41 | 56.428 | 0 | 61.444 | 23 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 8 | 146.449 | 0
46.859 | 0 | 48.64 | 47 | 56.626 | 0 | 61.642 | 37 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 0 | 146.647 | 4
47.056 | 0 | 48.838 | 19 | 56.824 | 0 | 61.84 | 84 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 0 | 146.845 | 0
47.254 | 0 | 49.036 | 8 | 57.022 | 8 | 62.038 | 108 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 0 | 122.03 | 5 | 147.043 | 0
47.452 | 0 | 49.234 | 16 | 57.22 | 0 | 62.236 | 65 | 67.252 | 4 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 0 | 147.241 | 0
47.65 | 0 | 49.432 | 23 | 57.418 | 13 | 62.434 | 49 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 5 | 122.426 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 8 | 57.616 | 4 | 62.632 | 108 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 6 | 122.624 | 0 | 147.637 | 0
48.046 | 0 | 49.828 | 19 | 57.814 | 3 | 62.83 | 77 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 0 | 122.822 | 4 | 147.835 | 0
48.244 | 0 | 50.026 | 4 | 58.012 | 0 | 63.028 | 58 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 0 | 148.033 | 0
48.442 | 0 | 50.224 | 2 | 58.21 | 13 | 63.226 | 51 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 16 | 58.408 | 0 | 63.424 | 33 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 0 | 148.429 | 0
48.838 | 0 | 50.62 | 19 | 58.606 | 0 | 63.622 | 37 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 0 | 123.614 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 24 | 58.804 | 5 | 63.82 | 23 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 15 | 59.002 | 6 | 64.018 | 11 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 6 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 9 | 59.2 | 0 | 64.216 | 9 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 0 | 124.208 | 0 | 149.221 | 0
|  | 51.412 | 0 | 59.398 | 0 | 64.414 | 16 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 0 | 149.419 | 0

## Analyte Polyalanine n7
0 |  | 2 |  | 10 |  | 20 |  | 32 |  | 50 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 30.029 | 0 | 41.975 | 0 | 59.992 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 30.227 | 0 | 42.173 | 0 | 60.19 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 30.425 | 0 | 42.371 | 0 | 60.388 | 0 | 110.414 | 0
10.824 | 4 | 12.606 | 0 | 20.591 | 0 | 30.623 | 0 | 42.569 | 0 | 60.586 | 0 | 110.612 | 0
11.022 | 8711 | 12.804 | 0 | 20.789 | 0 | 30.821 | 0 | 42.767 | 0 | 60.784 | 0 | 110.81 | 0
11.22 | 99836 | 13.002 | 0 | 20.987 | 1080 | 31.019 | 1185 | 42.965 | 89 | 60.982 | 282 | 111.008 | 263
11.418 | 202731 | 13.2 | 0 | 21.185 | 11262 | 31.217 | 5973 | 43.163 | 3683 | 61.18 | 3839 | 111.206 | 3074
11.616 | 380346 | 13.398 | 0 | 21.383 | 20830 | 31.415 | 10048 | 43.361 | 9169 | 61.378 | 8632 | 111.404 | 8007
11.814 | 531373 | 13.596 | 0 | 21.581 | 24090 | 31.613 | 11578 | 43.559 | 10308 | 61.576 | 8817 | 111.602 | 7944
12.012 | 204457 | 13.794 | 0 | 21.779 | 23642 | 31.811 | 14895 | 43.757 | 8236 | 61.774 | 8518 | 111.8 | 8135
12.21 | 25555 | 13.992 | 0 | 21.977 | 19608 | 32.009 | 14724 | 43.955 | 7798 | 61.972 | 8906 | 111.998 | 8773
12.408 | 2646 | 14.19 | 0 | 22.175 | 17253 | 32.207 | 14031 | 44.153 | 7778 | 62.17 | 9370 | 112.196 | 9254
12.606 | 251 | 14.388 | 0 | 22.373 | 17928 | 32.405 | 11942 | 44.351 | 8511 | 62.368 | 10248 | 112.394 | 10147
12.804 | 13 | 14.586 | 0 | 22.571 | 19080 | 32.603 | 10852 | 44.549 | 8676 | 62.566 | 10609 | 112.592 | 10511
13.002 | 10 | 14.784 | 0 | 22.769 | 18744 | 32.801 | 9149 | 44.747 | 8904 | 62.764 | 10756 | 112.79 | 11726
13.2 | 0 | 14.982 | 0 | 22.967 | 18583 | 32.999 | 8996 | 44.945 | 10228 | 62.962 | 10376 | 112.988 | 12562
13.398 | 0 | 15.18 | 0 | 23.165 | 18283 | 33.197 | 8892 | 45.143 | 10654 | 63.16 | 10599 | 113.186 | 14120
13.596 | 0 | 15.378 | 3 | 23.363 | 15919 | 33.395 | 9070 | 45.341 | 11375 | 63.358 | 10378 | 113.384 | 14482
13.794 | 0 | 15.576 | 32 | 23.561 | 18142 | 33.593 | 9417 | 45.539 | 11971 | 63.556 | 10129 | 113.582 | 14977
13.992 | 0 | 15.773 | 23 | 23.759 | 17437 | 33.791 | 10252 | 45.737 | 12387 | 63.754 | 9587 | 113.78 | 16477
14.19 | 0 | 15.971 | 7 | 23.957 | 17170 | 33.989 | 10587 | 45.935 | 12060 | 63.952 | 8913 | 113.978 | 17572
14.388 | 0 | 16.169 | 12 | 24.155 | 15813 | 34.187 | 11667 | 46.133 | 11677 | 64.15 | 8651 | 114.176 | 19790
14.586 | 0 | 16.367 | 49 | 24.353 | 16112 | 34.385 | 12189 | 46.331 | 10751 | 64.348 | 8731 | 114.374 | 19202
14.784 | 0 | 16.565 | 75 | 24.551 | 16600 | 34.583 | 13601 | 46.529 | 9984 | 64.546 | 7949 | 114.572 | 20675
14.982 | 0 | 16.763 | 148 | 24.749 | 15737 | 34.781 | 12740 | 46.727 | 9080 | 64.744 | 8252 | 114.77 | 21544
15.18 | 0 | 16.961 | 608 | 24.947 | 15314 | 34.979 | 12475 | 46.924 | 8570 | 64.942 | 7988 | 114.968 | 21004
15.378 | 0 | 17.159 | 653 | 25.145 | 14764 | 35.177 | 10846 | 47.122 | 8265 | 65.14 | 7961 | 115.166 | 20603
15.576 | 0 | 17.357 | 343 | 25.343 | 18193 | 35.375 | 9616 | 47.32 | 7867 | 65.338 | 8293 | 115.364 | 20374
15.773 | 0 | 17.555 | 535 | 25.541 | 27655 | 35.573 | 9758 | 47.518 | 7890 | 65.536 | 8204 | 115.562 | 19518
15.971 | 0 | 17.753 | 1007 | 25.739 | 65442 | 35.771 | 9941 | 47.716 | 8367 | 65.734 | 8474 | 115.76 | 19132
16.169 | 0 | 17.951 | 969 | 25.937 | 138479 | 35.969 | 10225 | 47.914 | 8704 | 65.932 | 8635 | 115.958 | 19880
16.367 | 0 | 18.149 | 1136 | 26.135 | 232190 | 36.167 | 10260 | 48.112 | 9697 | 66.13 | 8107 | 116.156 | 19506
16.565 | 0 | 18.347 | 2557 | 26.333 | 281858 | 36.365 | 9728 | 48.31 | 9275 | 66.328 | 8229 | 116.354 | 19059
16.763 | 0 | 18.545 | 4287 | 26.531 | 253637 | 36.563 | 10096 | 48.508 | 9892 | 66.526 | 8351 | 116.552 | 18272
16.961 | 0 | 18.743 | 3595 | 26.729 | 174130 | 36.761 | 10518 | 48.706 | 10531 | 66.724 | 8689 | 116.75 | 17088
17.159 | 0 | 18.941 | 3311 | 26.927 | 88685 | 36.959 | 12460 | 48.904 | 10551 | 66.922 | 8257 | 116.948 | 15330
17.357 | 0 | 19.139 | 6517 | 27.125 | 37657 | 37.157 | 13340 | 49.102 | 11552 | 67.12 | 9022 | 117.146 | 13984
17.555 | 0 | 19.337 | 10352 | 27.323 | 17057 | 37.355 | 14367 | 49.3 | 11805 | 67.318 | 8764 | 117.344 | 12596
17.753 | 0 | 19.535 | 7996 | 27.521 | 9951 | 37.553 | 13775 | 49.498 | 11675 | 67.516 | 9188 | 117.542 | 11206
17.951 | 0 | 19.733 | 5823 | 27.719 | 7661 | 37.751 | 11962 | 49.696 | 11662 | 67.714 | 9349 | 117.74 | 10806
18.149 | 0 | 19.931 | 7925 | 27.917 | 6847 | 37.949 | 10700 | 49.894 | 12131 | 67.912 | 9097 | 117.938 | 9624
18.347 | 0 | 20.129 | 13286 | 28.115 | 7573 | 38.147 | 9841 | 50.092 | 12236 | 68.11 | 9833 | 118.136 | 8548
18.545 | 0 | 20.327 | 12841 | 28.313 | 8859 | 38.345 | 8536 | 50.29 | 11965 | 68.308 | 10175 | 118.334 | 7583
18.743 | 0 | 20.525 | 9581 | 28.511 | 11539 | 38.543 | 8511 | 50.488 | 11508 | 68.506 | 10848 | 118.532 | 7368
18.941 | 0 | 20.723 | 10654 | 28.709 | 11867 | 38.741 | 8389 | 50.686 | 11417 | 68.704 | 12810 | 118.73 | 6476
19.139 | 0 | 20.921 | 16472 | 28.907 | 11602 | 38.939 | 8984 | 50.884 | 11041 | 68.902 | 15733 | 118.928 | 5976
19.337 | 0 | 21.119 | 20773 | 29.105 | 10090 | 39.137 | 9729 | 51.082 | 9925 | 69.1 | 18560 | 119.126 | 5492
19.535 | 0 | 21.317 | 19518 | 29.303 | 7795 | 39.335 | 10828 | 51.28 | 9888 | 69.298 | 22388 | 119.324 | 5206
19.733 | 0 | 21.515 | 19018 | 29.501 | 6647 | 39.533 | 12233 | 51.478 | 8922 | 69.496 | 27267 | 119.522 | 5169
19.931 | 0 | 21.713 | 22347 | 29.699 | 5574 | 39.731 | 14347 | 51.676 | 8797 | 69.694 | 31851 | 119.72 | 4947
20.129 | 0 | 21.911 | 23468 | 29.897 | 6613 | 39.929 | 19426 | 51.874 | 9072 | 69.892 | 36765 | 119.918 | 4955
20.327 | 0 | 22.109 | 21588 | 30.095 | 8519 | 40.127 | 29747 | 52.072 | 9271 | 70.09 | 41178 | 120.116 | 4892
20.525 | 0 | 22.307 | 18554 | 30.293 | 10996 | 40.325 | 47203 | 52.27 | 10122 | 70.288 | 45672 | 120.314 | 4515
20.723 | 0 | 22.505 | 17740 | 30.491 | 13033 | 40.523 | 69836 | 52.468 | 10506 | 70.486 | 50294 | 120.512 | 4755
20.921 | 0 | 22.703 | 18572 | 30.689 | 12684 | 40.721 | 97246 | 52.666 | 11206 | 70.684 | 51751 | 120.71 | 4734
21.119 | 0 | 22.901 | 19478 | 30.887 | 10564 | 40.919 | 125817 | 52.864 | 10950 | 70.882 | 51796 | 120.908 | 4467
21.317 | 0 | 23.099 | 19826 | 31.085 | 8013 | 41.117 | 146282 | 53.062 | 10131 | 71.08 | 52542 | 121.106 | 4289
21.515 | 0 | 23.297 | 18428 | 31.283 | 4620 | 41.315 | 152312 | 53.26 | 10158 | 71.278 | 51281 | 121.304 | 4692
21.713 | 0 | 23.495 | 17289 | 31.481 | 2291 | 41.513 | 143927 | 53.458 | 10053 | 71.476 | 48227 | 121.502 | 4241
21.911 | 0 | 23.693 | 17756 | 31.679 | 1528 | 41.711 | 122093 | 53.656 | 9285 | 71.674 | 42978 | 121.7 | 4248
22.109 | 0 | 23.891 | 18439 | 31.877 | 1184 | 41.909 | 91994 | 53.854 | 8204 | 71.872 | 38449 | 121.898 | 3975
22.307 | 0 | 24.089 | 17147 | 32.075 | 991 | 42.107 | 60795 | 54.052 | 6816 | 72.07 | 32739 | 122.096 | 3988
22.505 | 0 | 24.287 | 15757 | 32.273 | 936 | 42.305 | 35864 | 54.25 | 7901 | 72.268 | 27097 | 122.294 | 3821
22.703 | 0 | 24.485 | 16458 | 32.471 | 794 | 42.503 | 19155 | 54.448 | 11256 | 72.466 | 20960 | 122.492 | 3470
22.901 | 0 | 24.683 | 15922 | 32.669 | 679 | 42.701 | 9761 | 54.646 | 17642 | 72.664 | 16680 | 122.69 | 3153
23.099 | 0 | 24.881 | 15378 | 32.867 | 604 | 42.899 | 6000 | 54.844 | 25340 | 72.862 | 12826 | 122.888 | 2972
23.297 | 0 | 25.079 | 13940 | 33.065 | 462 | 43.097 | 4230 | 55.042 | 34937 | 73.06 | 9597 | 123.086 | 3346
23.495 | 0 | 25.277 | 14028 | 33.263 | 500 | 43.295 | 3518 | 55.24 | 46299 | 73.258 | 7909 | 123.284 | 2962
23.693 | 0 | 25.475 | 14857 | 33.461 | 360 | 43.493 | 3444 | 55.438 | 59385 | 73.456 | 5948 | 123.482 | 3213
23.891 | 0 | 25.673 | 19508 | 33.659 | 372 | 43.691 | 3413 | 55.636 | 70433 | 73.654 | 5097 | 123.68 | 2475
24.089 | 0 | 25.871 | 41641 | 33.857 | 377 | 43.889 | 3112 | 55.834 | 78382 | 73.852 | 3792 | 123.878 | 2371
24.287 | 0 | 26.069 | 103800 | 34.055 | 263 | 44.087 | 2898 | 56.032 | 84710 | 74.05 | 3130 | 124.076 | 2250
24.485 | 0 | 26.267 | 199337 | 34.253 | 323 | 44.285 | 2780 | 56.23 | 85129 | 74.248 | 2113 | 124.274 | 1736
24.683 | 0 | 26.465 | 283293 | 34.451 | 337 | 44.483 | 2879 | 56.428 | 84597 | 74.446 | 2031 | 124.472 | 1732
24.881 | 0 | 26.663 | 295496 | 34.649 | 318 | 44.681 | 2976 | 56.626 | 77109 | 74.644 | 1859 | 124.67 | 1899
25.079 | 0 | 26.861 | 223211 | 34.847 | 269 | 44.879 | 2699 | 56.824 | 69192 | 74.842 | 1924 | 124.868 | 1685
25.277 | 0 | 27.059 | 128131 | 35.045 | 230 | 45.077 | 2518 | 57.022 | 56961 | 75.04 | 1840 | 125.066 | 1527
25.475 | 0 | 27.257 | 52652 | 35.243 | 246 | 45.275 | 2247 | 57.22 | 43860 | 75.238 | 1538 | 125.264 | 1705
25.673 | 0 | 27.455 | 19718 | 35.441 | 172 | 45.473 | 2211 | 57.418 | 30946 | 75.436 | 1627 | 125.462 | 1951
25.871 | 0 | 27.653 | 9054 | 35.639 | 226 | 45.671 | 2035 | 57.616 | 20822 | 75.634 | 1515 | 125.66 | 1492
26.069 | 0 | 27.851 | 6324 | 35.837 | 214 | 45.869 | 2028 | 57.814 | 13134 | 75.832 | 1588 | 125.858 | 1276
26.267 | 0 | 28.049 | 6083 | 36.035 | 188 | 46.067 | 1839 | 58.012 | 8149 | 76.03 | 1483 | 126.056 | 995
26.465 | 0 | 28.247 | 5712 | 36.233 | 114 | 46.265 | 1982 | 58.21 | 5029 | 76.228 | 1484 | 126.254 | 975
26.663 | 0 | 28.445 | 5254 | 36.431 | 96 | 46.463 | 1718 | 58.408 | 3646 | 76.426 | 1356 | 126.452 | 754
26.861 | 0 | 28.643 | 4796 | 36.629 | 64 | 46.661 | 1777 | 58.606 | 2726 | 76.624 | 1555 | 126.65 | 773
27.059 | 0 | 28.841 | 4480 | 36.827 | 64 | 46.859 | 1891 | 58.804 | 2747 | 76.822 | 1438 | 126.848 | 697
27.257 | 0 | 29.039 | 4148 | 37.025 | 102 | 47.056 | 1722 | 59.002 | 2474 | 77.02 | 1303 | 127.046 | 725
27.455 | 0 | 29.237 | 4029 | 37.223 | 57 | 47.254 | 1675 | 59.2 | 2024 | 77.218 | 1188 | 127.244 | 584
27.653 | 0 | 29.435 | 4114 | 37.421 | 82 | 47.452 | 1685 | 59.398 | 1649 | 77.416 | 1195 | 127.442 | 611
27.851 | 0 | 29.633 | 3708 | 37.619 | 43 | 47.65 | 1662 | 59.596 | 968 | 77.614 | 1098 | 127.64 | 599
28.049 | 0 | 29.831 | 3138 | 37.817 | 30 | 47.848 | 1416 | 59.794 | 439 | 77.812 | 1020 | 127.838 | 626
28.247 | 0 | 30.029 | 2920 | 38.015 | 49 | 48.046 | 1412 | 59.992 | 197 | 78.01 | 865 | 128.036 | 562
28.445 | 0 | 30.227 | 2540 | 38.213 | 9 | 48.244 | 1144 | 60.19 | 74 | 78.208 | 707 | 128.234 | 414
28.643 | 0 | 30.425 | 2407 | 38.411 | 6 | 48.442 | 1139 | 60.388 | 38 | 78.405 | 408 | 128.432 | 249
28.841 | 0 | 30.623 | 2395 | 38.609 | 30 | 48.64 | 1065 | 60.586 | 40 | 78.603 | 167 | 128.63 | 234
29.039 | 0 | 30.821 | 2021 | 38.807 | 5 | 48.838 | 878 | 60.784 | 54 | 78.801 | 126 | 128.828 | 176
29.237 | 0 | 31.019 | 2039 | 39.005 | 31 | 49.036 | 1216 | 60.982 | 28 | 78.999 | 65 | 129.026 | 159
29.435 | 0 | 31.217 | 1967 | 39.203 | 15 | 49.234 | 1050 | 61.18 | 6 | 79.197 | 68 | 129.224 | 182
29.633 | 0 | 31.415 | 1402 | 39.401 | 51 | 49.432 | 1042 | 61.378 | 14 | 79.395 | 54 | 129.422 | 124
29.831 | 0 | 31.613 | 1530 | 39.599 | 20 | 49.63 | 897 | 61.576 | 18 | 79.593 | 67 | 129.62 | 158
30.029 | 0 | 31.811 | 1083 | 39.797 | 6 | 49.828 | 1043 | 61.774 | 8 | 79.791 | 81 | 129.818 | 113
30.227 | 0 | 32.009 | 1251 | 39.995 | 3 | 50.026 | 851 | 61.972 | 18 | 79.989 | 36 | 130.016 | 110
30.425 | 0 | 32.207 | 1174 | 40.193 | 9 | 50.224 | 905 | 62.17 | 25 | 80.187 | 41 | 130.214 | 121
30.623 | 0 | 32.405 | 976 | 40.391 | 8 | 50.422 | 706 | 62.368 | 0 | 80.385 | 44 | 130.412 | 118
30.821 | 0 | 32.603 | 817 | 40.589 | 9 | 50.62 | 599 | 62.566 | 17 | 80.583 | 59 | 130.61 | 72
31.019 | 0 | 32.801 | 918 | 40.787 | 0 | 50.818 | 519 | 62.764 | 17 | 80.781 | 35 | 130.808 | 146
31.217 | 0 | 32.999 | 631 | 40.985 | 8 | 51.016 | 597 | 62.962 | 7 | 80.979 | 33 | 131.006 | 94
31.415 | 0 | 33.197 | 467 | 41.183 | 9 | 51.214 | 565 | 63.16 | 22 | 81.177 | 54 | 131.204 | 115
31.613 | 0 | 33.395 | 466 | 41.381 | 4 | 51.412 | 295 | 63.358 | 0 | 81.375 | 49 | 131.402 | 73
31.811 | 0 | 33.593 | 503 | 41.579 | 4 | 51.61 | 282 | 63.556 | 15 | 81.573 | 35 | 131.6 | 109
32.009 | 0 | 33.791 | 466 | 41.777 | 5 | 51.808 | 119 | 63.754 | 7 | 81.771 | 23 | 131.798 | 35
32.207 | 0 | 33.989 | 402 | 41.975 | 8 | 52.006 | 109 | 63.952 | 0 | 81.969 | 46 | 131.996 | 19
32.405 | 0 | 34.187 | 311 | 42.173 | 0 | 52.204 | 54 | 64.15 | 2 | 82.167 | 51 | 132.194 | 18
32.603 | 0 | 34.385 | 265 | 42.371 | 23 | 52.402 | 33 | 64.348 | 12 | 82.365 | 12 | 132.392 | 0
32.801 | 0 | 34.583 | 190 | 42.569 | 3 | 52.6 | 20 | 64.546 | 9 | 82.563 | 17 | 132.59 | 0
32.999 | 0 | 34.781 | 265 | 42.767 | 7 | 52.798 | 27 | 64.744 | 5 | 82.761 | 37 | 132.788 | 20
33.197 | 0 | 34.979 | 202 | 42.965 | 4 | 52.996 | 0 | 64.942 | 8 | 82.959 | 35 | 132.986 | 10
33.395 | 0 | 35.177 | 195 | 43.163 | 18 | 53.194 | 3 | 65.14 | 8 | 83.157 | 47 | 133.184 | 11
33.593 | 0 | 35.375 | 199 | 43.361 | 19 | 53.392 | 9 | 65.338 | 8 | 83.355 | 22 | 133.382 | 13
33.791 | 0 | 35.573 | 133 | 43.559 | 25 | 53.59 | 10 | 65.536 | 3 | 83.553 | 6 | 133.58 | 0
33.989 | 0 | 35.771 | 106 | 43.757 | 10 | 53.788 | 5 | 65.734 | 6 | 83.751 | 0 | 133.778 | 4
34.187 | 0 | 35.969 | 196 | 43.955 | 35 | 53.986 | 0 | 65.932 | 0 | 83.949 | 28 | 133.976 | 21
34.385 | 0 | 36.167 | 50 | 44.153 | 22 | 54.184 | 0 | 66.13 | 8 | 84.147 | 27 | 134.174 | 23
34.583 | 0 | 36.365 | 77 | 44.351 | 16 | 54.382 | 0 | 66.328 | 0 | 84.345 | 27 | 134.372 | 12
34.781 | 0 | 36.563 | 123 | 44.549 | 18 | 54.58 | 0 | 66.526 | 0 | 84.543 | 18 | 134.57 | 10
34.979 | 0 | 36.761 | 54 | 44.747 | 11 | 54.778 | 0 | 66.724 | 14 | 84.741 | 0 | 134.768 | 0
35.177 | 0 | 36.959 | 95 | 44.945 | 34 | 54.976 | 0 | 66.922 | 3 | 84.939 | 47 | 134.966 | 3
35.375 | 0 | 37.157 | 47 | 45.143 | 11 | 55.174 | 0 | 67.12 | 6 | 85.137 | 3 | 135.164 | 11
35.573 | 0 | 37.355 | 28 | 45.341 | 22 | 55.372 | 0 | 67.318 | 12 | 85.335 | 8 | 135.362 | 13
35.771 | 0 | 37.553 | 37 | 45.539 | 9 | 55.57 | 6 | 67.516 | 2 | 85.533 | 24 | 135.56 | 6
35.969 | 0 | 37.751 | 49 | 45.737 | 8 | 55.768 | 0 | 67.714 | 4 | 85.731 | 22 | 135.758 | 18
36.167 | 0 | 37.949 | 59 | 45.935 | 5 | 55.966 | 0 | 67.912 | 0 | 85.929 | 6 | 135.956 | 9
36.365 | 0 | 38.147 | 12 | 46.133 | 4 | 56.164 | 0 | 68.11 | 0 | 86.127 | 7 | 136.154 | 2
36.563 | 0 | 38.345 | 40 | 46.331 | 7 | 56.362 | 0 | 68.308 | 0 | 86.325 | 19 | 136.352 | 10
36.761 | 0 | 38.543 | 17 | 46.529 | 9 | 56.56 | 0 | 68.506 | 0 | 86.523 | 13 | 136.55 | 12
36.959 | 0 | 38.741 | 34 | 46.727 | 0 | 56.758 | 0 | 68.704 | 12 | 86.721 | 7 | 136.748 | 0
37.157 | 0 | 38.939 | 39 | 46.924 | 0 | 56.956 | 0 | 68.902 | 6 | 86.919 | 0 | 136.946 | 24
37.355 | 0 | 39.137 | 10 | 47.122 | 0 | 57.154 | 0 | 69.1 | 0 | 87.117 | 0 | 137.144 | 0
37.553 | 0 | 39.335 | 19 | 47.32 | 0 | 57.352 | 0 | 69.298 | 5 | 87.315 | 0 | 137.342 | 6
37.751 | 0 | 39.533 | 28 | 47.518 | 3 | 57.55 | 0 | 69.496 | 0 | 87.513 | 0 | 137.54 | 5
37.949 | 0 | 39.731 | 15 | 47.716 | 0 | 57.748 | 0 | 69.694 | 7 | 87.711 | 0 | 137.738 | 0
38.147 | 0 | 39.929 | 30 | 47.914 | 0 | 57.946 | 0 | 69.892 | 12 | 87.909 | 0 | 137.936 | 0
38.345 | 0 | 40.127 | 10 | 48.112 | 0 | 58.144 | 0 | 70.09 | 6 | 88.107 | 0 | 138.134 | 0
38.543 | 0 | 40.325 | 12 | 48.31 | 0 | 58.342 | 0 | 70.288 | 13 | 88.305 | 0 | 138.332 | 0
38.741 | 0 | 40.523 | 9 | 48.508 | 0 | 58.54 | 0 | 70.486 | 9 | 88.503 | 0 | 138.53 | 0
38.939 | 0 | 40.721 | 15 | 48.706 | 5 | 58.738 | 0 | 70.684 | 0 | 88.701 | 5 | 138.728 | 8
39.137 | 0 | 40.919 | 21 | 48.904 | 0 | 58.936 | 0 | 70.882 | 15 | 88.899 | 0 | 138.926 | 0
39.335 | 0 | 41.117 | 25 | 49.102 | 0 | 59.134 | 0 | 71.08 | 3 | 89.097 | 0 | 139.124 | 0
39.533 | 0 | 41.315 | 13 | 49.3 | 4 | 59.332 | 0 | 71.278 | 15 | 89.295 | 0 | 139.322 | 7
39.731 | 0 | 41.513 | 5 | 49.498 | 0 | 59.53 | 0 | 71.476 | 7 | 89.493 | 0 | 139.52 | 0
39.929 | 0 | 41.711 | 4 | 49.696 | 6 | 59.728 | 0 | 71.674 | 8 | 89.691 | 0 | 139.718 | 0
40.127 | 0 | 41.909 | 5 | 49.894 | 8 | 59.926 | 0 | 71.872 | 0 | 89.889 | 0 | 139.916 | 7
40.325 | 0 | 42.107 | 16 | 50.092 | 0 | 60.124 | 0 | 72.07 | 9 | 90.087 | 0 | 140.114 | 0
40.523 | 0 | 42.305 | 3 | 50.29 | 0 | 60.322 | 0 | 72.268 | 0 | 90.285 | 0 | 140.312 | 0
40.721 | 0 | 42.503 | 22 | 50.488 | 4 | 60.52 | 0 | 72.466 | 5 | 90.483 | 0 | 140.51 | 7
40.919 | 0 | 42.701 | 16 | 50.686 | 2 | 60.718 | 0 | 72.664 | 0 | 90.681 | 0 | 140.708 | 0
41.117 | 0 | 42.899 | 0 | 50.884 | 5 | 60.916 | 0 | 72.862 | 4 | 90.879 | 0 | 140.906 | 0
41.315 | 0 | 43.097 | 23 | 51.082 | 4 | 61.114 | 0 | 73.06 | 0 | 91.077 | 0 | 141.103 | 0
41.513 | 0 | 43.295 | 18 | 51.28 | 7 | 61.312 | 0 | 73.258 | 0 | 91.275 | 0 | 141.301 | 6
41.711 | 0 | 43.493 | 32 | 51.478 | 0 | 61.51 | 0 | 73.456 | 4 | 91.473 | 0 | 141.499 | 0
41.909 | 0 | 43.691 | 22 | 51.676 | 0 | 61.708 | 0 | 73.654 | 0 | 91.671 | 0 | 141.697 | 13
42.107 | 0 | 43.889 | 3 | 51.874 | 4 | 61.906 | 0 | 73.852 | 7 | 91.869 | 0 | 141.895 | 0
42.305 | 0 | 44.087 | 37 | 52.072 | 0 | 62.104 | 0 | 74.05 | 0 | 92.067 | 0 | 142.093 | 0
42.503 | 0 | 44.285 | 32 | 52.27 | 0 | 62.302 | 0 | 74.248 | 0 | 92.265 | 0 | 142.291 | 0
42.701 | 0 | 44.483 | 25 | 52.468 | 0 | 62.5 | 0 | 74.446 | 3 | 92.463 | 0 | 142.489 | 0
42.899 | 0 | 44.681 | 40 | 52.666 | 0 | 62.698 | 0 | 74.644 | 6 | 92.661 | 0 | 142.687 | 0
43.097 | 0 | 44.879 | 28 | 52.864 | 0 | 62.896 | 0 | 74.842 | 0 | 92.859 | 0 | 142.885 | 8
43.295 | 0 | 45.077 | 17 | 53.062 | 9 | 63.094 | 0 | 75.04 | 0 | 93.057 | 0 | 143.083 | 0
43.493 | 0 | 45.275 | 10 | 53.26 | 0 | 63.292 | 0 | 75.238 | 6 | 93.255 | 0 | 143.281 | 0
43.691 | 0 | 45.473 | 30 | 53.458 | 0 | 63.49 | 0 | 75.436 | 8 | 93.453 | 0 | 143.479 | 0
43.889 | 0 | 45.671 | 6 | 53.656 | 6 | 63.688 | 0 | 75.634 | 4 | 93.651 | 0 | 143.677 | 0
44.087 | 0 | 45.869 | 14 | 53.854 | 0 | 63.886 | 0 | 75.832 | 0 | 93.849 | 0 | 143.875 | 0
44.285 | 0 | 46.067 | 0 | 54.052 | 0 | 64.084 | 0 | 76.03 | 0 | 94.047 | 0 | 144.073 | 0
44.483 | 0 | 46.265 | 11 | 54.25 | 4 | 64.282 | 0 | 76.228 | 0 | 94.245 | 0 | 144.271 | 0
44.681 | 0 | 46.463 | 16 | 54.448 | 4 | 64.48 | 0 | 76.426 | 0 | 94.443 | 0 | 144.469 | 0
44.879 | 0 | 46.661 | 41 | 54.646 | 0 | 64.678 | 0 | 76.624 | 0 | 94.641 | 0 | 144.667 | 0
45.077 | 0 | 46.859 | 8 | 54.844 | 0 | 64.876 | 0 | 76.822 | 0 | 94.839 | 0 | 144.865 | 4
45.275 | 0 | 47.056 | 9 | 55.042 | 0 | 65.074 | 0 | 77.02 | 0 | 95.037 | 0 | 145.063 | 6
45.473 | 0 | 47.254 | 10 | 55.24 | 0 | 65.272 | 0 | 77.218 | 0 | 95.235 | 0 | 145.261 | 0
45.671 | 0 | 47.452 | 0 | 55.438 | 0 | 65.47 | 0 | 77.416 | 0 | 95.433 | 0 | 145.459 | 10
45.869 | 0 | 47.65 | 0 | 55.636 | 0 | 65.668 | 0 | 77.614 | 0 | 95.631 | 0 | 145.657 | 0
46.067 | 0 | 47.848 | 4 | 55.834 | 0 | 65.866 | 0 | 77.812 | 0 | 95.829 | 0 | 145.855 | 8
46.265 | 0 | 48.046 | 0 | 56.032 | 7 | 66.064 | 0 | 78.01 | 0 | 96.027 | 0 | 146.053 | 0
46.463 | 0 | 48.244 | 0 | 56.23 | 0 | 66.262 | 0 | 78.208 | 0 | 96.225 | 0 | 146.251 | 0
46.661 | 0 | 48.442 | 14 | 56.428 | 0 | 66.46 | 0 | 78.405 | 0 | 96.423 | 0 | 146.449 | 0
46.859 | 0 | 48.64 | 7 | 56.626 | 0 | 66.658 | 0 | 78.603 | 0 | 96.621 | 0 | 146.647 | 0
47.056 | 0 | 48.838 | 15 | 56.824 | 0 | 66.856 | 0 | 78.801 | 0 | 96.819 | 0 | 146.845 | 0
47.254 | 0 | 49.036 | 0 | 57.022 | 0 | 67.054 | 0 | 78.999 | 0 | 97.017 | 0 | 147.043 | 0
47.452 | 0 | 49.234 | 0 | 57.22 | 0 | 67.252 | 0 | 79.197 | 0 | 97.215 | 0 | 147.241 | 0
47.65 | 0 | 49.432 | 0 | 57.418 | 0 | 67.45 | 0 | 79.395 | 0 | 97.413 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 13 | 57.616 | 0 | 67.648 | 0 | 79.593 | 0 | 97.611 | 0 | 147.637 | 0
48.046 | 0 | 49.828 | 0 | 57.814 | 0 | 67.846 | 0 | 79.791 | 0 | 97.809 | 0 | 147.835 | 0
48.244 | 0 | 50.026 | 7 | 58.012 | 0 | 68.044 | 0 | 79.989 | 0 | 98.007 | 0 | 148.033 | 0
48.442 | 0 | 50.224 | 0 | 58.21 | 0 | 68.242 | 0 | 80.187 | 0 | 98.205 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 5 | 58.408 | 0 | 68.44 | 0 | 80.385 | 0 | 98.403 | 0 | 148.429 | 0
48.838 | 0 | 50.62 | 9 | 58.606 | 0 | 68.638 | 0 | 80.583 | 0 | 98.601 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 0 | 58.804 | 0 | 68.836 | 0 | 80.781 | 0 | 98.799 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 2 | 59.002 | 0 | 69.034 | 0 | 80.979 | 0 | 98.997 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 18 | 59.2 | 0 | 69.232 | 0 | 81.177 | 0 | 99.195 | 0 | 149.221 | 0
|  | 51.412 | 6 | 59.398 | 0 | 69.43 | 0 | 81.375 | 0 | 99.393 | 0 | 149.419 | 0

## Analyte Leucine Enkephalin
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 10 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 457 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 5257 | 13.002 | 0 | 20.987 | 79 | 26.003 | 37 | 31.019 | 124 | 40.985 | 30 | 42.965 | 43 | 60.982 | 23 | 85.995 | 89 | 111.008 | 49
11.418 | 39534 | 13.2 | 0 | 21.185 | 853 | 26.201 | 214 | 31.217 | 427 | 41.183 | 361 | 43.163 | 714 | 61.18 | 1394 | 86.193 | 1046 | 111.206 | 698
11.616 | 73037 | 13.398 | 0 | 21.383 | 1069 | 26.399 | 400 | 31.415 | 521 | 41.381 | 630 | 43.361 | 833 | 61.378 | 2140 | 86.391 | 1189 | 111.404 | 814
11.814 | 411452 | 13.596 | 6 | 21.581 | 1653 | 26.597 | 549 | 31.613 | 1345 | 41.579 | 825 | 43.559 | 1317 | 61.576 | 1499 | 86.589 | 1821 | 111.602 | 752
12.012 | 1032925 | 13.794 | 0 | 21.779 | 1848 | 26.795 | 693 | 31.811 | 1566 | 41.777 | 1015 | 43.757 | 1256 | 61.774 | 1647 | 86.787 | 1084 | 111.8 | 956
12.21 | 789916 | 13.992 | 0 | 21.977 | 1813 | 26.993 | 541 | 32.009 | 1191 | 41.975 | 1132 | 43.955 | 1106 | 61.972 | 1281 | 86.985 | 1015 | 111.998 | 659
12.408 | 137444 | 14.19 | 0 | 22.175 | 1772 | 27.191 | 790 | 32.207 | 1198 | 42.173 | 1069 | 44.153 | 1100 | 62.17 | 1317 | 87.183 | 978 | 112.196 | 713
12.606 | 6442 | 14.388 | 0 | 22.373 | 2234 | 27.389 | 814 | 32.405 | 1392 | 42.371 | 967 | 44.351 | 1182 | 62.368 | 1059 | 87.381 | 706 | 112.394 | 752
12.804 | 390 | 14.586 | 8 | 22.571 | 3977 | 27.587 | 912 | 32.603 | 1067 | 42.569 | 1003 | 44.549 | 1043 | 62.566 | 966 | 87.579 | 676 | 112.592 | 812
13.002 | 24 | 14.784 | 0 | 22.769 | 7523 | 27.785 | 916 | 32.801 | 1213 | 42.767 | 991 | 44.747 | 1265 | 62.764 | 902 | 87.777 | 751 | 112.79 | 570
13.2 | 6 | 14.982 | 0 | 22.967 | 12279 | 27.983 | 946 | 32.999 | 1063 | 42.965 | 1103 | 44.945 | 1151 | 62.962 | 1176 | 87.975 | 688 | 112.988 | 671
13.398 | 0 | 15.18 | 0 | 23.165 | 16550 | 28.181 | 952 | 33.197 | 1097 | 43.163 | 991 | 45.143 | 1333 | 63.16 | 1053 | 88.173 | 710 | 113.186 | 606
13.596 | 0 | 15.378 | 0 | 23.363 | 15479 | 28.379 | 1066 | 33.395 | 1060 | 43.361 | 1094 | 45.341 | 1789 | 63.358 | 807 | 88.371 | 600 | 113.384 | 749
13.794 | 0 | 15.576 | 0 | 23.561 | 12616 | 28.577 | 1143 | 33.593 | 1264 | 43.559 | 1045 | 45.539 | 1829 | 63.556 | 968 | 88.569 | 603 | 113.582 | 742
13.992 | 0 | 15.773 | 4 | 23.759 | 8210 | 28.775 | 2142 | 33.791 | 1696 | 43.757 | 1109 | 45.737 | 2796 | 63.754 | 920 | 88.767 | 525 | 113.78 | 664
14.19 | 0 | 15.971 | 0 | 23.957 | 5113 | 28.973 | 12051 | 33.989 | 2195 | 43.955 | 1007 | 45.935 | 3628 | 63.952 | 1060 | 88.965 | 640 | 113.978 | 777
14.388 | 0 | 16.169 | 8 | 24.155 | 2927 | 29.171 | 68747 | 34.187 | 2962 | 44.153 | 1270 | 46.133 | 3532 | 64.15 | 777 | 89.163 | 518 | 114.176 | 782
14.586 | 0 | 16.367 | 26 | 24.353 | 1274 | 29.369 | 252264 | 34.385 | 4829 | 44.351 | 1082 | 46.331 | 4539 | 64.348 | 778 | 89.361 | 547 | 114.374 | 951
14.784 | 0 | 16.565 | 57 | 24.551 | 720 | 29.567 | 589250 | 34.583 | 5583 | 44.549 | 735 | 46.529 | 4795 | 64.546 | 773 | 89.559 | 542 | 114.572 | 980
14.982 | 0 | 16.763 | 20 | 24.749 | 504 | 29.765 | 906312 | 34.781 | 6896 | 44.747 | 1171 | 46.727 | 6401 | 64.744 | 1399 | 89.757 | 679 | 114.77 | 958
15.18 | 0 | 16.961 | 20 | 24.947 | 568 | 29.963 | 952315 | 34.979 | 8483 | 44.945 | 1239 | 46.924 | 10484 | 64.942 | 3280 | 89.955 | 760 | 114.968 | 961
15.378 | 0 | 17.159 | 8 | 25.145 | 365 | 30.161 | 693152 | 35.177 | 9250 | 45.143 | 1560 | 47.122 | 26371 | 65.14 | 9172 | 90.153 | 724 | 115.166 | 869
15.576 | 0 | 17.357 | 16 | 25.343 | 386 | 30.359 | 332920 | 35.375 | 9299 | 45.341 | 1726 | 47.32 | 67697 | 65.338 | 24727 | 90.351 | 769 | 115.364 | 887
15.773 | 0 | 17.555 | 33 | 25.541 | 646 | 30.557 | 106594 | 35.573 | 7926 | 45.539 | 2059 | 47.518 | 158392 | 65.536 | 57835 | 90.549 | 714 | 115.562 | 905
15.971 | 0 | 17.753 | 30 | 25.739 | 553 | 30.755 | 22461 | 35.771 | 7203 | 45.737 | 2483 | 47.716 | 295345 | 65.734 | 109595 | 90.747 | 681 | 115.76 | 950
16.169 | 0 | 17.951 | 20 | 25.937 | 599 | 30.953 | 3582 | 35.969 | 5888 | 45.935 | 3356 | 47.914 | 442110 | 65.932 | 182339 | 90.945 | 798 | 115.958 | 1068
16.367 | 0 | 18.149 | 24 | 26.135 | 494 | 31.151 | 1231 | 36.167 | 4007 | 46.133 | 3896 | 48.112 | 535661 | 66.13 | 258549 | 91.143 | 730 | 116.156 | 898
16.565 | 0 | 18.347 | 64 | 26.333 | 562 | 31.349 | 1047 | 36.365 | 3453 | 46.331 | 4528 | 48.31 | 522019 | 66.328 | 303859 | 91.341 | 827 | 116.354 | 1000
16.763 | 0 | 18.545 | 60 | 26.531 | 425 | 31.547 | 1012 | 36.563 | 3353 | 46.529 | 4961 | 48.508 | 403285 | 66.526 | 308446 | 91.539 | 748 | 116.552 | 959
16.961 | 0 | 18.743 | 125 | 26.729 | 462 | 31.745 | 1170 | 36.761 | 2366 | 46.727 | 5898 | 48.706 | 249291 | 66.724 | 270164 | 91.737 | 625 | 116.75 | 937
17.159 | 0 | 18.941 | 139 | 26.927 | 472 | 31.943 | 1214 | 36.959 | 1716 | 46.924 | 8754 | 48.904 | 126940 | 66.922 | 204392 | 91.935 | 843 | 116.948 | 1170
17.357 | 0 | 19.139 | 216 | 27.125 | 682 | 32.141 | 1341 | 37.157 | 1238 | 47.122 | 21328 | 49.102 | 51298 | 67.12 | 129530 | 92.133 | 1017 | 117.146 | 1252
17.555 | 0 | 19.337 | 271 | 27.323 | 808 | 32.339 | 1280 | 37.355 | 1174 | 47.32 | 56853 | 49.3 | 17364 | 67.318 | 73959 | 92.331 | 1205 | 117.344 | 1076
17.753 | 0 | 19.535 | 319 | 27.521 | 968 | 32.537 | 1195 | 37.553 | 922 | 47.518 | 136998 | 49.498 | 5343 | 67.516 | 34153 | 92.529 | 1096 | 117.542 | 1238
17.951 | 0 | 19.733 | 245 | 27.719 | 1003 | 32.735 | 1295 | 37.751 | 872 | 47.716 | 272661 | 49.696 | 2301 | 67.714 | 13780 | 92.727 | 1185 | 117.74 | 1234
18.149 | 0 | 19.931 | 301 | 27.917 | 959 | 32.933 | 1109 | 37.949 | 860 | 47.914 | 439102 | 49.894 | 1212 | 67.912 | 5287 | 92.925 | 1405 | 117.938 | 1408
18.347 | 0 | 20.129 | 293 | 28.115 | 854 | 33.131 | 1562 | 38.147 | 792 | 48.112 | 562950 | 50.092 | 1205 | 68.11 | 2106 | 93.123 | 1198 | 118.136 | 1296
18.545 | 0 | 20.327 | 408 | 28.313 | 1070 | 33.329 | 1428 | 38.345 | 656 | 48.31 | 577597 | 50.29 | 868 | 68.308 | 1441 | 93.321 | 1664 | 118.334 | 1468
18.743 | 0 | 20.525 | 470 | 28.511 | 1174 | 33.527 | 1182 | 38.543 | 477 | 48.508 | 476889 | 50.488 | 633 | 68.506 | 1183 | 93.519 | 1719 | 118.532 | 1504
18.941 | 0 | 20.723 | 526 | 28.709 | 1343 | 33.725 | 1252 | 38.741 | 534 | 48.706 | 315005 | 50.686 | 867 | 68.704 | 1308 | 93.717 | 1679 | 118.73 | 1362
19.139 | 0 | 20.921 | 538 | 28.907 | 2264 | 33.923 | 1658 | 38.939 | 385 | 48.904 | 169990 | 50.884 | 639 | 68.902 | 1506 | 93.915 | 1824 | 118.928 | 1498
19.337 | 0 | 21.119 | 553 | 29.105 | 12217 | 34.121 | 2345 | 39.137 | 399 | 49.102 | 72476 | 51.082 | 497 | 69.1 | 1581 | 94.113 | 1733 | 119.126 | 1876
19.535 | 0 | 21.317 | 670 | 29.303 | 72920 | 34.319 | 3030 | 39.335 | 331 | 49.3 | 25753 | 51.28 | 491 | 69.298 | 1628 | 94.311 | 2103 | 119.324 | 2426
19.733 | 0 | 21.515 | 924 | 29.501 | 268852 | 34.517 | 4224 | 39.533 | 416 | 49.498 | 7975 | 51.478 | 683 | 69.496 | 1945 | 94.509 | 2139 | 119.522 | 4122
19.931 | 0 | 21.713 | 1398 | 29.699 | 631840 | 34.715 | 5385 | 39.731 | 506 | 49.696 | 2474 | 51.676 | 527 | 69.694 | 1894 | 94.707 | 2202 | 119.72 | 6733
20.129 | 0 | 21.911 | 1549 | 29.897 | 982333 | 34.913 | 6899 | 39.929 | 386 | 49.894 | 1359 | 51.874 | 556 | 69.892 | 2474 | 94.905 | 2161 | 119.918 | 11890
20.327 | 0 | 22.109 | 1476 | 30.095 | 1043505 | 35.111 | 8573 | 40.127 | 395 | 50.092 | 1038 | 52.072 | 626 | 70.09 | 2615 | 95.103 | 2172 | 120.116 | 20204
20.525 | 0 | 22.307 | 1814 | 30.293 | 761439 | 35.309 | 8560 | 40.325 | 345 | 50.29 | 922 | 52.27 | 635 | 70.288 | 2493 | 95.301 | 2046 | 120.314 | 32498
20.723 | 0 | 22.505 | 2056 | 30.491 | 366935 | 35.507 | 9250 | 40.523 | 319 | 50.488 | 725 | 52.468 | 516 | 70.486 | 2657 | 95.499 | 1722 | 120.512 | 47652
20.921 | 0 | 22.703 | 2433 | 30.689 | 116580 | 35.705 | 8636 | 40.721 | 353 | 50.686 | 647 | 52.666 | 560 | 70.684 | 2705 | 95.697 | 1859 | 120.71 | 65809
21.119 | 0 | 22.901 | 5677 | 30.887 | 24564 | 35.903 | 7454 | 40.919 | 464 | 50.884 | 647 | 52.864 | 584 | 70.882 | 2797 | 95.895 | 2385 | 120.908 | 81341
21.317 | 0 | 23.099 | 9817 | 31.085 | 3827 | 36.101 | 5560 | 41.117 | 382 | 51.082 | 366 | 53.062 | 744 | 71.08 | 2958 | 96.093 | 1884 | 121.106 | 96342
21.515 | 0 | 23.297 | 15901 | 31.283 | 1045 | 36.299 | 4681 | 41.315 | 401 | 51.28 | 520 | 53.26 | 698 | 71.278 | 3328 | 96.291 | 2047 | 121.304 | 102331
21.713 | 0 | 23.495 | 16558 | 31.481 | 727 | 36.497 | 3633 | 41.513 | 439 | 51.478 | 488 | 53.458 | 693 | 71.476 | 3101 | 96.489 | 1909 | 121.502 | 100899
21.911 | 0 | 23.693 | 13831 | 31.679 | 721 | 36.695 | 3128 | 41.711 | 409 | 51.676 | 467 | 53.656 | 655 | 71.674 | 3449 | 96.687 | 1574 | 121.7 | 91488
22.109 | 0 | 23.891 | 10865 | 31.877 | 692 | 36.893 | 2374 | 41.909 | 443 | 51.874 | 355 | 53.854 | 667 | 71.872 | 3393 | 96.885 | 1468 | 121.898 | 78997
22.307 | 0 | 24.089 | 6757 | 32.075 | 517 | 37.091 | 1805 | 42.107 | 499 | 52.072 | 392 | 54.052 | 569 | 72.07 | 3202 | 97.083 | 1699 | 122.096 | 60113
22.505 | 0 | 24.287 | 4262 | 32.273 | 446 | 37.289 | 1241 | 42.305 | 700 | 52.27 | 290 | 54.25 | 337 | 72.268 | 2618 | 97.281 | 1486 | 122.294 | 42988
22.703 | 0 | 24.485 | 1964 | 32.471 | 307 | 37.487 | 1313 | 42.503 | 657 | 52.468 | 433 | 54.448 | 271 | 72.466 | 2779 | 97.479 | 1279 | 122.492 | 28021
22.901 | 0 | 24.683 | 927 | 32.669 | 573 | 37.685 | 930 | 42.701 | 604 | 52.666 | 351 | 54.646 | 279 | 72.664 | 2592 | 97.677 | 1254 | 122.69 | 17814
23.099 | 0 | 24.881 | 543 | 32.867 | 384 | 37.883 | 1001 | 42.899 | 538 | 52.864 | 327 | 54.844 | 308 | 72.862 | 2216 | 97.875 | 1211 | 122.888 | 9712
23.297 | 0 | 25.079 | 368 | 33.065 | 361 | 38.081 | 864 | 43.097 | 516 | 53.062 | 254 | 55.042 | 257 | 73.06 | 1940 | 98.073 | 1174 | 123.086 | 5435
23.495 | 0 | 25.277 | 465 | 33.263 | 407 | 38.279 | 658 | 43.295 | 597 | 53.26 | 344 | 55.24 | 214 | 73.258 | 1770 | 98.271 | 942 | 123.284 | 2694
23.693 | 0 | 25.475 | 412 | 33.461 | 404 | 38.477 | 470 | 43.493 | 676 | 53.458 | 223 | 55.438 | 312 | 73.456 | 1237 | 98.469 | 989 | 123.482 | 1294
23.891 | 0 | 25.673 | 392 | 33.659 | 423 | 38.675 | 334 | 43.691 | 536 | 53.656 | 371 | 55.636 | 229 | 73.654 | 1242 | 98.667 | 866 | 123.68 | 885
24.089 | 0 | 25.871 | 543 | 33.857 | 455 | 38.873 | 310 | 43.889 | 511 | 53.854 | 347 | 55.834 | 257 | 73.852 | 766 | 98.865 | 549 | 123.878 | 428
24.287 | 0 | 26.069 | 404 | 34.055 | 381 | 39.071 | 323 | 44.087 | 444 | 54.052 | 256 | 56.032 | 204 | 74.05 | 541 | 99.063 | 369 | 124.076 | 224
24.485 | 0 | 26.267 | 404 | 34.253 | 220 | 39.269 | 251 | 44.285 | 460 | 54.25 | 253 | 56.23 | 220 | 74.248 | 227 | 99.261 | 162 | 124.274 | 199
24.683 | 0 | 26.465 | 314 | 34.451 | 347 | 39.467 | 283 | 44.483 | 449 | 54.448 | 312 | 56.428 | 230 | 74.446 | 213 | 99.459 | 165 | 124.472 | 127
24.881 | 0 | 26.663 | 415 | 34.649 | 219 | 39.665 | 387 | 44.681 | 459 | 54.646 | 321 | 56.626 | 247 | 74.644 | 277 | 99.657 | 143 | 124.67 | 206
25.079 | 0 | 26.861 | 490 | 34.847 | 265 | 39.863 | 341 | 44.879 | 554 | 54.844 | 351 | 56.824 | 252 | 74.842 | 269 | 99.855 | 134 | 124.868 | 100
25.277 | 0 | 27.059 | 397 | 35.045 | 251 | 40.061 | 265 | 45.077 | 510 | 55.042 | 300 | 57.022 | 206 | 75.04 | 232 | 100.053 | 179 | 125.066 | 98
25.475 | 0 | 27.257 | 581 | 35.243 | 173 | 40.259 | 307 | 45.275 | 451 | 55.24 | 363 | 57.22 | 227 | 75.238 | 217 | 100.251 | 187 | 125.264 | 120
25.673 | 0 | 27.455 | 652 | 35.441 | 173 | 40.457 | 296 | 45.473 | 653 | 55.438 | 135 | 57.418 | 185 | 75.436 | 296 | 100.449 | 162 | 125.462 | 142
25.871 | 0 | 27.653 | 746 | 35.639 | 196 | 40.655 | 257 | 45.671 | 538 | 55.636 | 195 | 57.616 | 168 | 75.634 | 247 | 100.647 | 234 | 125.66 | 153
26.069 | 0 | 27.851 | 966 | 35.837 | 198 | 40.853 | 261 | 45.869 | 532 | 55.834 | 203 | 57.814 | 200 | 75.832 | 219 | 100.845 | 230 | 125.858 | 182
26.267 | 0 | 28.049 | 1030 | 36.035 | 155 | 41.051 | 237 | 46.067 | 679 | 56.032 | 186 | 58.012 | 233 | 76.03 | 197 | 101.043 | 393 | 126.056 | 185
26.465 | 0 | 28.247 | 953 | 36.233 | 150 | 41.249 | 209 | 46.265 | 664 | 56.23 | 219 | 58.21 | 236 | 76.228 | 231 | 101.241 | 881 | 126.254 | 129
26.663 | 0 | 28.445 | 998 | 36.431 | 190 | 41.447 | 306 | 46.463 | 612 | 56.428 | 186 | 58.408 | 249 | 76.426 | 204 | 101.439 | 1922 | 126.452 | 121
26.861 | 0 | 28.643 | 916 | 36.629 | 229 | 41.645 | 246 | 46.661 | 792 | 56.626 | 92 | 58.606 | 290 | 76.624 | 239 | 101.637 | 4898 | 126.65 | 152
27.059 | 0 | 28.841 | 1108 | 36.827 | 236 | 41.843 | 360 | 46.859 | 648 | 56.824 | 171 | 58.804 | 341 | 76.822 | 263 | 101.835 | 10689 | 126.848 | 98
27.257 | 0 | 29.039 | 1482 | 37.025 | 279 | 42.041 | 355 | 47.056 | 1690 | 57.022 | 96 | 59.002 | 261 | 77.02 | 239 | 102.033 | 20879 | 127.046 | 113
27.455 | 0 | 29.237 | 5456 | 37.223 | 261 | 42.239 | 381 | 47.254 | 6089 | 57.22 | 119 | 59.2 | 223 | 77.218 | 267 | 102.231 | 37157 | 127.244 | 92
27.653 | 0 | 29.435 | 36945 | 37.421 | 351 | 42.437 | 446 | 47.452 | 24797 | 57.418 | 88 | 59.398 | 225 | 77.416 | 248 | 102.429 | 60947 | 127.442 | 123
27.851 | 0 | 29.633 | 170476 | 37.619 | 464 | 42.635 | 311 | 47.65 | 79692 | 57.616 | 48 | 59.596 | 118 | 77.614 | 282 | 102.627 | 89385 | 127.64 | 149
28.049 | 0 | 29.831 | 490135 | 37.817 | 493 | 42.833 | 366 | 47.848 | 195386 | 57.814 | 106 | 59.794 | 85 | 77.812 | 212 | 102.825 | 119880 | 127.838 | 109
28.247 | 0 | 30.029 | 907130 | 38.015 | 554 | 43.031 | 247 | 48.046 | 373794 | 58.012 | 93 | 59.992 | 62 | 78.01 | 231 | 103.023 | 145074 | 128.036 | 76
28.445 | 0 | 30.227 | 1136720 | 38.213 | 729 | 43.229 | 234 | 48.244 | 571222 | 58.21 | 102 | 60.19 | 28 | 78.208 | 157 | 103.221 | 155805 | 128.234 | 137
28.643 | 0 | 30.425 | 974805 | 38.411 | 558 | 43.427 | 209 | 48.442 | 689544 | 58.408 | 104 | 60.388 | 63 | 78.405 | 139 | 103.419 | 153298 | 128.432 | 119
28.841 | 0 | 30.623 | 565036 | 38.609 | 333 | 43.625 | 166 | 48.64 | 675085 | 58.606 | 66 | 60.586 | 47 | 78.603 | 147 | 103.617 | 139721 | 128.63 | 70
29.039 | 0 | 30.821 | 212357 | 38.807 | 240 | 43.823 | 109 | 48.838 | 530221 | 58.804 | 82 | 60.784 | 91 | 78.801 | 58 | 103.815 | 115228 | 128.828 | 70
29.237 | 0 | 31.019 | 53554 | 39.005 | 157 | 44.021 | 138 | 49.036 | 335143 | 59.002 | 81 | 60.982 | 91 | 78.999 | 79 | 104.013 | 85804 | 129.026 | 105
29.435 | 0 | 31.217 | 8938 | 39.203 | 92 | 44.219 | 152 | 49.234 | 171067 | 59.2 | 71 | 61.18 | 59 | 79.197 | 95 | 104.211 | 58938 | 129.224 | 94
29.633 | 0 | 31.415 | 1548 | 39.401 | 126 | 44.417 | 107 | 49.432 | 68721 | 59.398 | 66 | 61.378 | 70 | 79.395 | 52 | 104.409 | 35631 | 129.422 | 130
29.831 | 0 | 31.613 | 706 | 39.599 | 72 | 44.615 | 119 | 49.63 | 22149 | 59.596 | 55 | 61.576 | 49 | 79.593 | 63 | 104.607 | 19851 | 129.62 | 48
30.029 | 0 | 31.811 | 690 | 39.797 | 83 | 44.813 | 92 | 49.828 | 5625 | 59.794 | 93 | 61.774 | 80 | 79.791 | 101 | 104.805 | 10097 | 129.818 | 24
30.227 | 0 | 32.009 | 552 | 39.995 | 46 | 45.011 | 97 | 50.026 | 1327 | 59.992 | 81 | 61.972 | 86 | 79.989 | 89 | 105.003 | 3817 | 130.016 | 83
30.425 | 0 | 32.207 | 722 | 40.193 | 65 | 45.209 | 116 | 50.224 | 441 | 60.19 | 34 | 62.17 | 87 | 80.187 | 66 | 105.201 | 1456 | 130.214 | 53
30.623 | 0 | 32.405 | 679 | 40.391 | 56 | 45.407 | 113 | 50.422 | 302 | 60.388 | 46 | 62.368 | 40 | 80.385 | 113 | 105.399 | 504 | 130.412 | 52
30.821 | 0 | 32.603 | 588 | 40.589 | 76 | 45.605 | 82 | 50.62 | 184 | 60.586 | 60 | 62.566 | 47 | 80.583 | 106 | 105.597 | 166 | 130.61 | 91
31.019 | 0 | 32.801 | 608 | 40.787 | 78 | 45.803 | 177 | 50.818 | 308 | 60.784 | 42 | 62.764 | 50 | 80.781 | 134 | 105.795 | 94 | 130.808 | 94
31.217 | 0 | 32.999 | 475 | 40.985 | 120 | 46.001 | 116 | 51.016 | 318 | 60.982 | 61 | 62.962 | 61 | 80.979 | 121 | 105.993 | 46 | 131.006 | 74
31.415 | 0 | 33.197 | 463 | 41.183 | 103 | 46.199 | 157 | 51.214 | 246 | 61.18 | 85 | 63.16 | 61 | 81.177 | 129 | 106.191 | 88 | 131.204 | 87
31.613 | 0 | 33.395 | 416 | 41.381 | 161 | 46.397 | 142 | 51.412 | 259 | 61.378 | 43 | 63.358 | 133 | 81.375 | 136 | 106.389 | 39 | 131.402 | 67
31.811 | 0 | 33.593 | 410 | 41.579 | 162 | 46.595 | 163 | 51.61 | 178 | 61.576 | 127 | 63.556 | 233 | 81.573 | 107 | 106.587 | 51 | 131.6 | 67
32.009 | 0 | 33.791 | 516 | 41.777 | 253 | 46.793 | 96 | 51.808 | 90 | 61.774 | 113 | 63.754 | 232 | 81.771 | 132 | 106.785 | 42 | 131.798 | 69
32.207 | 0 | 33.989 | 366 | 41.975 | 337 | 46.99 | 171 | 52.006 | 78 | 61.972 | 120 | 63.952 | 228 | 81.969 | 131 | 106.983 | 63 | 131.996 | 38
32.405 | 0 | 34.187 | 378 | 42.173 | 403 | 47.188 | 64 | 52.204 | 42 | 62.17 | 90 | 64.15 | 259 | 82.167 | 75 | 107.181 | 97 | 132.194 | 77
32.603 | 0 | 34.385 | 366 | 42.371 | 417 | 47.386 | 84 | 52.402 | 22 | 62.368 | 85 | 64.348 | 273 | 82.365 | 56 | 107.379 | 22 | 132.392 | 67
32.801 | 0 | 34.583 | 294 | 42.569 | 387 | 47.584 | 67 | 52.6 | 19 | 62.566 | 35 | 64.546 | 289 | 82.563 | 81 | 107.577 | 63 | 132.59 | 76
32.999 | 0 | 34.781 | 245 | 42.767 | 500 | 47.782 | 47 | 52.798 | 7 | 62.764 | 64 | 64.744 | 274 | 82.761 | 79 | 107.775 | 30 | 132.788 | 64
33.197 | 0 | 34.979 | 343 | 42.965 | 373 | 47.98 | 27 | 52.996 | 21 | 62.962 | 40 | 64.942 | 144 | 82.959 | 45 | 107.973 | 57 | 132.986 | 48
33.395 | 0 | 35.177 | 273 | 43.163 | 331 | 48.178 | 46 | 53.194 | 17 | 63.16 | 45 | 65.14 | 118 | 83.157 | 58 | 108.171 | 39 | 133.184 | 60
33.593 | 0 | 35.375 | 218 | 43.361 | 266 | 48.376 | 15 | 53.392 | 14 | 63.358 | 108 | 65.338 | 99 | 83.355 | 68 | 108.369 | 29 | 133.382 | 49
33.791 | 0 | 35.573 | 282 | 43.559 | 232 | 48.574 | 30 | 53.59 | 10 | 63.556 | 101 | 65.536 | 81 | 83.553 | 40 | 108.567 | 46 | 133.58 | 70
33.989 | 0 | 35.771 | 189 | 43.757 | 166 | 48.772 | 31 | 53.788 | 2 | 63.754 | 167 | 65.734 | 36 | 83.751 | 50 | 108.765 | 27 | 133.778 | 64
34.187 | 0 | 35.969 | 257 | 43.955 | 183 | 48.97 | 28 | 53.986 | 24 | 63.952 | 342 | 65.932 | 58 | 83.949 | 29 | 108.963 | 66 | 133.976 | 93
34.385 | 0 | 36.167 | 220 | 44.153 | 106 | 49.168 | 35 | 54.184 | 12 | 64.15 | 334 | 66.13 | 29 | 84.147 | 33 | 109.161 | 40 | 134.174 | 59
34.583 | 0 | 36.365 | 163 | 44.351 | 184 | 49.366 | 21 | 54.382 | 0 | 64.348 | 400 | 66.328 | 31 | 84.345 | 22 | 109.359 | 61 | 134.372 | 58
34.781 | 0 | 36.563 | 224 | 44.549 | 85 | 49.564 | 27 | 54.58 | 5 | 64.546 | 342 | 66.526 | 21 | 84.543 | 35 | 109.556 | 40 | 134.57 | 41
34.979 | 0 | 36.761 | 190 | 44.747 | 142 | 49.762 | 23 | 54.778 | 6 | 64.744 | 209 | 66.724 | 32 | 84.741 | 49 | 109.754 | 27 | 134.768 | 22
35.177 | 0 | 36.959 | 169 | 44.945 | 91 | 49.96 | 16 | 54.976 | 30 | 64.942 | 297 | 66.922 | 43 | 84.939 | 28 | 109.952 | 43 | 134.966 | 50
35.375 | 0 | 37.157 | 215 | 45.143 | 122 | 50.158 | 22 | 55.174 | 14 | 65.14 | 112 | 67.12 | 29 | 85.137 | 24 | 110.15 | 33 | 135.164 | 36
35.573 | 0 | 37.355 | 323 | 45.341 | 101 | 50.356 | 14 | 55.372 | 6 | 65.338 | 92 | 67.318 | 36 | 85.335 | 35 | 110.348 | 9 | 135.362 | 43
35.771 | 0 | 37.553 | 387 | 45.539 | 57 | 50.554 | 40 | 55.57 | 0 | 65.536 | 53 | 67.516 | 12 | 85.533 | 46 | 110.546 | 64 | 135.56 | 53
35.969 | 0 | 37.751 | 489 | 45.737 | 71 | 50.752 | 7 | 55.768 | 0 | 65.734 | 59 | 67.714 | 19 | 85.731 | 6 | 110.744 | 45 | 135.758 | 24
36.167 | 0 | 37.949 | 550 | 45.935 | 118 | 50.95 | 10 | 55.966 | 13 | 65.932 | 51 | 67.912 | 61 | 85.929 | 62 | 110.942 | 71 | 135.956 | 21
36.365 | 0 | 38.147 | 634 | 46.133 | 144 | 51.148 | 11 | 56.164 | 0 | 66.13 | 21 | 68.11 | 25 | 86.127 | 39 | 111.14 | 43 | 136.154 | 32
36.563 | 0 | 38.345 | 771 | 46.331 | 143 | 51.346 | 17 | 56.362 | 0 | 66.328 | 31 | 68.308 | 34 | 86.325 | 22 | 111.338 | 28 | 136.352 | 37
36.761 | 0 | 38.543 | 558 | 46.529 | 191 | 51.544 | 18 | 56.56 | 4 | 66.526 | 36 | 68.506 | 47 | 86.523 | 49 | 111.536 | 18 | 136.55 | 39
36.959 | 0 | 38.741 | 500 | 46.727 | 177 | 51.742 | 3 | 56.758 | 5 | 66.724 | 50 | 68.704 | 36 | 86.721 | 31 | 111.734 | 24 | 136.748 | 38
37.157 | 0 | 38.939 | 294 | 46.924 | 146 | 51.94 | 16 | 56.956 | 0 | 66.922 | 27 | 68.902 | 16 | 86.919 | 34 | 111.932 | 29 | 136.946 | 30
37.355 | 0 | 39.137 | 210 | 47.122 | 183 | 52.138 | 6 | 57.154 | 2 | 67.12 | 27 | 69.1 | 19 | 87.117 | 6 | 112.13 | 21 | 137.144 | 26
37.553 | 0 | 39.335 | 141 | 47.32 | 96 | 52.336 | 7 | 57.352 | 8 | 67.318 | 19 | 69.298 | 9 | 87.315 | 5 | 112.328 | 13 | 137.342 | 20
37.751 | 0 | 39.533 | 74 | 47.518 | 92 | 52.534 | 8 | 57.55 | 0 | 67.516 | 17 | 69.496 | 33 | 87.513 | 45 | 112.526 | 25 | 137.54 | 7
37.949 | 0 | 39.731 | 111 | 47.716 | 64 | 52.732 | 29 | 57.748 | 8 | 67.714 | 29 | 69.694 | 48 | 87.711 | 17 | 112.724 | 8 | 137.738 | 29
38.147 | 0 | 39.929 | 144 | 47.914 | 47 | 52.93 | 16 | 57.946 | 0 | 67.912 | 32 | 69.892 | 15 | 87.909 | 25 | 112.922 | 18 | 137.936 | 13
38.345 | 0 | 40.127 | 91 | 48.112 | 50 | 53.128 | 6 | 58.144 | 0 | 68.11 | 11 | 70.09 | 22 | 88.107 | 13 | 113.12 | 37 | 138.134 | 4
38.543 | 0 | 40.325 | 153 | 48.31 | 52 | 53.326 | 6 | 58.342 | 0 | 68.308 | 18 | 70.288 | 58 | 88.305 | 4 | 113.318 | 3 | 138.332 | 21
38.741 | 0 | 40.523 | 93 | 48.508 | 50 | 53.524 | 0 | 58.54 | 0 | 68.506 | 16 | 70.486 | 15 | 88.503 | 16 | 113.516 | 3 | 138.53 | 21
38.939 | 0 | 40.721 | 92 | 48.706 | 26 | 53.722 | 21 | 58.738 | 0 | 68.704 | 17 | 70.684 | 36 | 88.701 | 12 | 113.714 | 0 | 138.728 | 15
39.137 | 0 | 40.919 | 149 | 48.904 | 28 | 53.92 | 5 | 58.936 | 5 | 68.902 | 34 | 70.882 | 49 | 88.899 | 22 | 113.912 | 5 | 138.926 | 24
39.335 | 0 | 41.117 | 108 | 49.102 | 9 | 54.118 | 14 | 59.134 | 6 | 69.1 | 50 | 71.08 | 67 | 89.097 | 5 | 114.11 | 20 | 139.124 | 20
39.533 | 0 | 41.315 | 111 | 49.3 | 66 | 54.316 | 18 | 59.332 | 0 | 69.298 | 29 | 71.278 | 40 | 89.295 | 16 | 114.308 | 20 | 139.322 | 18
39.731 | 0 | 41.513 | 207 | 49.498 | 11 | 54.514 | 0 | 59.53 | 0 | 69.496 | 22 | 71.476 | 109 | 89.493 | 0 | 114.506 | 42 | 139.52 | 15
39.929 | 0 | 41.711 | 221 | 49.696 | 51 | 54.712 | 8 | 59.728 | 0 | 69.694 | 47 | 71.674 | 42 | 89.691 | 10 | 114.704 | 21 | 139.718 | 14
40.127 | 0 | 41.909 | 251 | 49.894 | 26 | 54.91 | 8 | 59.926 | 0 | 69.892 | 20 | 71.872 | 96 | 89.889 | 15 | 114.902 | 26 | 139.916 | 5
40.325 | 0 | 42.107 | 401 | 50.092 | 91 | 55.108 | 3 | 60.124 | 0 | 70.09 | 34 | 72.07 | 100 | 90.087 | 0 | 115.1 | 10 | 140.114 | 6
40.523 | 0 | 42.305 | 464 | 50.29 | 34 | 55.306 | 2 | 60.322 | 9 | 70.288 | 42 | 72.268 | 124 | 90.285 | 0 | 115.298 | 12 | 140.312 | 17
40.721 | 0 | 42.503 | 517 | 50.488 | 16 | 55.504 | 0 | 60.52 | 0 | 70.486 | 15 | 72.466 | 147 | 90.483 | 2 | 115.496 | 45 | 140.51 | 16
40.919 | 0 | 42.701 | 528 | 50.686 | 5 | 55.702 | 13 | 60.718 | 0 | 70.684 | 34 | 72.664 | 207 | 90.681 | 6 | 115.694 | 59 | 140.708 | 20
41.117 | 0 | 42.899 | 384 | 50.884 | 27 | 55.9 | 0 | 60.916 | 0 | 70.882 | 31 | 72.862 | 244 | 90.879 | 0 | 115.892 | 37 | 140.906 | 9
41.315 | 0 | 43.097 | 434 | 51.082 | 24 | 56.098 | 2 | 61.114 | 11 | 71.08 | 11 | 73.06 | 176 | 91.077 | 6 | 116.09 | 28 | 141.103 | 19
41.513 | 0 | 43.295 | 455 | 51.28 | 14 | 56.296 | 19 | 61.312 | 0 | 71.278 | 61 | 73.258 | 200 | 91.275 | 7 | 116.288 | 41 | 141.301 | 23
41.711 | 0 | 43.493 | 225 | 51.478 | 35 | 56.494 | 19 | 61.51 | 0 | 71.476 | 24 | 73.456 | 131 | 91.473 | 5 | 116.486 | 39 | 141.499 | 10
41.909 | 0 | 43.691 | 305 | 51.676 | 26 | 56.692 | 0 | 61.708 | 12 | 71.674 | 18 | 73.654 | 127 | 91.671 | 0 | 116.684 | 54 | 141.697 | 21
42.107 | 0 | 43.889 | 174 | 51.874 | 18 | 56.89 | 5 | 61.906 | 0 | 71.872 | 31 | 73.852 | 162 | 91.869 | 14 | 116.882 | 23 | 141.895 | 12
42.305 | 0 | 44.087 | 205 | 52.072 | 14 | 57.088 | 0 | 62.104 | 5 | 72.07 | 21 | 74.05 | 178 | 92.067 | 8 | 117.08 | 12 | 142.093 | 18
42.503 | 0 | 44.285 | 154 | 52.27 | 20 | 57.286 | 2 | 62.302 | 0 | 72.268 | 29 | 74.248 | 83 | 92.265 | 13 | 117.278 | 31 | 142.291 | 10
42.701 | 0 | 44.483 | 134 | 52.468 | 14 | 57.484 | 10 | 62.5 | 0 | 72.466 | 19 | 74.446 | 65 | 92.463 | 12 | 117.476 | 36 | 142.489 | 16
42.899 | 0 | 44.681 | 161 | 52.666 | 15 | 57.682 | 0 | 62.698 | 5 | 72.664 | 20 | 74.644 | 41 | 92.661 | 14 | 117.674 | 38 | 142.687 | 9
43.097 | 0 | 44.879 | 145 | 52.864 | 0 | 57.88 | 9 | 62.896 | 0 | 72.862 | 0 | 74.842 | 79 | 92.859 | 0 | 117.872 | 28 | 142.885 | 18
43.295 | 0 | 45.077 | 127 | 53.062 | 18 | 58.078 | 0 | 63.094 | 0 | 73.06 | 5 | 75.04 | 50 | 93.057 | 11 | 118.07 | 12 | 143.083 | 11
43.493 | 0 | 45.275 | 179 | 53.26 | 31 | 58.276 | 11 | 63.292 | 4 | 73.258 | 0 | 75.238 | 38 | 93.255 | 4 | 118.268 | 10 | 143.281 | 14
43.691 | 0 | 45.473 | 116 | 53.458 | 18 | 58.474 | 18 | 63.49 | 0 | 73.456 | 0 | 75.436 | 24 | 93.453 | 9 | 118.466 | 10 | 143.479 | 11
43.889 | 0 | 45.671 | 118 | 53.656 | 18 | 58.672 | 0 | 63.688 | 6 | 73.654 | 0 | 75.634 | 14 | 93.651 | 4 | 118.664 | 0 | 143.677 | 17
44.087 | 0 | 45.869 | 106 | 53.854 | 26 | 58.87 | 0 | 63.886 | 12 | 73.852 | 0 | 75.832 | 65 | 93.849 | 0 | 118.862 | 11 | 143.875 | 4
44.285 | 0 | 46.067 | 185 | 54.052 | 6 | 59.068 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 25 | 94.047 | 0 | 119.06 | 7 | 144.073 | 0
44.483 | 0 | 46.265 | 139 | 54.25 | 7 | 59.266 | 0 | 64.282 | 0 | 74.248 | 0 | 76.228 | 13 | 94.245 | 0 | 119.258 | 9 | 144.271 | 8
44.681 | 0 | 46.463 | 183 | 54.448 | 10 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 3 | 94.443 | 0 | 119.456 | 0 | 144.469 | 7
44.879 | 0 | 46.661 | 180 | 54.646 | 11 | 59.662 | 4 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 15 | 119.654 | 8 | 144.667 | 8
45.077 | 0 | 46.859 | 152 | 54.844 | 3 | 59.86 | 3 | 64.876 | 0 | 74.842 | 0 | 76.822 | 7 | 94.839 | 6 | 119.852 | 7 | 144.865 | 9
45.275 | 0 | 47.056 | 199 | 55.042 | 6 | 60.058 | 0 | 65.074 | 4 | 75.04 | 0 | 77.02 | 14 | 95.037 | 0 | 120.05 | 0 | 145.063 | 12
45.473 | 0 | 47.254 | 239 | 55.24 | 0 | 60.256 | 12 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 8 | 120.248 | 4 | 145.261 | 10
45.671 | 0 | 47.452 | 142 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 3 | 120.446 | 0 | 145.459 | 4
45.869 | 0 | 47.65 | 157 | 55.636 | 13 | 60.652 | 2 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 7 | 120.644 | 19 | 145.657 | 0
46.067 | 0 | 47.848 | 83 | 55.834 | 11 | 60.85 | 8 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 7 | 120.842 | 15 | 145.855 | 0
46.265 | 0 | 48.046 | 68 | 56.032 | 0 | 61.048 | 0 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 7 | 121.04 | 0 | 146.053 | 0
46.463 | 0 | 48.244 | 83 | 56.23 | 4 | 61.246 | 0 | 66.262 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 6 | 121.238 | 17 | 146.251 | 0
46.661 | 0 | 48.442 | 30 | 56.428 | 17 | 61.444 | 0 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 0 | 146.449 | 3
46.859 | 0 | 48.64 | 48 | 56.626 | 12 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 4 | 146.647 | 14
47.056 | 0 | 48.838 | 52 | 56.824 | 13 | 61.84 | 0 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 6 | 146.845 | 0
47.254 | 0 | 49.036 | 35 | 57.022 | 5 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 4 | 122.03 | 9 | 147.043 | 3
47.452 | 0 | 49.234 | 38 | 57.22 | 14 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 0 | 147.241 | 6
47.65 | 0 | 49.432 | 40 | 57.418 | 0 | 62.434 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 3 | 122.426 | 11 | 147.439 | 0
47.848 | 0 | 49.63 | 40 | 57.616 | 0 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 26 | 122.624 | 7 | 147.637 | 0
48.046 | 0 | 49.828 | 16 | 57.814 | 20 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 7 | 122.822 | 0 | 147.835 | 0
48.244 | 0 | 50.026 | 36 | 58.012 | 32 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 0 | 148.033 | 0
48.442 | 0 | 50.224 | 27 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 5 | 123.218 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 15 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 5 | 148.429 | 0
48.838 | 0 | 50.62 | 35 | 58.606 | 4 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 4 | 123.614 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 35 | 58.804 | 0 | 63.82 | 7 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 8 | 123.812 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 16 | 59.002 | 5 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 7 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 53 | 59.2 | 6 | 64.216 | 12 | 69.232 | 5 | 79.197 | 0 | 81.177 | 0 | 99.195 | 4 | 124.208 | 9 | 149.221 | 0
|  | 51.412 | 18 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 0 | 149.419 | 0

## Analyte Polyalanine n8
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 77 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 86.985 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 87.183 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 87.381 | 0 | 110.414 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 87.579 | 0 | 110.612 | 0
11.022 | 3695 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 87.777 | 0 | 110.81 | 0
11.22 | 58476 | 13.002 | 0 | 20.987 | 232 | 26.003 | 58 | 31.019 | 225 | 40.985 | 110 | 42.965 | 17 | 60.982 | 96 | 87.975 | 45 | 111.008 | 97
11.418 | 155330 | 13.2 | 0 | 21.185 | 5077 | 26.201 | 1822 | 31.217 | 2902 | 41.183 | 2183 | 43.163 | 2118 | 61.18 | 3678 | 88.173 | 2971 | 111.206 | 2940
11.616 | 203141 | 13.398 | 0 | 21.383 | 10911 | 26.399 | 9319 | 31.415 | 6014 | 41.381 | 6847 | 43.361 | 5826 | 61.378 | 7941 | 88.371 | 6834 | 111.404 | 5313
11.814 | 356883 | 13.596 | 0 | 21.581 | 20660 | 26.597 | 12530 | 31.613 | 9399 | 41.579 | 11886 | 43.559 | 9531 | 61.576 | 10487 | 88.569 | 9809 | 111.602 | 6575
12.012 | 597790 | 13.794 | 0 | 21.779 | 21106 | 26.795 | 12252 | 31.811 | 11342 | 41.777 | 11273 | 43.757 | 8694 | 61.774 | 10395 | 88.767 | 9164 | 111.8 | 6006
12.21 | 266051 | 13.992 | 0 | 21.977 | 19935 | 26.993 | 11450 | 32.009 | 12375 | 41.975 | 11497 | 43.955 | 7805 | 61.972 | 10507 | 88.965 | 8377 | 111.998 | 6101
12.408 | 28258 | 14.19 | 0 | 22.175 | 18022 | 27.191 | 10152 | 32.207 | 11763 | 42.173 | 12937 | 44.153 | 8293 | 62.17 | 11854 | 89.163 | 7980 | 112.196 | 6260
12.606 | 1628 | 14.388 | 0 | 22.373 | 18708 | 27.389 | 8672 | 32.405 | 11616 | 42.371 | 12580 | 44.351 | 8597 | 62.368 | 12294 | 89.361 | 8224 | 112.394 | 6752
12.804 | 63 | 14.586 | 0 | 22.571 | 20093 | 27.587 | 7639 | 32.603 | 9943 | 42.569 | 13513 | 44.549 | 8881 | 62.566 | 14323 | 89.559 | 7948 | 112.592 | 6755
13.002 | 17 | 14.784 | 0 | 22.769 | 21350 | 27.785 | 6639 | 32.801 | 9369 | 42.767 | 11779 | 44.747 | 9010 | 62.764 | 16935 | 89.757 | 7625 | 112.79 | 6579
13.2 | 0 | 14.982 | 0 | 22.967 | 22713 | 27.983 | 7420 | 32.999 | 9013 | 42.965 | 11074 | 44.945 | 10117 | 62.962 | 20283 | 89.955 | 7879 | 112.988 | 6459
13.398 | 0 | 15.18 | 0 | 23.165 | 22153 | 28.181 | 10825 | 33.197 | 8827 | 43.163 | 9994 | 45.143 | 11820 | 63.16 | 26485 | 90.153 | 7508 | 113.186 | 6905
13.596 | 0 | 15.378 | 0 | 23.363 | 22298 | 28.379 | 24881 | 33.395 | 9232 | 43.361 | 9350 | 45.341 | 14388 | 63.358 | 33031 | 90.351 | 7236 | 113.384 | 6980
13.794 | 0 | 15.576 | 0 | 23.561 | 24442 | 28.577 | 60022 | 33.593 | 10134 | 43.559 | 8376 | 45.539 | 19878 | 63.556 | 39772 | 90.549 | 7539 | 113.582 | 7654
13.992 | 0 | 15.773 | 0 | 23.759 | 28578 | 28.775 | 127822 | 33.791 | 10600 | 43.757 | 7994 | 45.737 | 29216 | 63.754 | 47621 | 90.747 | 6899 | 113.78 | 7260
14.19 | 0 | 15.971 | 0 | 23.957 | 28212 | 28.973 | 214046 | 33.989 | 11800 | 43.955 | 7900 | 45.935 | 43503 | 63.952 | 56073 | 90.945 | 6755 | 113.978 | 7349
14.388 | 0 | 16.169 | 9 | 24.155 | 26530 | 29.171 | 272059 | 34.187 | 11986 | 44.153 | 8505 | 46.133 | 62710 | 64.15 | 62863 | 91.143 | 6473 | 114.176 | 7950
14.586 | 0 | 16.367 | 125 | 24.353 | 25105 | 29.369 | 263013 | 34.385 | 13146 | 44.351 | 8633 | 46.331 | 85113 | 64.348 | 69766 | 91.341 | 6475 | 114.374 | 8546
14.784 | 0 | 16.565 | 103 | 24.551 | 22747 | 29.567 | 204698 | 34.583 | 13906 | 44.549 | 9169 | 46.529 | 107329 | 64.546 | 73028 | 91.539 | 6649 | 114.572 | 9014
14.982 | 0 | 16.763 | 28 | 24.749 | 20157 | 29.765 | 134954 | 34.781 | 14195 | 44.747 | 9253 | 46.727 | 125725 | 64.744 | 73623 | 91.737 | 6651 | 114.77 | 9651
15.18 | 0 | 16.961 | 133 | 24.947 | 17810 | 29.963 | 84802 | 34.979 | 13130 | 44.945 | 10699 | 46.924 | 136028 | 64.942 | 71190 | 91.935 | 5642 | 114.968 | 10426
15.378 | 0 | 17.159 | 293 | 25.145 | 16707 | 30.161 | 55478 | 35.177 | 12882 | 45.143 | 11062 | 47.122 | 133105 | 65.14 | 65017 | 92.133 | 6144 | 115.166 | 11838
15.576 | 0 | 17.357 | 196 | 25.343 | 15984 | 30.359 | 40270 | 35.375 | 12306 | 45.341 | 13524 | 47.32 | 118087 | 65.338 | 57643 | 92.331 | 6437 | 115.364 | 12978
15.773 | 0 | 17.555 | 343 | 25.541 | 15481 | 30.557 | 28953 | 35.573 | 12900 | 45.539 | 18381 | 47.518 | 95979 | 65.536 | 48336 | 92.529 | 6148 | 115.562 | 13418
15.971 | 0 | 17.753 | 1586 | 25.739 | 14582 | 30.755 | 19966 | 35.771 | 13581 | 45.737 | 27266 | 47.716 | 75886 | 65.734 | 39540 | 92.727 | 6402 | 115.76 | 14809
16.169 | 0 | 17.951 | 2500 | 25.937 | 12629 | 30.953 | 12412 | 35.969 | 14357 | 45.935 | 40882 | 47.914 | 57278 | 65.932 | 32817 | 92.925 | 6861 | 115.958 | 15240
16.367 | 0 | 18.149 | 1673 | 26.135 | 11181 | 31.151 | 8607 | 36.167 | 15300 | 46.133 | 59782 | 48.112 | 44130 | 66.13 | 27402 | 93.123 | 6278 | 116.156 | 15262
16.565 | 0 | 18.347 | 1324 | 26.333 | 10460 | 31.349 | 7439 | 36.365 | 14866 | 46.331 | 83834 | 48.31 | 35422 | 66.328 | 22751 | 93.321 | 6681 | 116.354 | 16345
16.763 | 0 | 18.545 | 2257 | 26.531 | 9743 | 31.547 | 7774 | 36.563 | 15663 | 46.529 | 108618 | 48.508 | 30953 | 66.526 | 20481 | 93.519 | 6979 | 116.552 | 17374
16.961 | 0 | 18.743 | 2546 | 26.729 | 9400 | 31.745 | 9215 | 36.761 | 16589 | 46.727 | 130854 | 48.706 | 27414 | 66.724 | 18690 | 93.717 | 6186 | 116.75 | 18248
17.159 | 0 | 18.941 | 1903 | 26.927 | 11610 | 31.943 | 10864 | 36.959 | 18043 | 46.924 | 143918 | 48.904 | 25341 | 66.922 | 17620 | 93.915 | 6981 | 116.948 | 19238
17.357 | 0 | 19.139 | 2751 | 27.125 | 11556 | 32.141 | 11623 | 37.157 | 17281 | 47.122 | 144973 | 49.102 | 24581 | 67.12 | 15694 | 94.113 | 6571 | 117.146 | 20013
17.555 | 0 | 19.337 | 5944 | 27.323 | 11627 | 32.339 | 12672 | 37.355 | 17375 | 47.32 | 131310 | 49.3 | 22945 | 67.318 | 14733 | 94.311 | 6508 | 117.344 | 20341
17.753 | 0 | 19.535 | 7812 | 27.521 | 9475 | 32.537 | 12221 | 37.553 | 17171 | 47.518 | 111827 | 49.498 | 20458 | 67.516 | 14938 | 94.509 | 7058 | 117.542 | 21737
17.951 | 0 | 19.733 | 6064 | 27.719 | 7734 | 32.735 | 10029 | 37.751 | 16050 | 47.716 | 87041 | 49.696 | 18676 | 67.714 | 14218 | 94.707 | 7398 | 117.74 | 21314
18.149 | 0 | 19.931 | 6385 | 27.917 | 6987 | 32.933 | 9264 | 37.949 | 14340 | 47.914 | 65259 | 49.894 | 17507 | 67.912 | 13846 | 94.905 | 7251 | 117.938 | 21524
18.347 | 0 | 20.129 | 10028 | 28.115 | 7277 | 33.131 | 8601 | 38.147 | 12871 | 48.112 | 50010 | 50.092 | 16007 | 68.11 | 13168 | 95.103 | 7212 | 118.136 | 21613
18.545 | 0 | 20.327 | 12880 | 28.313 | 11275 | 33.329 | 9045 | 38.345 | 10827 | 48.31 | 38149 | 50.29 | 14595 | 68.308 | 12070 | 95.301 | 7393 | 118.334 | 21525
18.743 | 0 | 20.525 | 11360 | 28.511 | 25293 | 33.527 | 9480 | 38.543 | 9883 | 48.508 | 32557 | 50.488 | 13434 | 68.506 | 11792 | 95.499 | 7762 | 118.532 | 21200
18.941 | 0 | 20.723 | 9656 | 28.709 | 62186 | 33.725 | 10519 | 38.741 | 9622 | 48.706 | 30284 | 50.686 | 12942 | 68.704 | 10867 | 95.697 | 7977 | 118.73 | 20423
19.139 | 0 | 20.921 | 12653 | 28.907 | 138071 | 33.923 | 11296 | 38.939 | 9404 | 48.904 | 27193 | 50.884 | 11499 | 68.902 | 9793 | 95.895 | 7714 | 118.928 | 19432
19.337 | 0 | 21.119 | 16714 | 29.105 | 230383 | 34.121 | 11095 | 39.137 | 10170 | 49.102 | 24742 | 51.082 | 11262 | 69.1 | 9016 | 96.093 | 8063 | 119.126 | 16984
19.535 | 0 | 21.317 | 17772 | 29.303 | 290009 | 34.319 | 11708 | 39.335 | 10170 | 49.3 | 23834 | 51.28 | 10205 | 69.298 | 8273 | 96.291 | 7946 | 119.324 | 16918
19.733 | 0 | 21.515 | 16214 | 29.501 | 285960 | 34.517 | 11260 | 39.533 | 11101 | 49.498 | 21072 | 51.478 | 9639 | 69.496 | 8195 | 96.489 | 7781 | 119.522 | 14801
19.931 | 0 | 21.713 | 17576 | 29.699 | 221352 | 34.715 | 10674 | 39.731 | 11762 | 49.696 | 19627 | 51.676 | 8762 | 69.694 | 7520 | 96.687 | 8091 | 119.72 | 13594
20.129 | 0 | 21.911 | 19967 | 29.897 | 144816 | 34.913 | 10821 | 39.929 | 12090 | 49.894 | 17764 | 51.874 | 9101 | 69.892 | 7010 | 96.885 | 8590 | 119.918 | 12755
20.327 | 0 | 22.109 | 20336 | 30.095 | 91673 | 35.111 | 10804 | 40.127 | 12337 | 50.092 | 14599 | 52.072 | 9070 | 70.09 | 7384 | 97.083 | 8406 | 120.116 | 10996
20.525 | 0 | 22.307 | 19145 | 30.293 | 58056 | 35.309 | 10487 | 40.325 | 12239 | 50.29 | 13174 | 52.27 | 9447 | 70.288 | 7359 | 97.281 | 8842 | 120.314 | 10230
20.723 | 0 | 22.505 | 18150 | 30.491 | 41300 | 35.507 | 10919 | 40.523 | 11068 | 50.488 | 11805 | 52.468 | 10040 | 70.486 | 7144 | 97.479 | 9733 | 120.512 | 8857
20.921 | 0 | 22.703 | 19420 | 30.689 | 30171 | 35.705 | 11943 | 40.721 | 10238 | 50.686 | 11596 | 52.666 | 10235 | 70.684 | 7305 | 97.677 | 10396 | 120.71 | 8687
21.119 | 0 | 22.901 | 22045 | 30.887 | 20369 | 35.903 | 12926 | 40.919 | 9466 | 50.884 | 10992 | 52.864 | 9926 | 70.882 | 6606 | 97.875 | 11924 | 120.908 | 8498
21.317 | 0 | 23.099 | 22838 | 31.085 | 12740 | 36.101 | 13452 | 41.117 | 8416 | 51.082 | 10430 | 53.062 | 10630 | 71.08 | 7082 | 98.073 | 13378 | 121.106 | 8129
21.515 | 0 | 23.297 | 22664 | 31.283 | 6902 | 36.299 | 13470 | 41.315 | 7352 | 51.28 | 9231 | 53.26 | 10073 | 71.278 | 6854 | 98.271 | 14681 | 121.304 | 8033
21.713 | 0 | 23.495 | 21821 | 31.481 | 3207 | 36.497 | 14026 | 41.513 | 6139 | 51.478 | 7731 | 53.458 | 9994 | 71.476 | 6218 | 98.469 | 16711 | 121.502 | 7720
21.911 | 0 | 23.693 | 24340 | 31.679 | 1634 | 36.695 | 13691 | 41.711 | 4601 | 51.676 | 6816 | 53.656 | 9109 | 71.674 | 5731 | 98.667 | 18532 | 121.7 | 7000
22.109 | 0 | 23.891 | 26818 | 31.877 | 1185 | 36.893 | 13323 | 41.909 | 4393 | 51.874 | 6244 | 53.854 | 8112 | 71.872 | 5595 | 98.865 | 21267 | 121.898 | 6646
22.307 | 0 | 24.089 | 28451 | 32.075 | 1030 | 37.091 | 13257 | 42.107 | 3892 | 52.072 | 6040 | 54.052 | 7019 | 72.07 | 5424 | 99.063 | 23002 | 122.096 | 6819
22.505 | 0 | 24.287 | 27957 | 32.273 | 795 | 37.289 | 13280 | 42.305 | 4141 | 52.27 | 5652 | 54.25 | 5379 | 72.268 | 5926 | 99.261 | 24758 | 122.294 | 6788
22.703 | 0 | 24.485 | 25753 | 32.471 | 808 | 37.487 | 12413 | 42.503 | 4085 | 52.468 | 5443 | 54.448 | 4446 | 72.466 | 6281 | 99.459 | 26542 | 122.492 | 6247
22.901 | 0 | 24.683 | 23724 | 32.669 | 822 | 37.685 | 11710 | 42.701 | 3821 | 52.666 | 5775 | 54.646 | 3944 | 72.664 | 6231 | 99.657 | 26726 | 122.69 | 5532
23.099 | 0 | 24.881 | 20416 | 32.867 | 716 | 37.883 | 11030 | 42.899 | 3966 | 52.864 | 5299 | 54.844 | 3749 | 72.862 | 6259 | 99.855 | 28659 | 122.888 | 5276
23.297 | 0 | 25.079 | 18688 | 33.065 | 714 | 38.081 | 10386 | 43.097 | 3587 | 53.062 | 5095 | 55.042 | 4101 | 73.06 | 6505 | 100.053 | 29703 | 123.086 | 4989
23.495 | 0 | 25.277 | 16819 | 33.263 | 667 | 38.279 | 9236 | 43.295 | 3265 | 53.26 | 5499 | 55.24 | 3718 | 73.258 | 6252 | 100.251 | 30121 | 123.284 | 4586
23.693 | 0 | 25.475 | 15596 | 33.461 | 591 | 38.477 | 8759 | 43.493 | 3005 | 53.458 | 5189 | 55.438 | 3972 | 73.456 | 6229 | 100.449 | 30433 | 123.482 | 4837
23.891 | 0 | 25.673 | 14892 | 33.659 | 763 | 38.675 | 7867 | 43.691 | 3063 | 53.656 | 5015 | 55.636 | 3598 | 73.654 | 5584 | 100.647 | 30015 | 123.68 | 4627
24.089 | 0 | 25.871 | 13231 | 33.857 | 658 | 38.873 | 8114 | 43.889 | 2911 | 53.854 | 4662 | 55.834 | 3749 | 73.852 | 4618 | 100.845 | 29143 | 123.878 | 3959
24.287 | 0 | 26.069 | 11835 | 34.055 | 750 | 39.071 | 8001 | 44.087 | 2872 | 54.052 | 4516 | 56.032 | 3868 | 74.05 | 3119 | 101.043 | 27439 | 124.076 | 3212
24.485 | 0 | 26.267 | 11030 | 34.253 | 550 | 39.269 | 7438 | 44.285 | 2898 | 54.25 | 4604 | 56.23 | 3192 | 74.248 | 2143 | 101.241 | 24333 | 124.274 | 2940
24.683 | 0 | 26.465 | 10498 | 34.451 | 511 | 39.467 | 7144 | 44.483 | 2782 | 54.448 | 4168 | 56.428 | 3419 | 74.446 | 1799 | 101.439 | 21365 | 124.472 | 2885
24.881 | 0 | 26.663 | 9574 | 34.649 | 553 | 39.665 | 7118 | 44.681 | 2740 | 54.646 | 4188 | 56.626 | 3328 | 74.644 | 1758 | 101.637 | 18700 | 124.67 | 2735
25.079 | 0 | 26.861 | 8609 | 34.847 | 460 | 39.863 | 6557 | 44.879 | 2478 | 54.844 | 4157 | 56.824 | 2974 | 74.842 | 1760 | 101.835 | 15440 | 124.868 | 2571
25.277 | 0 | 27.059 | 8252 | 35.045 | 428 | 40.061 | 6401 | 45.077 | 2637 | 55.042 | 3716 | 57.022 | 3080 | 75.04 | 1829 | 102.033 | 13278 | 125.066 | 2328
25.475 | 0 | 27.257 | 8322 | 35.243 | 411 | 40.259 | 6089 | 45.275 | 2553 | 55.24 | 3577 | 57.22 | 2809 | 75.238 | 1524 | 102.231 | 11486 | 125.264 | 2213
25.673 | 0 | 27.455 | 8021 | 35.441 | 230 | 40.457 | 5952 | 45.473 | 3089 | 55.438 | 3770 | 57.418 | 3097 | 75.436 | 1555 | 102.429 | 10012 | 125.462 | 2128
25.871 | 0 | 27.653 | 7069 | 35.639 | 233 | 40.655 | 5177 | 45.671 | 5288 | 55.636 | 3574 | 57.616 | 2610 | 75.634 | 1585 | 102.627 | 8917 | 125.66 | 1968
26.069 | 0 | 27.851 | 6584 | 35.837 | 153 | 40.853 | 5109 | 45.869 | 9947 | 55.834 | 3439 | 57.814 | 2610 | 75.832 | 1638 | 102.825 | 7252 | 125.858 | 1441
26.267 | 0 | 28.049 | 6130 | 36.035 | 168 | 41.051 | 4585 | 46.067 | 21189 | 56.032 | 3538 | 58.012 | 2763 | 76.03 | 1613 | 103.023 | 6678 | 126.056 | 1252
26.465 | 0 | 28.247 | 5568 | 36.233 | 112 | 41.249 | 3758 | 46.265 | 39589 | 56.23 | 2991 | 58.21 | 2672 | 76.228 | 1449 | 103.221 | 5861 | 126.254 | 1176
26.663 | 0 | 28.445 | 6850 | 36.431 | 129 | 41.447 | 2653 | 46.463 | 64598 | 56.428 | 2193 | 58.408 | 2679 | 76.426 | 1228 | 103.419 | 5267 | 126.452 | 909
26.861 | 0 | 28.643 | 13501 | 36.629 | 122 | 41.645 | 1578 | 46.661 | 93892 | 56.626 | 1512 | 58.606 | 2609 | 76.624 | 1530 | 103.617 | 5058 | 126.65 | 824
27.059 | 0 | 28.841 | 38157 | 36.827 | 113 | 41.843 | 704 | 46.859 | 122400 | 56.824 | 857 | 58.804 | 2472 | 76.822 | 1585 | 103.815 | 4388 | 126.848 | 833
27.257 | 0 | 29.039 | 97680 | 37.025 | 96 | 42.041 | 306 | 47.056 | 147982 | 57.022 | 492 | 59.002 | 2188 | 77.02 | 1490 | 104.013 | 4009 | 127.046 | 858
27.455 | 0 | 29.237 | 190241 | 37.223 | 115 | 42.239 | 110 | 47.254 | 159640 | 57.22 | 318 | 59.2 | 2118 | 77.218 | 1549 | 104.211 | 3598 | 127.244 | 815
27.653 | 0 | 29.435 | 280166 | 37.421 | 102 | 42.437 | 38 | 47.452 | 157750 | 57.418 | 277 | 59.398 | 1663 | 77.416 | 1404 | 104.409 | 3284 | 127.442 | 629
27.851 | 0 | 29.633 | 307845 | 37.619 | 108 | 42.635 | 36 | 47.65 | 139404 | 57.616 | 311 | 59.596 | 1190 | 77.614 | 1309 | 104.607 | 2835 | 127.64 | 709
28.049 | 0 | 29.831 | 261937 | 37.817 | 134 | 42.833 | 73 | 47.848 | 111283 | 57.814 | 324 | 59.794 | 548 | 77.812 | 1148 | 104.805 | 2718 | 127.838 | 601
28.247 | 0 | 30.029 | 182374 | 38.015 | 103 | 43.031 | 67 | 48.046 | 83953 | 58.012 | 230 | 59.992 | 278 | 78.01 | 1305 | 105.003 | 2462 | 128.036 | 488
28.445 | 0 | 30.227 | 114683 | 38.213 | 67 | 43.229 | 47 | 48.244 | 61377 | 58.21 | 194 | 60.19 | 136 | 78.208 | 971 | 105.201 | 2257 | 128.234 | 439
28.643 | 0 | 30.425 | 67160 | 38.411 | 67 | 43.427 | 56 | 48.442 | 42676 | 58.408 | 154 | 60.388 | 123 | 78.405 | 494 | 105.399 | 2406 | 128.432 | 438
28.841 | 0 | 30.623 | 41979 | 38.609 | 91 | 43.625 | 45 | 48.64 | 31586 | 58.606 | 164 | 60.586 | 114 | 78.603 | 340 | 105.597 | 2237 | 128.63 | 316
29.039 | 0 | 30.821 | 26250 | 38.807 | 54 | 43.823 | 28 | 48.838 | 24157 | 58.804 | 130 | 60.784 | 104 | 78.801 | 267 | 105.795 | 2161 | 128.828 | 225
29.237 | 0 | 31.019 | 15399 | 39.005 | 66 | 44.021 | 49 | 49.036 | 19236 | 59.002 | 111 | 60.982 | 95 | 78.999 | 147 | 105.993 | 1771 | 129.026 | 237
29.435 | 0 | 31.217 | 7808 | 39.203 | 69 | 44.219 | 36 | 49.234 | 16086 | 59.2 | 152 | 61.18 | 54 | 79.197 | 125 | 106.191 | 1798 | 129.224 | 247
29.633 | 0 | 31.415 | 3589 | 39.401 | 79 | 44.417 | 29 | 49.432 | 13097 | 59.398 | 98 | 61.378 | 128 | 79.395 | 231 | 106.389 | 1531 | 129.422 | 183
29.831 | 0 | 31.613 | 1996 | 39.599 | 15 | 44.615 | 42 | 49.63 | 10908 | 59.596 | 63 | 61.576 | 61 | 79.593 | 175 | 106.587 | 1341 | 129.62 | 113
30.029 | 0 | 31.811 | 1382 | 39.797 | 25 | 44.813 | 11 | 49.828 | 8631 | 59.794 | 63 | 61.774 | 51 | 79.791 | 114 | 106.785 | 1222 | 129.818 | 126
30.227 | 0 | 32.009 | 1138 | 39.995 | 58 | 45.011 | 13 | 50.026 | 5558 | 59.992 | 76 | 61.972 | 36 | 79.989 | 126 | 106.983 | 987 | 130.016 | 189
30.425 | 0 | 32.207 | 1000 | 40.193 | 15 | 45.209 | 5 | 50.224 | 4198 | 60.19 | 55 | 62.17 | 39 | 80.187 | 100 | 107.181 | 969 | 130.214 | 131
30.623 | 0 | 32.405 | 947 | 40.391 | 22 | 45.407 | 9 | 50.422 | 2383 | 60.388 | 78 | 62.368 | 35 | 80.385 | 88 | 107.379 | 651 | 130.412 | 106
30.821 | 0 | 32.603 | 928 | 40.589 | 16 | 45.605 | 3 | 50.62 | 1459 | 60.586 | 60 | 62.566 | 45 | 80.583 | 65 | 107.577 | 500 | 130.61 | 124
31.019 | 0 | 32.801 | 832 | 40.787 | 14 | 45.803 | 10 | 50.818 | 790 | 60.784 | 44 | 62.764 | 47 | 80.781 | 58 | 107.775 | 361 | 130.808 | 116
31.217 | 0 | 32.999 | 897 | 40.985 | 44 | 46.001 | 0 | 51.016 | 731 | 60.982 | 43 | 62.962 | 16 | 80.979 | 122 | 107.973 | 144 | 131.006 | 100
31.415 | 0 | 33.197 | 890 | 41.183 | 19 | 46.199 | 22 | 51.214 | 532 | 61.18 | 70 | 63.16 | 45 | 81.177 | 35 | 108.171 | 97 | 131.204 | 152
31.613 | 0 | 33.395 | 739 | 41.381 | 20 | 46.397 | 21 | 51.412 | 399 | 61.378 | 34 | 63.358 | 49 | 81.375 | 90 | 108.369 | 49 | 131.402 | 111
31.811 | 0 | 33.593 | 710 | 41.579 | 14 | 46.595 | 22 | 51.61 | 337 | 61.576 | 29 | 63.556 | 31 | 81.573 | 90 | 108.567 | 11 | 131.6 | 102
32.009 | 0 | 33.791 | 794 | 41.777 | 0 | 46.793 | 16 | 51.808 | 163 | 61.774 | 42 | 63.754 | 29 | 81.771 | 45 | 108.765 | 5 | 131.798 | 31
32.207 | 0 | 33.989 | 697 | 41.975 | 19 | 46.99 | 2 | 52.006 | 64 | 61.972 | 32 | 63.952 | 39 | 81.969 | 37 | 108.963 | 7 | 131.996 | 35
32.405 | 0 | 34.187 | 776 | 42.173 | 7 | 47.188 | 2 | 52.204 | 27 | 62.17 | 44 | 64.15 | 25 | 82.167 | 40 | 109.161 | 11 | 132.194 | 45
32.603 | 0 | 34.385 | 707 | 42.371 | 3 | 47.386 | 10 | 52.402 | 40 | 62.368 | 10 | 64.348 | 28 | 82.365 | 31 | 109.359 | 0 | 132.392 | 22
32.801 | 0 | 34.583 | 587 | 42.569 | 0 | 47.584 | 0 | 52.6 | 4 | 62.566 | 35 | 64.546 | 33 | 82.563 | 68 | 109.556 | 4 | 132.59 | 29
32.999 | 0 | 34.781 | 434 | 42.767 | 10 | 47.782 | 0 | 52.798 | 9 | 62.764 | 19 | 64.744 | 2 | 82.761 | 49 | 109.754 | 10 | 132.788 | 23
33.197 | 0 | 34.979 | 553 | 42.965 | 7 | 47.98 | 0 | 52.996 | 0 | 62.962 | 25 | 64.942 | 19 | 82.959 | 71 | 109.952 | 9 | 132.986 | 0
33.395 | 0 | 35.177 | 410 | 43.163 | 3 | 48.178 | 11 | 53.194 | 0 | 63.16 | 57 | 65.14 | 26 | 83.157 | 68 | 110.15 | 2 | 133.184 | 9
33.593 | 0 | 35.375 | 300 | 43.361 | 0 | 48.376 | 4 | 53.392 | 0 | 63.358 | 14 | 65.338 | 22 | 83.355 | 19 | 110.348 | 29 | 133.382 | 19
33.791 | 0 | 35.573 | 349 | 43.559 | 7 | 48.574 | 3 | 53.59 | 6 | 63.556 | 22 | 65.536 | 24 | 83.553 | 31 | 110.546 | 17 | 133.58 | 12
33.989 | 0 | 35.771 | 242 | 43.757 | 2 | 48.772 | 0 | 53.788 | 0 | 63.754 | 18 | 65.734 | 21 | 83.751 | 32 | 110.744 | 17 | 133.778 | 12
34.187 | 0 | 35.969 | 199 | 43.955 | 4 | 48.97 | 3 | 53.986 | 0 | 63.952 | 8 | 65.932 | 28 | 83.949 | 38 | 110.942 | 26 | 133.976 | 7
34.385 | 0 | 36.167 | 171 | 44.153 | 0 | 49.168 | 4 | 54.184 | 0 | 64.15 | 20 | 66.13 | 15 | 84.147 | 36 | 111.14 | 34 | 134.174 | 3
34.583 | 0 | 36.365 | 174 | 44.351 | 0 | 49.366 | 0 | 54.382 | 0 | 64.348 | 26 | 66.328 | 6 | 84.345 | 28 | 111.338 | 10 | 134.372 | 18
34.781 | 0 | 36.563 | 129 | 44.549 | 0 | 49.564 | 0 | 54.58 | 0 | 64.546 | 29 | 66.526 | 21 | 84.543 | 21 | 111.536 | 40 | 134.57 | 14
34.979 | 0 | 36.761 | 125 | 44.747 | 2 | 49.762 | 0 | 54.778 | 0 | 64.744 | 42 | 66.724 | 40 | 84.741 | 14 | 111.734 | 22 | 134.768 | 11
35.177 | 0 | 36.959 | 101 | 44.945 | 0 | 49.96 | 12 | 54.976 | 5 | 64.942 | 25 | 66.922 | 0 | 84.939 | 23 | 111.932 | 9 | 134.966 | 18
35.375 | 0 | 37.157 | 76 | 45.143 | 0 | 50.158 | 0 | 55.174 | 0 | 65.14 | 11 | 67.12 | 13 | 85.137 | 22 | 112.13 | 0 | 135.164 | 0
35.573 | 0 | 37.355 | 79 | 45.341 | 8 | 50.356 | 0 | 55.372 | 0 | 65.338 | 14 | 67.318 | 0 | 85.335 | 22 | 112.328 | 17 | 135.362 | 9
35.771 | 0 | 37.553 | 88 | 45.539 | 5 | 50.554 | 0 | 55.57 | 0 | 65.536 | 10 | 67.516 | 4 | 85.533 | 17 | 112.526 | 13 | 135.56 | 26
35.969 | 0 | 37.751 | 62 | 45.737 | 4 | 50.752 | 0 | 55.768 | 0 | 65.734 | 16 | 67.714 | 7 | 85.731 | 7 | 112.724 | 5 | 135.758 | 4
36.167 | 0 | 37.949 | 52 | 45.935 | 3 | 50.95 | 0 | 55.966 | 0 | 65.932 | 13 | 67.912 | 0 | 85.929 | 12 | 112.922 | 7 | 135.956 | 10
36.365 | 0 | 38.147 | 108 | 46.133 | 0 | 51.148 | 0 | 56.164 | 0 | 66.13 | 18 | 68.11 | 11 | 86.127 | 14 | 113.12 | 5 | 136.154 | 14
36.563 | 0 | 38.345 | 91 | 46.331 | 2 | 51.346 | 0 | 56.362 | 0 | 66.328 | 12 | 68.308 | 11 | 86.325 | 38 | 113.318 | 10 | 136.352 | 0
36.761 | 0 | 38.543 | 45 | 46.529 | 8 | 51.544 | 0 | 56.56 | 0 | 66.526 | 27 | 68.506 | 0 | 86.523 | 8 | 113.516 | 8 | 136.55 | 11
36.959 | 0 | 38.741 | 66 | 46.727 | 24 | 51.742 | 0 | 56.758 | 0 | 66.724 | 11 | 68.704 | 0 | 86.721 | 4 | 113.714 | 10 | 136.748 | 6
37.157 | 0 | 38.939 | 36 | 46.924 | 3 | 51.94 | 0 | 56.956 | 0 | 66.922 | 22 | 68.902 | 0 | 86.919 | 25 | 113.912 | 15 | 136.946 | 0
37.355 | 0 | 39.137 | 43 | 47.122 | 15 | 52.138 | 0 | 57.154 | 0 | 67.12 | 0 | 69.1 | 9 | 87.117 | 0 | 114.11 | 13 | 137.144 | 5
37.553 | 0 | 39.335 | 31 | 47.32 | 12 | 52.336 | 0 | 57.352 | 0 | 67.318 | 0 | 69.298 | 9 | 87.315 | 5 | 114.308 | 20 | 137.342 | 0
37.751 | 0 | 39.533 | 54 | 47.518 | 5 | 52.534 | 0 | 57.55 | 0 | 67.516 | 4 | 69.496 | 18 | 87.513 | 8 | 114.506 | 15 | 137.54 | 0
37.949 | 0 | 39.731 | 53 | 47.716 | 28 | 52.732 | 0 | 57.748 | 0 | 67.714 | 11 | 69.694 | 0 | 87.711 | 3 | 114.704 | 0 | 137.738 | 0
38.147 | 0 | 39.929 | 27 | 47.914 | 9 | 52.93 | 11 | 57.946 | 0 | 67.912 | 5 | 69.892 | 11 | 87.909 | 6 | 114.902 | 6 | 137.936 | 0
38.345 | 0 | 40.127 | 25 | 48.112 | 0 | 53.128 | 4 | 58.144 | 0 | 68.11 | 8 | 70.09 | 11 | 88.107 | 9 | 115.1 | 2 | 138.134 | 0
38.543 | 0 | 40.325 | 47 | 48.31 | 5 | 53.326 | 0 | 58.342 | 0 | 68.308 | 0 | 70.288 | 10 | 88.305 | 6 | 115.298 | 3 | 138.332 | 0
38.741 | 0 | 40.523 | 25 | 48.508 | 7 | 53.524 | 0 | 58.54 | 0 | 68.506 | 3 | 70.486 | 0 | 88.503 | 13 | 115.496 | 0 | 138.53 | 0
38.939 | 0 | 40.721 | 19 | 48.706 | 0 | 53.722 | 0 | 58.738 | 0 | 68.704 | 0 | 70.684 | 6 | 88.701 | 0 | 115.694 | 0 | 138.728 | 0
39.137 | 0 | 40.919 | 10 | 48.904 | 11 | 53.92 | 0 | 58.936 | 0 | 68.902 | 11 | 70.882 | 10 | 88.899 | 0 | 115.892 | 0 | 138.926 | 0
39.335 | 0 | 41.117 | 45 | 49.102 | 6 | 54.118 | 0 | 59.134 | 0 | 69.1 | 0 | 71.08 | 4 | 89.097 | 0 | 116.09 | 0 | 139.124 | 0
39.533 | 0 | 41.315 | 22 | 49.3 | 0 | 54.316 | 0 | 59.332 | 0 | 69.298 | 7 | 71.278 | 16 | 89.295 | 0 | 116.288 | 0 | 139.322 | 0
39.731 | 0 | 41.513 | 3 | 49.498 | 0 | 54.514 | 0 | 59.53 | 0 | 69.496 | 4 | 71.476 | 0 | 89.493 | 0 | 116.486 | 0 | 139.52 | 0
39.929 | 0 | 41.711 | 43 | 49.696 | 5 | 54.712 | 0 | 59.728 | 0 | 69.694 | 0 | 71.674 | 10 | 89.691 | 2 | 116.684 | 0 | 139.718 | 0
40.127 | 0 | 41.909 | 10 | 49.894 | 10 | 54.91 | 0 | 59.926 | 0 | 69.892 | 4 | 71.872 | 21 | 89.889 | 0 | 116.882 | 0 | 139.916 | 0
40.325 | 0 | 42.107 | 7 | 50.092 | 0 | 55.108 | 0 | 60.124 | 0 | 70.09 | 20 | 72.07 | 10 | 90.087 | 0 | 117.08 | 14 | 140.114 | 0
40.523 | 0 | 42.305 | 14 | 50.29 | 5 | 55.306 | 7 | 60.322 | 0 | 70.288 | 10 | 72.268 | 15 | 90.285 | 0 | 117.278 | 0 | 140.312 | 0
40.721 | 0 | 42.503 | 27 | 50.488 | 0 | 55.504 | 3 | 60.52 | 0 | 70.486 | 15 | 72.466 | 14 | 90.483 | 0 | 117.476 | 0 | 140.51 | 0
40.919 | 0 | 42.701 | 3 | 50.686 | 0 | 55.702 | 0 | 60.718 | 0 | 70.684 | 29 | 72.664 | 0 | 90.681 | 0 | 117.674 | 0 | 140.708 | 0
41.117 | 0 | 42.899 | 0 | 50.884 | 0 | 55.9 | 0 | 60.916 | 0 | 70.882 | 8 | 72.862 | 0 | 90.879 | 0 | 117.872 | 0 | 140.906 | 4
41.315 | 0 | 43.097 | 6 | 51.082 | 0 | 56.098 | 0 | 61.114 | 0 | 71.08 | 3 | 73.06 | 4 | 91.077 | 0 | 118.07 | 0 | 141.103 | 0
41.513 | 0 | 43.295 | 9 | 51.28 | 0 | 56.296 | 6 | 61.312 | 0 | 71.278 | 0 | 73.258 | 0 | 91.275 | 0 | 118.268 | 8 | 141.301 | 0
41.711 | 0 | 43.493 | 4 | 51.478 | 0 | 56.494 | 0 | 61.51 | 0 | 71.476 | 10 | 73.456 | 0 | 91.473 | 0 | 118.466 | 0 | 141.499 | 0
41.909 | 0 | 43.691 | 5 | 51.676 | 0 | 56.692 | 0 | 61.708 | 0 | 71.674 | 0 | 73.654 | 3 | 91.671 | 0 | 118.664 | 0 | 141.697 | 0
42.107 | 0 | 43.889 | 12 | 51.874 | 4 | 56.89 | 0 | 61.906 | 0 | 71.872 | 6 | 73.852 | 0 | 91.869 | 3 | 118.862 | 4 | 141.895 | 0
42.305 | 0 | 44.087 | 0 | 52.072 | 3 | 57.088 | 0 | 62.104 | 0 | 72.07 | 0 | 74.05 | 0 | 92.067 | 0 | 119.06 | 0 | 142.093 | 0
42.503 | 0 | 44.285 | 0 | 52.27 | 0 | 57.286 | 0 | 62.302 | 0 | 72.268 | 0 | 74.248 | 0 | 92.265 | 0 | 119.258 | 4 | 142.291 | 0
42.701 | 0 | 44.483 | 9 | 52.468 | 0 | 57.484 | 0 | 62.5 | 0 | 72.466 | 1 | 74.446 | 0 | 92.463 | 0 | 119.456 | 0 | 142.489 | 0
42.899 | 0 | 44.681 | 0 | 52.666 | 0 | 57.682 | 0 | 62.698 | 0 | 72.664 | 0 | 74.644 | 3 | 92.661 | 0 | 119.654 | 0 | 142.687 | 0
43.097 | 0 | 44.879 | 0 | 52.864 | 0 | 57.88 | 0 | 62.896 | 0 | 72.862 | 0 | 74.842 | 0 | 92.859 | 0 | 119.852 | 0 | 142.885 | 0
43.295 | 0 | 45.077 | 3 | 53.062 | 0 | 58.078 | 0 | 63.094 | 0 | 73.06 | 0 | 75.04 | 0 | 93.057 | 4 | 120.05 | 0 | 143.083 | 0
43.493 | 0 | 45.275 | 4 | 53.26 | 0 | 58.276 | 0 | 63.292 | 0 | 73.258 | 0 | 75.238 | 0 | 93.255 | 0 | 120.248 | 0 | 143.281 | 0
43.691 | 0 | 45.473 | 0 | 53.458 | 0 | 58.474 | 0 | 63.49 | 0 | 73.456 | 0 | 75.436 | 0 | 93.453 | 0 | 120.446 | 6 | 143.479 | 0
43.889 | 0 | 45.671 | 0 | 53.656 | 0 | 58.672 | 0 | 63.688 | 0 | 73.654 | 0 | 75.634 | 0 | 93.651 | 0 | 120.644 | 4 | 143.677 | 0
44.087 | 0 | 45.869 | 0 | 53.854 | 0 | 58.87 | 0 | 63.886 | 0 | 73.852 | 0 | 75.832 | 0 | 93.849 | 0 | 120.842 | 0 | 143.875 | 0
44.285 | 0 | 46.067 | 0 | 54.052 | 0 | 59.068 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 0 | 94.047 | 0 | 121.04 | 0 | 144.073 | 0
44.483 | 0 | 46.265 | 35 | 54.25 | 9 | 59.266 | 0 | 64.282 | 7 | 74.248 | 0 | 76.228 | 0 | 94.245 | 0 | 121.238 | 0 | 144.271 | 0
44.681 | 0 | 46.463 | 7 | 54.448 | 4 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 0 | 94.443 | 0 | 121.436 | 0 | 144.469 | 0
44.879 | 0 | 46.661 | 23 | 54.646 | 0 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 0 | 121.634 | 5 | 144.667 | 0
45.077 | 0 | 46.859 | 18 | 54.844 | 0 | 59.86 | 0 | 64.876 | 0 | 74.842 | 0 | 76.822 | 0 | 94.839 | 0 | 121.832 | 0 | 144.865 | 0
45.275 | 0 | 47.056 | 0 | 55.042 | 0 | 60.058 | 0 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 0 | 122.03 | 0 | 145.063 | 0
45.473 | 0 | 47.254 | 0 | 55.24 | 0 | 60.256 | 0 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 0 | 122.228 | 0 | 145.261 | 0
45.671 | 0 | 47.452 | 12 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 122.426 | 0 | 145.459 | 0
45.869 | 0 | 47.65 | 8 | 55.636 | 0 | 60.652 | 0 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 122.624 | 0 | 145.657 | 0
46.067 | 0 | 47.848 | 0 | 55.834 | 0 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 0 | 122.822 | 5 | 145.855 | 0
46.265 | 0 | 48.046 | 0 | 56.032 | 0 | 61.048 | 0 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 0 | 123.02 | 0 | 146.053 | 0
46.463 | 0 | 48.244 | 16 | 56.23 | 0 | 61.246 | 0 | 66.262 | 6 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 123.218 | 0 | 146.251 | 0
46.661 | 0 | 48.442 | 17 | 56.428 | 0 | 61.444 | 0 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 7 | 123.416 | 9 | 146.449 | 0
46.859 | 0 | 48.64 | 0 | 56.626 | 0 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 123.614 | 0 | 146.647 | 0
47.056 | 0 | 48.838 | 6 | 56.824 | 6 | 61.84 | 0 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 123.812 | 0 | 146.845 | 0
47.254 | 0 | 49.036 | 0 | 57.022 | 0 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 0 | 124.01 | 0 | 147.043 | 0
47.452 | 0 | 49.234 | 9 | 57.22 | 6 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 124.208 | 0 | 147.241 | 0
47.65 | 0 | 49.432 | 0 | 57.418 | 0 | 62.434 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 124.406 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 11 | 57.616 | 0 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 124.604 | 0 | 147.637 | 0
48.046 | 0 | 49.828 | 8 | 57.814 | 0 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 0 | 124.802 | 0 | 147.835 | 0
48.244 | 0 | 50.026 | 3 | 58.012 | 2 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 125 | 0 | 148.033 | 0
48.442 | 0 | 50.224 | 6 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 125.198 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 0 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 125.396 | 0 | 148.429 | 9
48.838 | 0 | 50.62 | 0 | 58.606 | 7 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 0 | 125.594 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 0 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 125.792 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 0 | 59.002 | 0 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 0 | 125.99 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 0 | 59.2 | 0 | 64.216 | 0 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 7 | 126.188 | 0 | 149.221 | 0
|  | 51.412 | 2 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 126.386 | 0 | 149.419 | 0

## Analyte Reserpine
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 34 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 8061 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 8 | 110.81 | 0
11.22 | 31983 | 13.002 | 0 | 20.987 | 156 | 26.003 | 30 | 31.019 | 181 | 40.985 | 160 | 42.965 | 159 | 60.982 | 184 | 85.995 | 69 | 111.008 | 118
11.418 | 41206 | 13.2 | 0 | 21.185 | 1369 | 26.201 | 343 | 31.217 | 784 | 41.183 | 821 | 43.163 | 1115 | 61.18 | 1684 | 86.193 | 1329 | 111.206 | 1089
11.616 | 59357 | 13.398 | 0 | 21.383 | 1906 | 26.399 | 1033 | 31.415 | 3106 | 41.381 | 1489 | 43.361 | 1657 | 61.378 | 2253 | 86.391 | 2329 | 111.404 | 1703
11.814 | 39299 | 13.596 | 14 | 21.581 | 3089 | 26.597 | 1550 | 31.613 | 12982 | 41.579 | 2770 | 43.559 | 3329 | 61.576 | 3455 | 86.589 | 6984 | 111.602 | 2778
12.012 | 292590 | 13.794 | 13 | 21.779 | 3249 | 26.795 | 2045 | 31.811 | 11724 | 41.777 | 4346 | 43.757 | 4425 | 61.774 | 3598 | 86.787 | 7259 | 111.8 | 3245
12.21 | 952711 | 13.992 | 0 | 21.977 | 3658 | 26.993 | 2814 | 32.009 | 5971 | 41.975 | 4856 | 43.955 | 4330 | 61.972 | 3321 | 86.985 | 5809 | 111.998 | 2599
12.408 | 914103 | 14.19 | 3 | 22.175 | 4494 | 27.191 | 3978 | 32.207 | 5721 | 42.173 | 5476 | 44.153 | 4078 | 62.17 | 3219 | 87.183 | 5648 | 112.196 | 2567
12.606 | 207554 | 14.388 | 0 | 22.373 | 6151 | 27.389 | 3755 | 32.405 | 7191 | 42.371 | 6401 | 44.351 | 3564 | 62.368 | 3284 | 87.381 | 5707 | 112.394 | 2495
12.804 | 9002 | 14.586 | 0 | 22.571 | 8152 | 27.587 | 3795 | 32.603 | 8924 | 42.569 | 7905 | 44.549 | 3510 | 62.566 | 3058 | 87.579 | 5529 | 112.592 | 2408
13.002 | 579 | 14.784 | 22 | 22.769 | 7761 | 27.785 | 3859 | 32.801 | 11242 | 42.767 | 8616 | 44.747 | 3325 | 62.764 | 3397 | 87.777 | 5284 | 112.79 | 2473
13.2 | 83 | 14.982 | 12 | 22.967 | 6921 | 27.983 | 3490 | 32.999 | 14903 | 42.965 | 8138 | 44.945 | 3422 | 62.962 | 3905 | 87.975 | 5115 | 112.988 | 2745
13.398 | 3 | 15.18 | 21 | 23.165 | 6767 | 28.181 | 3321 | 33.197 | 36190 | 43.163 | 7255 | 45.143 | 3463 | 63.16 | 3768 | 88.173 | 5428 | 113.186 | 2646
13.596 | 0 | 15.378 | 10 | 23.363 | 7200 | 28.379 | 2420 | 33.395 | 110897 | 43.361 | 6016 | 45.341 | 3455 | 63.358 | 3615 | 88.371 | 4973 | 113.384 | 2871
13.794 | 0 | 15.576 | 7 | 23.561 | 8640 | 28.577 | 2487 | 33.593 | 269378 | 43.559 | 5046 | 45.539 | 3391 | 63.556 | 4212 | 88.569 | 4918 | 113.582 | 2777
13.992 | 0 | 15.773 | 21 | 23.759 | 10211 | 28.775 | 2814 | 33.791 | 484711 | 43.757 | 4279 | 45.737 | 3926 | 63.754 | 4373 | 88.767 | 4423 | 113.78 | 2421
14.19 | 0 | 15.971 | 54 | 23.957 | 12066 | 28.973 | 3725 | 33.989 | 626873 | 43.955 | 3730 | 45.935 | 4280 | 63.952 | 4562 | 88.965 | 4757 | 113.978 | 2605
14.388 | 0 | 16.169 | 34 | 24.155 | 13606 | 29.171 | 5027 | 34.187 | 610326 | 44.153 | 3457 | 46.133 | 4405 | 64.15 | 4766 | 89.163 | 5062 | 114.176 | 2463
14.586 | 0 | 16.367 | 100 | 24.353 | 15542 | 29.369 | 6001 | 34.385 | 450889 | 44.351 | 3416 | 46.331 | 5215 | 64.348 | 4887 | 89.361 | 4950 | 114.374 | 2334
14.784 | 0 | 16.565 | 161 | 24.551 | 15182 | 29.567 | 8479 | 34.583 | 251226 | 44.549 | 3712 | 46.529 | 5272 | 64.546 | 4480 | 89.559 | 4893 | 114.572 | 2547
14.982 | 0 | 16.763 | 387 | 24.749 | 12086 | 29.765 | 11860 | 34.781 | 109290 | 44.747 | 3395 | 46.727 | 5352 | 64.744 | 4896 | 89.757 | 5285 | 114.77 | 2783
15.18 | 0 | 16.961 | 1210 | 24.947 | 8946 | 29.963 | 17245 | 34.979 | 40078 | 44.945 | 3079 | 46.924 | 5363 | 64.942 | 4540 | 89.955 | 4835 | 114.968 | 2775
15.378 | 0 | 17.159 | 3211 | 25.145 | 7027 | 30.161 | 23784 | 35.177 | 15225 | 45.143 | 3147 | 47.122 | 5769 | 65.14 | 4828 | 90.153 | 5153 | 115.166 | 2668
15.576 | 0 | 17.357 | 5029 | 25.343 | 5849 | 30.359 | 29600 | 35.375 | 8529 | 45.341 | 3537 | 47.32 | 4907 | 65.338 | 4999 | 90.351 | 4551 | 115.364 | 2826
15.773 | 0 | 17.555 | 4960 | 25.541 | 5119 | 30.557 | 31049 | 35.573 | 6917 | 45.539 | 3731 | 47.518 | 5500 | 65.536 | 4930 | 90.549 | 4870 | 115.562 | 3124
15.971 | 0 | 17.753 | 7255 | 25.739 | 5198 | 30.755 | 27374 | 35.771 | 6593 | 45.737 | 3920 | 47.716 | 6395 | 65.734 | 5024 | 90.747 | 4770 | 115.76 | 2846
16.169 | 0 | 17.951 | 9089 | 25.937 | 4202 | 30.953 | 22686 | 35.969 | 6754 | 45.935 | 4326 | 47.914 | 6929 | 65.932 | 5219 | 90.945 | 4828 | 115.958 | 3145
16.367 | 0 | 18.149 | 6770 | 26.135 | 2832 | 31.151 | 17594 | 36.167 | 6383 | 46.133 | 4560 | 48.112 | 7857 | 66.13 | 5999 | 91.143 | 4981 | 116.156 | 3456
16.565 | 0 | 18.347 | 2766 | 26.333 | 1825 | 31.349 | 14291 | 36.365 | 6717 | 46.331 | 5105 | 48.31 | 8547 | 66.328 | 6365 | 91.341 | 4669 | 116.354 | 3562
16.763 | 0 | 18.545 | 2644 | 26.531 | 1164 | 31.547 | 11504 | 36.563 | 6842 | 46.529 | 5282 | 48.508 | 10178 | 66.526 | 6643 | 91.539 | 4219 | 116.552 | 3272
16.961 | 0 | 18.743 | 4303 | 26.729 | 1215 | 31.745 | 9000 | 36.761 | 6765 | 46.727 | 5634 | 48.706 | 11759 | 66.724 | 6351 | 91.737 | 4462 | 116.75 | 3478
17.159 | 0 | 18.941 | 4588 | 26.927 | 1877 | 31.943 | 7030 | 36.959 | 6718 | 46.924 | 5746 | 48.904 | 14010 | 66.922 | 7142 | 91.935 | 4137 | 116.948 | 3561
17.357 | 0 | 19.139 | 2827 | 27.125 | 2613 | 32.141 | 5902 | 37.157 | 6725 | 47.122 | 5476 | 49.102 | 15143 | 67.12 | 7233 | 92.133 | 3589 | 117.146 | 3694
17.555 | 0 | 19.337 | 3045 | 27.323 | 3593 | 32.339 | 6316 | 37.355 | 7714 | 47.32 | 5377 | 49.3 | 16243 | 67.318 | 8030 | 92.331 | 3679 | 117.344 | 3277
17.753 | 0 | 19.535 | 4651 | 27.521 | 3723 | 32.537 | 7047 | 37.553 | 7914 | 47.518 | 5236 | 49.498 | 16236 | 67.516 | 8622 | 92.529 | 3272 | 117.542 | 3602
17.951 | 0 | 19.733 | 5676 | 27.719 | 3568 | 32.735 | 9393 | 37.751 | 8942 | 47.716 | 5047 | 49.696 | 15608 | 67.714 | 8693 | 92.727 | 2895 | 117.74 | 3493
18.149 | 0 | 19.931 | 4271 | 27.917 | 2770 | 32.933 | 11493 | 37.949 | 7929 | 47.914 | 5242 | 49.894 | 13942 | 67.912 | 8635 | 92.925 | 2487 | 117.938 | 2958
18.347 | 0 | 20.129 | 3527 | 28.115 | 1982 | 33.131 | 15998 | 38.147 | 7896 | 48.112 | 6447 | 50.092 | 13391 | 68.11 | 8698 | 93.123 | 2277 | 118.136 | 2463
18.545 | 0 | 20.327 | 4088 | 28.313 | 1495 | 33.329 | 39081 | 38.345 | 6110 | 48.31 | 7364 | 50.29 | 12325 | 68.308 | 8869 | 93.321 | 2071 | 118.334 | 2245
18.743 | 0 | 20.525 | 4167 | 28.511 | 1482 | 33.527 | 118696 | 38.543 | 6039 | 48.508 | 8381 | 50.488 | 12392 | 68.506 | 8175 | 93.519 | 2087 | 118.532 | 2314
18.941 | 0 | 20.723 | 4042 | 28.709 | 1970 | 33.725 | 289191 | 38.741 | 4677 | 48.706 | 10340 | 50.686 | 12139 | 68.704 | 7513 | 93.717 | 2367 | 118.73 | 2100
19.139 | 0 | 20.921 | 3070 | 28.907 | 2359 | 33.923 | 514864 | 38.939 | 3902 | 48.904 | 12477 | 50.884 | 11561 | 68.902 | 7179 | 93.915 | 2068 | 118.928 | 1933
19.337 | 0 | 21.119 | 2897 | 29.105 | 3158 | 34.121 | 670646 | 39.137 | 3321 | 49.102 | 14111 | 51.082 | 11194 | 69.1 | 6814 | 94.113 | 1618 | 119.126 | 1823
19.535 | 0 | 21.317 | 3117 | 29.303 | 3558 | 34.319 | 652503 | 39.335 | 2562 | 49.3 | 14686 | 51.28 | 10129 | 69.298 | 6029 | 94.311 | 1653 | 119.324 | 1592
19.733 | 0 | 21.515 | 3548 | 29.501 | 4794 | 34.517 | 478328 | 39.533 | 2879 | 49.498 | 16152 | 51.478 | 10030 | 69.496 | 5312 | 94.509 | 1567 | 119.522 | 1932
19.931 | 0 | 21.713 | 2959 | 29.699 | 6285 | 34.715 | 263552 | 39.731 | 2445 | 49.696 | 15283 | 51.676 | 8037 | 69.694 | 5152 | 94.707 | 1651 | 119.72 | 1663
20.129 | 0 | 21.911 | 3236 | 29.897 | 10081 | 34.913 | 115192 | 39.929 | 2492 | 49.894 | 14785 | 51.874 | 6526 | 69.892 | 4616 | 94.905 | 1730 | 119.918 | 1848
20.327 | 0 | 22.109 | 3053 | 30.095 | 15397 | 35.111 | 40927 | 40.127 | 2432 | 50.092 | 13241 | 52.072 | 5677 | 70.09 | 3868 | 95.103 | 1513 | 120.116 | 1580
20.525 | 0 | 22.307 | 3671 | 30.293 | 21440 | 35.309 | 12469 | 40.325 | 2411 | 50.29 | 12717 | 52.27 | 5169 | 70.288 | 3863 | 95.301 | 1683 | 120.314 | 1538
20.723 | 0 | 22.505 | 4214 | 30.491 | 27599 | 35.507 | 4624 | 40.523 | 2466 | 50.488 | 11631 | 52.468 | 4095 | 70.486 | 3828 | 95.499 | 1817 | 120.512 | 1788
20.921 | 0 | 22.703 | 5554 | 30.689 | 28894 | 35.705 | 3166 | 40.721 | 2136 | 50.686 | 11122 | 52.666 | 3529 | 70.684 | 3943 | 95.697 | 2084 | 120.71 | 1782
21.119 | 0 | 22.901 | 5882 | 30.887 | 28327 | 35.903 | 3092 | 40.919 | 1767 | 50.884 | 10250 | 52.864 | 3265 | 70.882 | 3639 | 95.895 | 2045 | 120.908 | 1866
21.317 | 0 | 23.099 | 6540 | 31.085 | 23704 | 36.101 | 3028 | 41.117 | 1710 | 51.082 | 10772 | 53.062 | 3120 | 71.08 | 3650 | 96.093 | 2205 | 121.106 | 1823
21.515 | 0 | 23.297 | 5497 | 31.283 | 18251 | 36.299 | 3405 | 41.315 | 1284 | 51.28 | 10044 | 53.26 | 2898 | 71.278 | 3301 | 96.291 | 2122 | 121.304 | 2353
21.713 | 0 | 23.495 | 5314 | 31.481 | 13799 | 36.497 | 3329 | 41.513 | 1000 | 51.478 | 9175 | 53.458 | 3374 | 71.476 | 3452 | 96.489 | 2196 | 121.502 | 2766
21.911 | 0 | 23.693 | 5333 | 31.679 | 11012 | 36.695 | 3657 | 41.711 | 828 | 51.676 | 7905 | 53.656 | 3728 | 71.674 | 2662 | 96.687 | 1812 | 121.7 | 4009
22.109 | 0 | 23.891 | 5913 | 31.877 | 8232 | 36.893 | 4009 | 41.909 | 960 | 51.874 | 6443 | 53.854 | 4466 | 71.872 | 2453 | 96.885 | 2069 | 121.898 | 6493
22.307 | 0 | 24.089 | 6570 | 32.075 | 5755 | 37.091 | 4195 | 42.107 | 1198 | 52.072 | 5529 | 54.052 | 4742 | 72.07 | 2314 | 97.083 | 1664 | 122.096 | 10202
22.505 | 0 | 24.287 | 7771 | 32.273 | 3927 | 37.289 | 5184 | 42.305 | 1955 | 52.27 | 4514 | 54.25 | 4313 | 72.268 | 2382 | 97.281 | 1474 | 122.294 | 16059
22.703 | 0 | 24.485 | 10876 | 32.471 | 3697 | 37.487 | 5358 | 42.503 | 2615 | 52.468 | 3896 | 54.448 | 3968 | 72.466 | 1818 | 97.479 | 1471 | 122.492 | 23335
22.901 | 0 | 24.683 | 12556 | 32.669 | 4719 | 37.685 | 6247 | 42.701 | 3792 | 52.666 | 3212 | 54.646 | 3342 | 72.664 | 1875 | 97.677 | 1102 | 122.69 | 35516
23.099 | 0 | 24.881 | 12571 | 32.867 | 6712 | 37.883 | 6372 | 42.899 | 4522 | 52.864 | 3031 | 54.844 | 2457 | 72.862 | 2053 | 97.875 | 1065 | 122.888 | 47322
23.297 | 0 | 25.079 | 10255 | 33.065 | 8282 | 38.081 | 6965 | 43.097 | 4719 | 53.062 | 2291 | 55.042 | 2772 | 73.06 | 2311 | 98.073 | 1066 | 123.086 | 57573
23.495 | 0 | 25.277 | 7256 | 33.263 | 13749 | 38.279 | 6044 | 43.295 | 4139 | 53.26 | 2523 | 55.24 | 5748 | 73.258 | 2509 | 98.271 | 995 | 123.284 | 68593
23.693 | 0 | 25.475 | 4946 | 33.461 | 40250 | 38.477 | 5289 | 43.493 | 3146 | 53.458 | 2856 | 55.438 | 14336 | 73.456 | 2613 | 98.469 | 960 | 123.482 | 73611
23.891 | 0 | 25.673 | 3369 | 33.659 | 124912 | 38.675 | 4339 | 43.691 | 2336 | 53.656 | 3328 | 55.636 | 38018 | 73.654 | 2636 | 98.667 | 855 | 123.68 | 74283
24.089 | 0 | 25.871 | 2783 | 33.857 | 310649 | 38.873 | 3187 | 43.889 | 1576 | 53.854 | 3617 | 55.834 | 84614 | 73.852 | 2668 | 98.865 | 873 | 123.878 | 72631
24.287 | 0 | 26.069 | 1703 | 34.055 | 561397 | 39.071 | 3117 | 44.087 | 891 | 54.052 | 4335 | 56.032 | 162040 | 74.05 | 3102 | 99.063 | 939 | 124.076 | 64172
24.485 | 0 | 26.267 | 1392 | 34.253 | 734315 | 39.269 | 2420 | 44.285 | 761 | 54.25 | 4322 | 56.23 | 253721 | 74.248 | 2703 | 99.261 | 1328 | 124.274 | 53901
24.683 | 0 | 26.465 | 984 | 34.451 | 715115 | 39.467 | 2424 | 44.483 | 677 | 54.448 | 4419 | 56.428 | 336479 | 74.446 | 2426 | 99.459 | 2186 | 124.472 | 41206
24.881 | 0 | 26.663 | 887 | 34.649 | 527616 | 39.665 | 1863 | 44.681 | 862 | 54.646 | 3316 | 56.626 | 373683 | 74.644 | 2630 | 99.657 | 4904 | 124.67 | 30836
25.079 | 0 | 26.861 | 782 | 34.847 | 293244 | 39.863 | 1992 | 44.879 | 721 | 54.844 | 2694 | 56.824 | 358253 | 74.842 | 2735 | 99.855 | 9174 | 124.868 | 21973
25.277 | 0 | 27.059 | 748 | 35.045 | 125276 | 40.061 | 1575 | 45.077 | 786 | 55.042 | 2833 | 57.022 | 291272 | 75.04 | 2533 | 100.053 | 16761 | 125.066 | 14493
25.475 | 0 | 27.257 | 775 | 35.243 | 41550 | 40.259 | 1484 | 45.275 | 843 | 55.24 | 4644 | 57.22 | 203787 | 75.238 | 2390 | 100.251 | 28532 | 125.264 | 9824
25.673 | 0 | 27.455 | 682 | 35.441 | 11127 | 40.457 | 1672 | 45.473 | 768 | 55.438 | 11328 | 57.418 | 126109 | 75.436 | 2371 | 100.449 | 46119 | 125.462 | 6275
25.871 | 0 | 27.653 | 660 | 35.639 | 2693 | 40.655 | 1371 | 45.671 | 648 | 55.636 | 31044 | 57.616 | 66280 | 75.634 | 1947 | 100.647 | 65764 | 125.66 | 5204
26.069 | 0 | 27.851 | 616 | 35.837 | 868 | 40.853 | 1214 | 45.869 | 850 | 55.834 | 73067 | 57.814 | 30363 | 75.832 | 1826 | 100.845 | 87494 | 125.858 | 3344
26.267 | 0 | 28.049 | 579 | 36.035 | 503 | 41.051 | 836 | 46.067 | 669 | 56.032 | 146156 | 58.012 | 13851 | 76.03 | 1437 | 101.043 | 105139 | 126.056 | 2938
26.465 | 0 | 28.247 | 543 | 36.233 | 534 | 41.249 | 603 | 46.265 | 583 | 56.23 | 246442 | 58.21 | 6533 | 76.228 | 1217 | 101.241 | 116675 | 126.254 | 2630
26.663 | 0 | 28.445 | 462 | 36.431 | 466 | 41.447 | 402 | 46.463 | 782 | 56.428 | 341038 | 58.408 | 3015 | 76.426 | 804 | 101.439 | 117130 | 126.452 | 2196
26.861 | 0 | 28.643 | 359 | 36.629 | 642 | 41.645 | 322 | 46.661 | 732 | 56.626 | 404032 | 58.606 | 1449 | 76.624 | 893 | 101.637 | 109942 | 126.65 | 2234
27.059 | 0 | 28.841 | 438 | 36.827 | 492 | 41.843 | 286 | 46.859 | 701 | 56.824 | 410250 | 58.804 | 825 | 76.822 | 870 | 101.835 | 94575 | 126.848 | 1857
27.257 | 0 | 29.039 | 619 | 37.025 | 507 | 42.041 | 590 | 47.056 | 801 | 57.022 | 345525 | 59.002 | 412 | 77.02 | 904 | 102.033 | 74162 | 127.046 | 1570
27.455 | 0 | 29.237 | 761 | 37.223 | 279 | 42.239 | 937 | 47.254 | 885 | 57.22 | 252752 | 59.2 | 353 | 77.218 | 1699 | 102.231 | 54348 | 127.244 | 1364
27.653 | 0 | 29.435 | 1129 | 37.421 | 195 | 42.437 | 1768 | 47.452 | 937 | 57.418 | 158863 | 59.398 | 227 | 77.416 | 2965 | 102.429 | 36753 | 127.442 | 1207
27.851 | 0 | 29.633 | 1968 | 37.619 | 151 | 42.635 | 2913 | 47.65 | 1495 | 57.616 | 86532 | 59.596 | 209 | 77.614 | 7463 | 102.627 | 23036 | 127.64 | 1038
28.049 | 0 | 29.831 | 3454 | 37.817 | 79 | 42.833 | 3801 | 47.848 | 2072 | 57.814 | 43242 | 59.794 | 113 | 77.812 | 16565 | 102.825 | 13764 | 127.838 | 855
28.247 | 0 | 30.029 | 7127 | 38.015 | 114 | 43.031 | 4596 | 48.046 | 2655 | 58.012 | 19169 | 59.992 | 141 | 78.01 | 33311 | 103.023 | 8332 | 128.036 | 840
28.445 | 0 | 30.227 | 11983 | 38.213 | 106 | 43.229 | 4696 | 48.244 | 3561 | 58.21 | 8370 | 60.19 | 137 | 78.208 | 62248 | 103.221 | 5498 | 128.234 | 747
28.643 | 0 | 30.425 | 18491 | 38.411 | 112 | 43.427 | 4174 | 48.442 | 4442 | 58.408 | 3746 | 60.388 | 142 | 78.405 | 103093 | 103.419 | 3837 | 128.432 | 749
28.841 | 0 | 30.623 | 25809 | 38.609 | 124 | 43.625 | 3565 | 48.64 | 5892 | 58.606 | 1729 | 60.586 | 174 | 78.603 | 150368 | 103.617 | 2900 | 128.63 | 568
29.039 | 0 | 30.821 | 29423 | 38.807 | 126 | 43.823 | 2321 | 48.838 | 8101 | 58.804 | 889 | 60.784 | 222 | 78.801 | 191539 | 103.815 | 2850 | 128.828 | 638
29.237 | 0 | 31.019 | 28490 | 39.005 | 93 | 44.021 | 1322 | 49.036 | 10145 | 59.002 | 370 | 60.982 | 313 | 78.999 | 217882 | 104.013 | 2656 | 129.026 | 588
29.435 | 0 | 31.217 | 24376 | 39.203 | 76 | 44.219 | 868 | 49.234 | 11852 | 59.2 | 241 | 61.18 | 293 | 79.197 | 221756 | 104.211 | 2685 | 129.224 | 470
29.633 | 0 | 31.415 | 20332 | 39.401 | 118 | 44.417 | 598 | 49.432 | 13705 | 59.398 | 130 | 61.378 | 336 | 79.395 | 204271 | 104.409 | 2710 | 129.422 | 690
29.831 | 0 | 31.613 | 16200 | 39.599 | 102 | 44.615 | 552 | 49.63 | 14194 | 59.596 | 130 | 61.576 | 262 | 79.593 | 167082 | 104.607 | 2447 | 129.62 | 449
30.029 | 0 | 31.811 | 13155 | 39.797 | 110 | 44.813 | 623 | 49.828 | 14725 | 59.794 | 99 | 61.774 | 242 | 79.791 | 124077 | 104.805 | 2591 | 129.818 | 643
30.227 | 0 | 32.009 | 9564 | 39.995 | 215 | 45.011 | 603 | 50.026 | 13805 | 59.992 | 96 | 61.972 | 133 | 79.989 | 81908 | 105.003 | 3012 | 130.016 | 502
30.425 | 0 | 32.207 | 7171 | 40.193 | 403 | 45.209 | 627 | 50.224 | 11997 | 60.19 | 178 | 62.17 | 110 | 80.187 | 49610 | 105.201 | 2862 | 130.214 | 735
30.623 | 0 | 32.405 | 5016 | 40.391 | 437 | 45.407 | 706 | 50.422 | 10449 | 60.388 | 94 | 62.368 | 110 | 80.385 | 28187 | 105.399 | 2550 | 130.412 | 580
30.821 | 0 | 32.603 | 3652 | 40.589 | 540 | 45.605 | 669 | 50.62 | 8934 | 60.586 | 216 | 62.566 | 88 | 80.583 | 14548 | 105.597 | 1807 | 130.61 | 605
31.019 | 0 | 32.801 | 4339 | 40.787 | 564 | 45.803 | 628 | 50.818 | 7898 | 60.784 | 185 | 62.764 | 68 | 80.781 | 7587 | 105.795 | 1280 | 130.808 | 566
31.217 | 0 | 32.999 | 6221 | 40.985 | 458 | 46.001 | 624 | 51.016 | 6233 | 60.982 | 221 | 62.962 | 45 | 80.979 | 4666 | 105.993 | 1039 | 131.006 | 616
31.415 | 0 | 33.197 | 8289 | 41.183 | 271 | 46.199 | 497 | 51.214 | 5066 | 61.18 | 341 | 63.16 | 48 | 81.177 | 2883 | 106.191 | 715 | 131.204 | 528
31.613 | 0 | 33.395 | 10328 | 41.381 | 333 | 46.397 | 450 | 51.412 | 3771 | 61.378 | 344 | 63.358 | 101 | 81.375 | 2078 | 106.389 | 407 | 131.402 | 561
31.811 | 0 | 33.593 | 24709 | 41.579 | 288 | 46.595 | 414 | 51.61 | 2488 | 61.576 | 282 | 63.556 | 54 | 81.573 | 1159 | 106.587 | 425 | 131.6 | 425
32.009 | 0 | 33.791 | 81353 | 41.777 | 270 | 46.793 | 455 | 51.808 | 1580 | 61.774 | 215 | 63.754 | 43 | 81.771 | 716 | 106.785 | 309 | 131.798 | 276
32.207 | 0 | 33.989 | 232964 | 41.975 | 398 | 46.99 | 374 | 52.006 | 1030 | 61.972 | 205 | 63.952 | 60 | 81.969 | 406 | 106.983 | 176 | 131.996 | 83
32.405 | 0 | 34.187 | 477036 | 42.173 | 521 | 47.188 | 316 | 52.204 | 402 | 62.17 | 170 | 64.15 | 48 | 82.167 | 285 | 107.181 | 153 | 132.194 | 82
32.603 | 0 | 34.385 | 715430 | 42.371 | 1009 | 47.386 | 401 | 52.402 | 144 | 62.368 | 79 | 64.348 | 61 | 82.365 | 218 | 107.379 | 76 | 132.392 | 71
32.801 | 0 | 34.583 | 783500 | 42.569 | 1773 | 47.584 | 444 | 52.6 | 97 | 62.566 | 86 | 64.546 | 27 | 82.563 | 213 | 107.577 | 69 | 132.59 | 47
32.999 | 0 | 34.781 | 654419 | 42.767 | 2680 | 47.782 | 310 | 52.798 | 87 | 62.764 | 40 | 64.744 | 44 | 82.761 | 209 | 107.775 | 74 | 132.788 | 30
33.197 | 0 | 34.979 | 408177 | 42.965 | 3917 | 47.98 | 422 | 52.996 | 80 | 62.962 | 85 | 64.942 | 37 | 82.959 | 200 | 107.973 | 67 | 132.986 | 65
33.395 | 0 | 35.177 | 194621 | 43.163 | 4946 | 48.178 | 316 | 53.194 | 19 | 63.16 | 40 | 65.14 | 73 | 83.157 | 154 | 108.171 | 65 | 133.184 | 57
33.593 | 0 | 35.375 | 72643 | 43.361 | 5119 | 48.376 | 214 | 53.392 | 38 | 63.358 | 72 | 65.338 | 62 | 83.355 | 212 | 108.369 | 83 | 133.382 | 48
33.791 | 0 | 35.573 | 21238 | 43.559 | 4666 | 48.574 | 250 | 53.59 | 52 | 63.556 | 54 | 65.536 | 42 | 83.553 | 161 | 108.567 | 24 | 133.58 | 58
33.989 | 0 | 35.771 | 5655 | 43.757 | 3586 | 48.772 | 257 | 53.788 | 60 | 63.754 | 46 | 65.734 | 18 | 83.751 | 171 | 108.765 | 80 | 133.778 | 77
34.187 | 0 | 35.969 | 1278 | 43.955 | 2242 | 48.97 | 199 | 53.986 | 41 | 63.952 | 43 | 65.932 | 22 | 83.949 | 116 | 108.963 | 58 | 133.976 | 76
34.385 | 0 | 36.167 | 606 | 44.153 | 1474 | 49.168 | 153 | 54.184 | 51 | 64.15 | 19 | 66.13 | 28 | 84.147 | 173 | 109.161 | 113 | 134.174 | 82
34.583 | 0 | 36.365 | 468 | 44.351 | 884 | 49.366 | 103 | 54.382 | 23 | 64.348 | 59 | 66.328 | 31 | 84.345 | 163 | 109.359 | 86 | 134.372 | 122
34.781 | 0 | 36.563 | 523 | 44.549 | 772 | 49.564 | 145 | 54.58 | 12 | 64.546 | 33 | 66.526 | 20 | 84.543 | 172 | 109.556 | 118 | 134.57 | 125
34.979 | 0 | 36.761 | 681 | 44.747 | 620 | 49.762 | 146 | 54.778 | 30 | 64.744 | 42 | 66.724 | 49 | 84.741 | 153 | 109.754 | 137 | 134.768 | 255
35.177 | 0 | 36.959 | 622 | 44.945 | 634 | 49.96 | 170 | 54.976 | 21 | 64.942 | 29 | 66.922 | 22 | 84.939 | 114 | 109.952 | 159 | 134.966 | 185
35.375 | 0 | 37.157 | 507 | 45.143 | 676 | 50.158 | 210 | 55.174 | 28 | 65.14 | 25 | 67.12 | 49 | 85.137 | 144 | 110.15 | 146 | 135.164 | 277
35.573 | 0 | 37.355 | 377 | 45.341 | 731 | 50.356 | 167 | 55.372 | 15 | 65.338 | 22 | 67.318 | 37 | 85.335 | 160 | 110.348 | 149 | 135.362 | 458
35.771 | 0 | 37.553 | 257 | 45.539 | 724 | 50.554 | 119 | 55.57 | 21 | 65.536 | 24 | 67.516 | 50 | 85.533 | 204 | 110.546 | 126 | 135.56 | 601
35.969 | 0 | 37.751 | 137 | 45.737 | 652 | 50.752 | 173 | 55.768 | 33 | 65.734 | 35 | 67.714 | 37 | 85.731 | 221 | 110.744 | 208 | 135.758 | 527
36.167 | 0 | 37.949 | 166 | 45.935 | 696 | 50.95 | 166 | 55.966 | 24 | 65.932 | 28 | 67.912 | 47 | 85.929 | 220 | 110.942 | 165 | 135.956 | 550
36.365 | 0 | 38.147 | 99 | 46.133 | 571 | 51.148 | 119 | 56.164 | 43 | 66.13 | 46 | 68.11 | 41 | 86.127 | 206 | 111.14 | 176 | 136.154 | 605
36.563 | 0 | 38.345 | 172 | 46.331 | 670 | 51.346 | 194 | 56.362 | 28 | 66.328 | 46 | 68.308 | 106 | 86.325 | 150 | 111.338 | 156 | 136.352 | 792
36.761 | 0 | 38.543 | 94 | 46.529 | 523 | 51.544 | 147 | 56.56 | 20 | 66.526 | 27 | 68.506 | 130 | 86.523 | 199 | 111.536 | 168 | 136.55 | 586
36.959 | 0 | 38.741 | 99 | 46.727 | 581 | 51.742 | 135 | 56.758 | 19 | 66.724 | 19 | 68.704 | 179 | 86.721 | 100 | 111.734 | 144 | 136.748 | 730
37.157 | 0 | 38.939 | 62 | 46.924 | 530 | 51.94 | 168 | 56.956 | 21 | 66.922 | 11 | 68.902 | 222 | 86.919 | 76 | 111.932 | 105 | 136.946 | 610
37.355 | 0 | 39.137 | 89 | 47.122 | 500 | 52.138 | 103 | 57.154 | 10 | 67.12 | 22 | 69.1 | 239 | 87.117 | 88 | 112.13 | 105 | 137.144 | 514
37.553 | 0 | 39.335 | 78 | 47.32 | 470 | 52.336 | 140 | 57.352 | 20 | 67.318 | 32 | 69.298 | 267 | 87.315 | 89 | 112.328 | 80 | 137.342 | 475
37.751 | 0 | 39.533 | 100 | 47.518 | 421 | 52.534 | 86 | 57.55 | 24 | 67.516 | 31 | 69.496 | 209 | 87.513 | 99 | 112.526 | 58 | 137.54 | 334
37.949 | 0 | 39.731 | 125 | 47.716 | 414 | 52.732 | 55 | 57.748 | 14 | 67.714 | 25 | 69.694 | 227 | 87.711 | 57 | 112.724 | 123 | 137.738 | 300
38.147 | 0 | 39.929 | 183 | 47.914 | 521 | 52.93 | 56 | 57.946 | 24 | 67.912 | 69 | 69.892 | 151 | 87.909 | 72 | 112.922 | 104 | 137.936 | 244
38.345 | 0 | 40.127 | 204 | 48.112 | 385 | 53.128 | 68 | 58.144 | 11 | 68.11 | 26 | 70.09 | 136 | 88.107 | 11 | 113.12 | 122 | 138.134 | 189
38.543 | 0 | 40.325 | 327 | 48.31 | 367 | 53.326 | 159 | 58.342 | 9 | 68.308 | 90 | 70.288 | 117 | 88.305 | 43 | 113.318 | 83 | 138.332 | 149
38.741 | 0 | 40.523 | 406 | 48.508 | 228 | 53.524 | 78 | 58.54 | 9 | 68.506 | 151 | 70.486 | 109 | 88.503 | 76 | 113.516 | 150 | 138.53 | 183
38.939 | 0 | 40.721 | 516 | 48.706 | 347 | 53.722 | 72 | 58.738 | 11 | 68.704 | 171 | 70.684 | 60 | 88.701 | 26 | 113.714 | 145 | 138.728 | 50
39.137 | 0 | 40.919 | 543 | 48.904 | 239 | 53.92 | 71 | 58.936 | 21 | 68.902 | 168 | 70.882 | 71 | 88.899 | 36 | 113.912 | 117 | 138.926 | 12
39.335 | 0 | 41.117 | 491 | 49.102 | 274 | 54.118 | 70 | 59.134 | 2 | 69.1 | 171 | 71.08 | 92 | 89.097 | 54 | 114.11 | 136 | 139.124 | 18
39.533 | 0 | 41.315 | 443 | 49.3 | 229 | 54.316 | 42 | 59.332 | 16 | 69.298 | 223 | 71.278 | 34 | 89.295 | 24 | 114.308 | 101 | 139.322 | 35
39.731 | 0 | 41.513 | 389 | 49.498 | 272 | 54.514 | 38 | 59.53 | 9 | 69.496 | 255 | 71.476 | 71 | 89.493 | 26 | 114.506 | 125 | 139.52 | 10
39.929 | 0 | 41.711 | 359 | 49.696 | 194 | 54.712 | 61 | 59.728 | 5 | 69.694 | 254 | 71.674 | 40 | 89.691 | 49 | 114.704 | 67 | 139.718 | 17
40.127 | 0 | 41.909 | 293 | 49.894 | 196 | 54.91 | 26 | 59.926 | 9 | 69.892 | 188 | 71.872 | 69 | 89.889 | 63 | 114.902 | 39 | 139.916 | 9
40.325 | 0 | 42.107 | 288 | 50.092 | 158 | 55.108 | 18 | 60.124 | 0 | 70.09 | 122 | 72.07 | 75 | 90.087 | 45 | 115.1 | 85 | 140.114 | 26
40.523 | 0 | 42.305 | 534 | 50.29 | 125 | 55.306 | 7 | 60.322 | 7 | 70.288 | 83 | 72.268 | 59 | 90.285 | 51 | 115.298 | 68 | 140.312 | 5
40.721 | 0 | 42.503 | 894 | 50.488 | 86 | 55.504 | 31 | 60.52 | 0 | 70.486 | 110 | 72.466 | 97 | 90.483 | 50 | 115.496 | 61 | 140.51 | 32
40.919 | 0 | 42.701 | 1348 | 50.686 | 125 | 55.702 | 32 | 60.718 | 3 | 70.684 | 83 | 72.664 | 101 | 90.681 | 53 | 115.694 | 22 | 140.708 | 18
41.117 | 0 | 42.899 | 2479 | 50.884 | 107 | 55.9 | 7 | 60.916 | 0 | 70.882 | 56 | 72.862 | 126 | 90.879 | 52 | 115.892 | 49 | 140.906 | 19
41.315 | 0 | 43.097 | 3752 | 51.082 | 119 | 56.098 | 41 | 61.114 | 5 | 71.08 | 80 | 73.06 | 314 | 91.077 | 44 | 116.09 | 38 | 141.103 | 26
41.513 | 0 | 43.295 | 5136 | 51.28 | 127 | 56.296 | 40 | 61.312 | 0 | 71.278 | 88 | 73.258 | 380 | 91.275 | 38 | 116.288 | 21 | 141.301 | 19
41.711 | 0 | 43.493 | 5948 | 51.478 | 119 | 56.494 | 6 | 61.51 | 6 | 71.476 | 43 | 73.456 | 671 | 91.473 | 14 | 116.486 | 32 | 141.499 | 8
41.909 | 0 | 43.691 | 5770 | 51.676 | 144 | 56.692 | 38 | 61.708 | 2 | 71.674 | 45 | 73.654 | 1247 | 91.671 | 42 | 116.684 | 15 | 141.697 | 35
42.107 | 0 | 43.889 | 4643 | 51.874 | 180 | 56.89 | 25 | 61.906 | 0 | 71.872 | 66 | 73.852 | 1534 | 91.869 | 30 | 116.882 | 74 | 141.895 | 26
42.305 | 0 | 44.087 | 3197 | 52.072 | 113 | 57.088 | 28 | 62.104 | 0 | 72.07 | 17 | 74.05 | 2029 | 92.067 | 29 | 117.08 | 56 | 142.093 | 13
42.503 | 0 | 44.285 | 2233 | 52.27 | 193 | 57.286 | 8 | 62.302 | 0 | 72.268 | 5 | 74.248 | 2294 | 92.265 | 44 | 117.278 | 70 | 142.291 | 49
42.701 | 0 | 44.483 | 1365 | 52.468 | 107 | 57.484 | 13 | 62.5 | 0 | 72.466 | 24 | 74.446 | 2459 | 92.463 | 70 | 117.476 | 47 | 142.489 | 51
42.899 | 0 | 44.681 | 746 | 52.666 | 120 | 57.682 | 32 | 62.698 | 7 | 72.664 | 2 | 74.644 | 2627 | 92.661 | 33 | 117.674 | 63 | 142.687 | 33
43.097 | 0 | 44.879 | 596 | 52.864 | 142 | 57.88 | 10 | 62.896 | 0 | 72.862 | 6 | 74.842 | 2358 | 92.859 | 44 | 117.872 | 87 | 142.885 | 50
43.295 | 0 | 45.077 | 649 | 53.062 | 87 | 58.078 | 33 | 63.094 | 8 | 73.06 | 0 | 75.04 | 2201 | 93.057 | 45 | 118.07 | 89 | 143.083 | 54
43.493 | 0 | 45.275 | 697 | 53.26 | 86 | 58.276 | 13 | 63.292 | 5 | 73.258 | 0 | 75.238 | 1437 | 93.255 | 43 | 118.268 | 65 | 143.281 | 106
43.691 | 0 | 45.473 | 719 | 53.458 | 124 | 58.474 | 6 | 63.49 | 13 | 73.456 | 0 | 75.436 | 950 | 93.453 | 34 | 118.466 | 60 | 143.479 | 64
43.889 | 0 | 45.671 | 930 | 53.656 | 87 | 58.672 | 8 | 63.688 | 3 | 73.654 | 0 | 75.634 | 475 | 93.651 | 44 | 118.664 | 86 | 143.677 | 95
44.087 | 0 | 45.869 | 999 | 53.854 | 70 | 58.87 | 14 | 63.886 | 0 | 73.852 | 0 | 75.832 | 259 | 93.849 | 41 | 118.862 | 68 | 143.875 | 70
44.285 | 0 | 46.067 | 592 | 54.052 | 61 | 59.068 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 196 | 94.047 | 50 | 119.06 | 38 | 144.073 | 53
44.483 | 0 | 46.265 | 715 | 54.25 | 73 | 59.266 | 12 | 64.282 | 0 | 74.248 | 0 | 76.228 | 52 | 94.245 | 33 | 119.258 | 79 | 144.271 | 68
44.681 | 0 | 46.463 | 598 | 54.448 | 41 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 26 | 94.443 | 11 | 119.456 | 88 | 144.469 | 28
44.879 | 0 | 46.661 | 772 | 54.646 | 69 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 5 | 94.641 | 38 | 119.654 | 47 | 144.667 | 57
45.077 | 0 | 46.859 | 655 | 54.844 | 15 | 59.86 | 10 | 64.876 | 6 | 74.842 | 0 | 76.822 | 3 | 94.839 | 30 | 119.852 | 68 | 144.865 | 48
45.275 | 0 | 47.056 | 526 | 55.042 | 17 | 60.058 | 0 | 65.074 | 6 | 75.04 | 0 | 77.02 | 0 | 95.037 | 48 | 120.05 | 57 | 145.063 | 26
45.473 | 0 | 47.254 | 516 | 55.24 | 39 | 60.256 | 21 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 19 | 120.248 | 24 | 145.261 | 24
45.671 | 0 | 47.452 | 514 | 55.438 | 39 | 60.454 | 3 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 22 | 120.446 | 42 | 145.459 | 6
45.869 | 0 | 47.65 | 552 | 55.636 | 38 | 60.652 | 0 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 9 | 120.644 | 67 | 145.657 | 12
46.067 | 0 | 47.848 | 410 | 55.834 | 28 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 10 | 95.829 | 33 | 120.842 | 23 | 145.855 | 26
46.265 | 0 | 48.046 | 480 | 56.032 | 10 | 61.048 | 6 | 66.064 | 2 | 76.03 | 0 | 78.01 | 0 | 96.027 | 18 | 121.04 | 23 | 146.053 | 7
46.463 | 0 | 48.244 | 473 | 56.23 | 51 | 61.246 | 4 | 66.262 | 3 | 76.228 | 0 | 78.208 | 0 | 96.225 | 5 | 121.238 | 59 | 146.251 | 4
46.661 | 0 | 48.442 | 342 | 56.428 | 22 | 61.444 | 4 | 66.46 | 3 | 76.426 | 0 | 78.405 | 0 | 96.423 | 6 | 121.436 | 10 | 146.449 | 8
46.859 | 0 | 48.64 | 386 | 56.626 | 46 | 61.642 | 5 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 6 | 121.634 | 36 | 146.647 | 4
47.056 | 0 | 48.838 | 323 | 56.824 | 6 | 61.84 | 0 | 66.856 | 5 | 76.822 | 0 | 78.801 | 0 | 96.819 | 26 | 121.832 | 30 | 146.845 | 7
47.254 | 0 | 49.036 | 294 | 57.022 | 49 | 62.038 | 3 | 67.054 | 5 | 77.02 | 0 | 78.999 | 4 | 97.017 | 0 | 122.03 | 18 | 147.043 | 0
47.452 | 0 | 49.234 | 294 | 57.22 | 13 | 62.236 | 3 | 67.252 | 7 | 77.218 | 0 | 79.197 | 0 | 97.215 | 20 | 122.228 | 21 | 147.241 | 5
47.65 | 0 | 49.432 | 218 | 57.418 | 11 | 62.434 | 22 | 67.45 | 6 | 77.416 | 0 | 79.395 | 0 | 97.413 | 7 | 122.426 | 31 | 147.439 | 9
47.848 | 0 | 49.63 | 185 | 57.616 | 17 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 5 | 122.624 | 15 | 147.637 | 11
48.046 | 0 | 49.828 | 224 | 57.814 | 8 | 62.83 | 11 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 7 | 122.822 | 30 | 147.835 | 4
48.244 | 0 | 50.026 | 230 | 58.012 | 11 | 63.028 | 12 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 10 | 123.02 | 31 | 148.033 | 0
48.442 | 0 | 50.224 | 274 | 58.21 | 18 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 28 | 123.218 | 42 | 148.231 | 0
48.64 | 0 | 50.422 | 184 | 58.408 | 30 | 63.424 | 15 | 68.44 | 2 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 55 | 148.429 | 0
48.838 | 0 | 50.62 | 163 | 58.606 | 18 | 63.622 | 0 | 68.638 | 6 | 78.603 | 0 | 80.583 | 0 | 98.601 | 7 | 123.614 | 29 | 148.627 | 6
49.036 | 0 | 50.818 | 140 | 58.804 | 20 | 63.82 | 5 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 12 | 123.812 | 22 | 148.825 | 0
49.234 | 0 | 51.016 | 115 | 59.002 | 0 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 15 | 124.01 | 14 | 149.023 | 8
49.432 | 0 | 51.214 | 134 | 59.2 | 6 | 64.216 | 3 | 69.232 | 17 | 79.197 | 0 | 81.177 | 0 | 99.195 | 19 | 124.208 | 0 | 149.221 | 0
|  | 51.412 | 129 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 21 | 124.406 | 14 | 149.419 | 0

## Analyte Polyalanine n9
0 |  | 2 |  | 10 |  | 15 |  | 30 |  | 32 |  | 50 |  | 75
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0
11.022 | 1073 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0
11.22 | 26298 | 13.002 | 0 | 20.987 | 137 | 26.003 | 47 | 40.985 | 52 | 42.965 | 7 | 60.982 | 32 | 85.995 | 81
11.418 | 86999 | 13.2 | 0 | 21.185 | 2928 | 26.201 | 4765 | 41.183 | 2293 | 43.163 | 1061 | 61.18 | 1338 | 86.193 | 1751
11.616 | 134453 | 13.398 | 0 | 21.383 | 5840 | 26.399 | 24498 | 41.381 | 15439 | 43.361 | 5433 | 61.378 | 3883 | 86.391 | 6764
11.814 | 158253 | 13.596 | 0 | 21.581 | 11954 | 26.597 | 21802 | 41.579 | 15587 | 43.559 | 7005 | 61.576 | 6212 | 86.589 | 7080
12.012 | 410933 | 13.794 | 0 | 21.779 | 15752 | 26.795 | 17702 | 41.777 | 14442 | 43.757 | 7898 | 61.774 | 6665 | 86.787 | 7799
12.21 | 621217 | 13.992 | 0 | 21.977 | 15729 | 26.993 | 14925 | 41.975 | 13526 | 43.955 | 6817 | 61.972 | 6795 | 86.985 | 7548
12.408 | 215124 | 14.19 | 0 | 22.175 | 14297 | 27.191 | 12178 | 42.173 | 14020 | 44.153 | 6213 | 62.17 | 6571 | 87.183 | 7549
12.606 | 15226 | 14.388 | 0 | 22.373 | 14675 | 27.389 | 9858 | 42.371 | 14209 | 44.351 | 6205 | 62.368 | 7386 | 87.381 | 7537
12.804 | 341 | 14.586 | 0 | 22.571 | 15773 | 27.587 | 7636 | 42.569 | 12720 | 44.549 | 7034 | 62.566 | 7695 | 87.579 | 7477
13.002 | 8 | 14.784 | 0 | 22.769 | 17047 | 27.785 | 6149 | 42.767 | 11948 | 44.747 | 7363 | 62.764 | 8354 | 87.777 | 7341
13.2 | 0 | 14.982 | 0 | 22.967 | 17099 | 27.983 | 5801 | 42.965 | 10965 | 44.945 | 7862 | 62.962 | 8316 | 87.975 | 7683
13.398 | 0 | 15.18 | 0 | 23.165 | 16529 | 28.181 | 5810 | 43.163 | 9185 | 45.143 | 8509 | 63.16 | 8269 | 88.173 | 7690
13.596 | 0 | 15.378 | 0 | 23.363 | 15645 | 28.379 | 6115 | 43.361 | 8235 | 45.341 | 8882 | 63.358 | 8145 | 88.371 | 7906
13.794 | 0 | 15.576 | 0 | 23.561 | 17124 | 28.577 | 6922 | 43.559 | 7536 | 45.539 | 9677 | 63.556 | 8178 | 88.569 | 7802
13.992 | 0 | 15.773 | 0 | 23.759 | 16893 | 28.775 | 7756 | 43.757 | 7070 | 45.737 | 9577 | 63.754 | 8082 | 88.767 | 7705
14.19 | 0 | 15.971 | 0 | 23.957 | 16670 | 28.973 | 7326 | 43.955 | 6572 | 45.935 | 9838 | 63.952 | 7880 | 88.965 | 7858
14.388 | 0 | 16.169 | 0 | 24.155 | 15860 | 29.171 | 6474 | 44.153 | 6824 | 46.133 | 8975 | 64.15 | 6885 | 89.163 | 7960
14.586 | 0 | 16.367 | 0 | 24.353 | 16324 | 29.369 | 4745 | 44.351 | 6892 | 46.331 | 9078 | 64.348 | 7467 | 89.361 | 7958
14.784 | 0 | 16.565 | 0 | 24.551 | 17162 | 29.567 | 3881 | 44.549 | 6820 | 46.529 | 8782 | 64.546 | 6587 | 89.559 | 8540
14.982 | 0 | 16.763 | 0 | 24.749 | 17281 | 29.765 | 4321 | 44.747 | 7637 | 46.727 | 7640 | 64.744 | 6557 | 89.757 | 8942
15.18 | 0 | 16.961 | 24 | 24.947 | 16848 | 29.963 | 4914 | 44.945 | 7494 | 46.924 | 7182 | 64.942 | 6381 | 89.955 | 9409
15.378 | 0 | 17.159 | 65 | 25.145 | 18367 | 30.161 | 6827 | 45.143 | 8633 | 47.122 | 6618 | 65.14 | 6291 | 90.153 | 10611
15.576 | 0 | 17.357 | 113 | 25.343 | 18650 | 30.359 | 8265 | 45.341 | 8931 | 47.32 | 6725 | 65.338 | 6295 | 90.351 | 11263
15.773 | 0 | 17.555 | 79 | 25.541 | 19083 | 30.557 | 9793 | 45.539 | 9401 | 47.518 | 6954 | 65.536 | 6354 | 90.549 | 11600
15.971 | 0 | 17.753 | 114 | 25.739 | 20292 | 30.755 | 12291 | 45.737 | 9961 | 47.716 | 6816 | 65.734 | 6384 | 90.747 | 12375
16.169 | 0 | 17.951 | 429 | 25.937 | 21548 | 30.953 | 20591 | 45.935 | 10023 | 47.914 | 7312 | 65.932 | 6912 | 90.945 | 14271
16.367 | 0 | 18.149 | 436 | 26.135 | 21694 | 31.151 | 37575 | 46.133 | 9664 | 48.112 | 7763 | 66.13 | 6774 | 91.143 | 14734
16.565 | 0 | 18.347 | 679 | 26.333 | 22153 | 31.349 | 69629 | 46.331 | 8939 | 48.31 | 8162 | 66.328 | 6944 | 91.341 | 15977
16.763 | 0 | 18.545 | 1727 | 26.531 | 21712 | 31.547 | 114237 | 46.529 | 9040 | 48.508 | 8035 | 66.526 | 6899 | 91.539 | 16799
16.961 | 0 | 18.743 | 3598 | 26.729 | 19120 | 31.745 | 163986 | 46.727 | 7760 | 48.706 | 8668 | 66.724 | 7685 | 91.737 | 18339
17.159 | 0 | 18.941 | 3184 | 26.927 | 17264 | 31.943 | 209574 | 46.924 | 7569 | 48.904 | 8939 | 66.922 | 7908 | 91.935 | 19610
17.357 | 0 | 19.139 | 2047 | 27.125 | 15362 | 32.141 | 229258 | 47.122 | 7036 | 49.102 | 9237 | 67.12 | 8117 | 92.133 | 22146
17.555 | 0 | 19.337 | 2138 | 27.323 | 12545 | 32.339 | 210521 | 47.32 | 7100 | 49.3 | 9494 | 67.318 | 7635 | 92.331 | 23891
17.753 | 0 | 19.535 | 3857 | 27.521 | 9798 | 32.537 | 167900 | 47.518 | 6813 | 49.498 | 9205 | 67.516 | 8154 | 92.529 | 25371
17.951 | 0 | 19.733 | 3983 | 27.719 | 7725 | 32.735 | 120425 | 47.716 | 6984 | 49.696 | 9905 | 67.714 | 8733 | 92.727 | 27854
18.149 | 0 | 19.931 | 3273 | 27.917 | 6200 | 32.933 | 78618 | 47.914 | 7733 | 49.894 | 9948 | 67.912 | 8839 | 92.925 | 29424
18.347 | 0 | 20.129 | 4132 | 28.115 | 5692 | 33.131 | 50158 | 48.112 | 7975 | 50.092 | 10455 | 68.11 | 8450 | 93.123 | 30266
18.545 | 0 | 20.327 | 7235 | 28.313 | 5915 | 33.329 | 31508 | 48.31 | 7801 | 50.29 | 11520 | 68.308 | 8817 | 93.321 | 31824
18.743 | 0 | 20.525 | 8152 | 28.511 | 6611 | 33.527 | 20516 | 48.508 | 8257 | 50.488 | 14215 | 68.506 | 9143 | 93.519 | 31425
18.941 | 0 | 20.723 | 7658 | 28.709 | 7361 | 33.725 | 13685 | 48.706 | 8470 | 50.686 | 18297 | 68.704 | 8891 | 93.717 | 29579
19.139 | 0 | 20.921 | 8982 | 28.907 | 8136 | 33.923 | 10961 | 48.904 | 8730 | 50.884 | 22860 | 68.902 | 8988 | 93.915 | 29066
19.337 | 0 | 21.119 | 11835 | 29.105 | 7448 | 34.121 | 9594 | 49.102 | 9018 | 51.082 | 29645 | 69.1 | 9129 | 94.113 | 26175
19.535 | 0 | 21.317 | 14062 | 29.303 | 6668 | 34.319 | 9009 | 49.3 | 9207 | 51.28 | 37954 | 69.298 | 9013 | 94.311 | 24123
19.733 | 0 | 21.515 | 13236 | 29.501 | 5171 | 34.517 | 8699 | 49.498 | 9035 | 51.478 | 47233 | 69.496 | 9017 | 94.509 | 21772
19.931 | 0 | 21.713 | 12651 | 29.699 | 3985 | 34.715 | 8347 | 49.696 | 9198 | 51.676 | 59619 | 69.694 | 9267 | 94.707 | 19777
20.129 | 0 | 21.911 | 14452 | 29.897 | 4199 | 34.913 | 7857 | 49.894 | 9986 | 51.874 | 72219 | 69.892 | 9834 | 94.905 | 18580
20.327 | 0 | 22.109 | 15394 | 30.095 | 5452 | 35.111 | 7482 | 50.092 | 9278 | 52.072 | 86129 | 70.09 | 10796 | 95.103 | 16485
20.525 | 0 | 22.307 | 15389 | 30.293 | 6337 | 35.309 | 7714 | 50.29 | 11460 | 52.27 | 98897 | 70.288 | 12480 | 95.301 | 15391
20.723 | 0 | 22.505 | 13959 | 30.491 | 8611 | 35.507 | 7609 | 50.488 | 12860 | 52.468 | 107374 | 70.486 | 14767 | 95.499 | 14098
20.921 | 0 | 22.703 | 14676 | 30.689 | 10145 | 35.705 | 8496 | 50.686 | 17436 | 52.666 | 110862 | 70.684 | 16921 | 95.697 | 13819
21.119 | 0 | 22.901 | 16121 | 30.887 | 13328 | 35.903 | 8907 | 50.884 | 22088 | 52.864 | 107015 | 70.882 | 20526 | 95.895 | 12698
21.317 | 0 | 23.099 | 16764 | 31.085 | 21267 | 36.101 | 9139 | 51.082 | 28998 | 53.062 | 94580 | 71.08 | 22230 | 96.093 | 12026
21.515 | 0 | 23.297 | 16861 | 31.283 | 39702 | 36.299 | 9158 | 51.28 | 36883 | 53.26 | 80610 | 71.278 | 26125 | 96.291 | 10615
21.713 | 0 | 23.495 | 16302 | 31.481 | 72031 | 36.497 | 8422 | 51.478 | 47148 | 53.458 | 64688 | 71.476 | 29108 | 96.489 | 9925
21.911 | 0 | 23.693 | 16702 | 31.679 | 119753 | 36.695 | 7896 | 51.676 | 58236 | 53.656 | 54515 | 71.674 | 32461 | 96.687 | 9183
22.109 | 0 | 23.891 | 17463 | 31.877 | 175301 | 36.893 | 7561 | 51.874 | 70978 | 53.854 | 44412 | 71.872 | 35148 | 96.885 | 8928
22.307 | 0 | 24.089 | 16747 | 32.075 | 221415 | 37.091 | 7250 | 52.072 | 86965 | 54.052 | 35664 | 72.07 | 41238 | 97.083 | 8000
22.505 | 0 | 24.287 | 16357 | 32.273 | 240850 | 37.289 | 7482 | 52.27 | 99332 | 54.25 | 29772 | 72.268 | 46255 | 97.281 | 7321
22.703 | 0 | 24.485 | 16746 | 32.471 | 218828 | 37.487 | 8261 | 52.468 | 110635 | 54.448 | 24184 | 72.466 | 51431 | 97.479 | 6486
22.901 | 0 | 24.683 | 16179 | 32.669 | 172055 | 37.685 | 8550 | 52.666 | 114364 | 54.646 | 18789 | 72.664 | 55663 | 97.677 | 6187
23.099 | 0 | 24.881 | 16985 | 32.867 | 119243 | 37.883 | 8998 | 52.864 | 112767 | 54.844 | 15361 | 72.862 | 58474 | 97.875 | 5787
23.297 | 0 | 25.079 | 16526 | 33.065 | 75677 | 38.081 | 8647 | 53.062 | 102449 | 55.042 | 12652 | 73.06 | 59687 | 98.073 | 5562
23.495 | 0 | 25.277 | 16148 | 33.263 | 48063 | 38.279 | 8017 | 53.26 | 86775 | 55.24 | 11339 | 73.258 | 59641 | 98.271 | 5702
23.693 | 0 | 25.475 | 17870 | 33.461 | 27215 | 38.477 | 7462 | 53.458 | 71587 | 55.438 | 9892 | 73.456 | 54988 | 98.469 | 4912
23.891 | 0 | 25.673 | 17886 | 33.659 | 14542 | 38.675 | 7863 | 53.656 | 58255 | 55.636 | 9880 | 73.654 | 50095 | 98.667 | 5044
24.089 | 0 | 25.871 | 18588 | 33.857 | 7236 | 38.873 | 7532 | 53.854 | 45783 | 55.834 | 8272 | 73.852 | 45316 | 98.865 | 4396
24.287 | 0 | 26.069 | 18759 | 34.055 | 3103 | 39.071 | 8246 | 54.052 | 37589 | 56.032 | 8384 | 74.05 | 37421 | 99.063 | 3912
24.485 | 0 | 26.267 | 20918 | 34.253 | 1497 | 39.269 | 8615 | 54.25 | 30752 | 56.23 | 7034 | 74.248 | 31773 | 99.261 | 3578
24.683 | 0 | 26.465 | 21786 | 34.451 | 582 | 39.467 | 8780 | 54.448 | 24901 | 56.428 | 7778 | 74.446 | 26852 | 99.459 | 3436
24.881 | 0 | 26.663 | 21475 | 34.649 | 476 | 39.665 | 8072 | 54.646 | 20663 | 56.626 | 7472 | 74.644 | 22352 | 99.657 | 3255
25.079 | 0 | 26.861 | 20148 | 34.847 | 480 | 39.863 | 8041 | 54.844 | 17119 | 56.824 | 7150 | 74.842 | 19846 | 99.855 | 2850
25.277 | 0 | 27.059 | 17356 | 35.045 | 312 | 40.061 | 8050 | 55.042 | 14343 | 57.022 | 6890 | 75.04 | 17147 | 100.053 | 2999
25.475 | 0 | 27.257 | 15017 | 35.243 | 395 | 40.259 | 8191 | 55.24 | 12138 | 57.22 | 6360 | 75.238 | 15606 | 100.251 | 3054
25.673 | 0 | 27.455 | 11992 | 35.441 | 340 | 40.457 | 8252 | 55.438 | 10243 | 57.418 | 6401 | 75.436 | 13066 | 100.449 | 2947
25.871 | 0 | 27.653 | 9091 | 35.639 | 328 | 40.655 | 9471 | 55.636 | 9258 | 57.616 | 5962 | 75.634 | 11385 | 100.647 | 3243
26.069 | 0 | 27.851 | 7282 | 35.837 | 277 | 40.853 | 9205 | 55.834 | 8609 | 57.814 | 6138 | 75.832 | 9717 | 100.845 | 2883
26.267 | 0 | 28.049 | 6466 | 36.035 | 339 | 41.051 | 9277 | 56.032 | 8067 | 58.012 | 5049 | 76.03 | 8186 | 101.043 | 3426
26.465 | 0 | 28.247 | 5633 | 36.233 | 353 | 41.249 | 7748 | 56.23 | 7118 | 58.21 | 5409 | 76.228 | 6821 | 101.241 | 2635
26.663 | 0 | 28.445 | 5081 | 36.431 | 258 | 41.447 | 5304 | 56.428 | 4696 | 58.408 | 5099 | 76.426 | 6050 | 101.439 | 1962
26.861 | 0 | 28.643 | 4278 | 36.629 | 328 | 41.645 | 3338 | 56.626 | 2808 | 58.606 | 4805 | 76.624 | 4938 | 101.637 | 1297
27.059 | 0 | 28.841 | 4041 | 36.827 | 271 | 41.843 | 1251 | 56.824 | 1339 | 58.804 | 4252 | 76.822 | 3658 | 101.835 | 970
27.257 | 0 | 29.039 | 3458 | 37.025 | 355 | 42.041 | 534 | 57.022 | 452 | 59.002 | 3809 | 77.02 | 3519 | 102.033 | 587
27.455 | 0 | 29.237 | 3610 | 37.223 | 348 | 42.239 | 270 | 57.22 | 358 | 59.2 | 3325 | 77.218 | 2930 | 102.231 | 518
27.653 | 0 | 29.435 | 3111 | 37.421 | 274 | 42.437 | 55 | 57.418 | 153 | 59.398 | 2364 | 77.416 | 2601 | 102.429 | 414
27.851 | 0 | 29.633 | 2936 | 37.619 | 288 | 42.635 | 29 | 57.616 | 210 | 59.596 | 1610 | 77.614 | 1969 | 102.627 | 446
28.049 | 0 | 29.831 | 2162 | 37.817 | 199 | 42.833 | 36 | 57.814 | 199 | 59.794 | 1023 | 77.812 | 1605 | 102.825 | 433
28.247 | 0 | 30.029 | 2264 | 38.015 | 216 | 43.031 | 31 | 58.012 | 155 | 59.992 | 444 | 78.01 | 1508 | 103.023 | 447
28.445 | 0 | 30.227 | 1821 | 38.213 | 178 | 43.229 | 18 | 58.21 | 147 | 60.19 | 210 | 78.208 | 881 | 103.221 | 436
28.643 | 0 | 30.425 | 1783 | 38.411 | 158 | 43.427 | 43 | 58.408 | 203 | 60.388 | 179 | 78.405 | 575 | 103.419 | 343
28.841 | 0 | 30.623 | 1885 | 38.609 | 159 | 43.625 | 24 | 58.606 | 202 | 60.586 | 183 | 78.603 | 371 | 103.617 | 364
29.039 | 0 | 30.821 | 2200 | 38.807 | 165 | 43.823 | 36 | 58.804 | 152 | 60.784 | 178 | 78.801 | 246 | 103.815 | 477
29.237 | 0 | 31.019 | 4474 | 39.005 | 138 | 44.021 | 27 | 59.002 | 161 | 60.982 | 132 | 78.999 | 221 | 104.013 | 343
29.435 | 0 | 31.217 | 10442 | 39.203 | 77 | 44.219 | 27 | 59.2 | 169 | 61.18 | 98 | 79.197 | 156 | 104.211 | 390
29.633 | 0 | 31.415 | 26428 | 39.401 | 160 | 44.417 | 17 | 59.398 | 177 | 61.378 | 99 | 79.395 | 123 | 104.409 | 364
29.831 | 0 | 31.613 | 56042 | 39.599 | 39 | 44.615 | 33 | 59.596 | 187 | 61.576 | 158 | 79.593 | 168 | 104.607 | 364
30.029 | 0 | 31.811 | 98860 | 39.797 | 86 | 44.813 | 9 | 59.794 | 152 | 61.774 | 148 | 79.791 | 137 | 104.805 | 351
30.227 | 0 | 32.009 | 155535 | 39.995 | 48 | 45.011 | 3 | 59.992 | 163 | 61.972 | 104 | 79.989 | 83 | 105.003 | 353
30.425 | 0 | 32.207 | 210247 | 40.193 | 56 | 45.209 | 40 | 60.19 | 173 | 62.17 | 116 | 80.187 | 142 | 105.201 | 281
30.623 | 0 | 32.405 | 247618 | 40.391 | 86 | 45.407 | 18 | 60.388 | 69 | 62.368 | 104 | 80.385 | 112 | 105.399 | 303
30.821 | 0 | 32.603 | 242680 | 40.589 | 52 | 45.605 | 30 | 60.586 | 194 | 62.566 | 207 | 80.583 | 168 | 105.597 | 168
31.019 | 0 | 32.801 | 203458 | 40.787 | 32 | 45.803 | 9 | 60.784 | 116 | 62.764 | 48 | 80.781 | 115 | 105.795 | 185
31.217 | 0 | 32.999 | 146699 | 40.985 | 124 | 46.001 | 31 | 60.982 | 122 | 62.962 | 133 | 80.979 | 167 | 105.993 | 112
31.415 | 0 | 33.197 | 98133 | 41.183 | 76 | 46.199 | 12 | 61.18 | 156 | 63.16 | 132 | 81.177 | 118 | 106.191 | 109
31.613 | 0 | 33.395 | 61654 | 41.381 | 50 | 46.397 | 14 | 61.378 | 81 | 63.358 | 170 | 81.375 | 100 | 106.389 | 75
31.811 | 0 | 33.593 | 36083 | 41.579 | 19 | 46.595 | 17 | 61.576 | 181 | 63.556 | 67 | 81.573 | 120 | 106.587 | 99
32.009 | 0 | 33.791 | 19924 | 41.777 | 33 | 46.793 | 18 | 61.774 | 91 | 63.754 | 120 | 81.771 | 48 | 106.785 | 88
32.207 | 0 | 33.989 | 10263 | 41.975 | 68 | 46.99 | 23 | 61.972 | 136 | 63.952 | 75 | 81.969 | 93 | 106.983 | 84
32.405 | 0 | 34.187 | 4965 | 42.173 | 29 | 47.188 | 13 | 62.17 | 121 | 64.15 | 46 | 82.167 | 114 | 107.181 | 52
32.603 | 0 | 34.385 | 2054 | 42.371 | 41 | 47.386 | 6 | 62.368 | 148 | 64.348 | 42 | 82.365 | 65 | 107.379 | 62
32.801 | 0 | 34.583 | 893 | 42.569 | 12 | 47.584 | 3 | 62.566 | 129 | 64.546 | 73 | 82.563 | 77 | 107.577 | 16
32.999 | 0 | 34.781 | 582 | 42.767 | 10 | 47.782 | 0 | 62.764 | 129 | 64.744 | 66 | 82.761 | 89 | 107.775 | 11
33.197 | 0 | 34.979 | 409 | 42.965 | 7 | 47.98 | 0 | 62.962 | 88 | 64.942 | 44 | 82.959 | 102 | 107.973 | 67
33.395 | 0 | 35.177 | 387 | 43.163 | 3 | 48.178 | 0 | 63.16 | 123 | 65.14 | 88 | 83.157 | 59 | 108.171 | 40
33.593 | 0 | 35.375 | 415 | 43.361 | 0 | 48.376 | 0 | 63.358 | 90 | 65.338 | 49 | 83.355 | 95 | 108.369 | 65
33.791 | 0 | 35.573 | 379 | 43.559 | 4 | 48.574 | 0 | 63.556 | 79 | 65.536 | 21 | 83.553 | 143 | 108.567 | 34
33.989 | 0 | 35.771 | 398 | 43.757 | 0 | 48.772 | 11 | 63.754 | 30 | 65.734 | 18 | 83.751 | 40 | 108.765 | 8
34.187 | 0 | 35.969 | 258 | 43.955 | 6 | 48.97 | 2 | 63.952 | 55 | 65.932 | 36 | 83.949 | 76 | 108.963 | 9
34.385 | 0 | 36.167 | 265 | 44.153 | 9 | 49.168 | 0 | 64.15 | 33 | 66.13 | 38 | 84.147 | 86 | 109.161 | 17
34.583 | 0 | 36.365 | 263 | 44.351 | 0 | 49.366 | 0 | 64.348 | 56 | 66.328 | 49 | 84.345 | 71 | 109.359 | 26
34.781 | 0 | 36.563 | 342 | 44.549 | 16 | 49.564 | 0 | 64.546 | 20 | 66.526 | 39 | 84.543 | 71 | 109.556 | 23
34.979 | 0 | 36.761 | 271 | 44.747 | 5 | 49.762 | 6 | 64.744 | 74 | 66.724 | 18 | 84.741 | 41 | 109.754 | 15
35.177 | 0 | 36.959 | 239 | 44.945 | 0 | 49.96 | 4 | 64.942 | 34 | 66.922 | 35 | 84.939 | 65 | 109.952 | 32
35.375 | 0 | 37.157 | 361 | 45.143 | 0 | 50.158 | 0 | 65.14 | 74 | 67.12 | 39 | 85.137 | 68 | 110.15 | 14
35.573 | 0 | 37.355 | 359 | 45.341 | 11 | 50.356 | 4 | 65.338 | 35 | 67.318 | 7 | 85.335 | 65 | 110.348 | 41
35.771 | 0 | 37.553 | 300 | 45.539 | 0 | 50.554 | 9 | 65.536 | 100 | 67.516 | 23 | 85.533 | 63 | 110.546 | 21
35.969 | 0 | 37.751 | 245 | 45.737 | 5 | 50.752 | 7 | 65.734 | 76 | 67.714 | 39 | 85.731 | 57 | 110.744 | 12
36.167 | 0 | 37.949 | 212 | 45.935 | 3 | 50.95 | 3 | 65.932 | 75 | 67.912 | 28 | 85.929 | 83 | 110.942 | 30
36.365 | 0 | 38.147 | 194 | 46.133 | 7 | 51.148 | 6 | 66.13 | 33 | 68.11 | 11 | 86.127 | 28 | 111.14 | 31
36.563 | 0 | 38.345 | 221 | 46.331 | 2 | 51.346 | 5 | 66.328 | 17 | 68.308 | 49 | 86.325 | 19 | 111.338 | 27
36.761 | 0 | 38.543 | 147 | 46.529 | 0 | 51.544 | 2 | 66.526 | 26 | 68.506 | 12 | 86.523 | 52 | 111.536 | 19
36.959 | 0 | 38.741 | 134 | 46.727 | 10 | 51.742 | 0 | 66.724 | 20 | 68.704 | 24 | 86.721 | 17 | 111.734 | 22
37.157 | 0 | 38.939 | 102 | 46.924 | 7 | 51.94 | 3 | 66.922 | 27 | 68.902 | 23 | 86.919 | 22 | 111.932 | 13
37.355 | 0 | 39.137 | 123 | 47.122 | 5 | 52.138 | 0 | 67.12 | 11 | 69.1 | 5 | 87.117 | 53 | 112.13 | 0
37.553 | 0 | 39.335 | 95 | 47.32 | 0 | 52.336 | 0 | 67.318 | 37 | 69.298 | 7 | 87.315 | 17 | 112.328 | 0
37.751 | 0 | 39.533 | 88 | 47.518 | 0 | 52.534 | 0 | 67.516 | 45 | 69.496 | 13 | 87.513 | 25 | 112.526 | 0
37.949 | 0 | 39.731 | 44 | 47.716 | 0 | 52.732 | 2 | 67.714 | 41 | 69.694 | 7 | 87.711 | 5 | 112.724 | 0
38.147 | 0 | 39.929 | 57 | 47.914 | 7 | 52.93 | 0 | 67.912 | 55 | 69.892 | 16 | 87.909 | 0 | 112.922 | 8
38.345 | 0 | 40.127 | 107 | 48.112 | 8 | 53.128 | 0 | 68.11 | 9 | 70.09 | 3 | 88.107 | 3 | 113.12 | 0
38.543 | 0 | 40.325 | 54 | 48.31 | 0 | 53.326 | 0 | 68.308 | 10 | 70.288 | 16 | 88.305 | 5 | 113.318 | 0
38.741 | 0 | 40.523 | 38 | 48.508 | 15 | 53.524 | 6 | 68.506 | 27 | 70.486 | 5 | 88.503 | 6 | 113.516 | 0
38.939 | 0 | 40.721 | 18 | 48.706 | 0 | 53.722 | 0 | 68.704 | 14 | 70.684 | 5 | 88.701 | 0 | 113.714 | 0
39.137 | 0 | 40.919 | 97 | 48.904 | 3 | 53.92 | 0 | 68.902 | 24 | 70.882 | 0 | 88.899 | 0 | 113.912 | 7
39.335 | 0 | 41.117 | 49 | 49.102 | 2 | 54.118 | 0 | 69.1 | 34 | 71.08 | 5 | 89.097 | 0 | 114.11 | 0
39.533 | 0 | 41.315 | 60 | 49.3 | 0 | 54.316 | 5 | 69.298 | 19 | 71.278 | 11 | 89.295 | 0 | 114.308 | 0
39.731 | 0 | 41.513 | 44 | 49.498 | 0 | 54.514 | 0 | 69.496 | 17 | 71.476 | 0 | 89.493 | 0 | 114.506 | 0
39.929 | 0 | 41.711 | 59 | 49.696 | 9 | 54.712 | 0 | 69.694 | 15 | 71.674 | 25 | 89.691 | 6 | 114.704 | 0
40.127 | 0 | 41.909 | 33 | 49.894 | 8 | 54.91 | 0 | 69.892 | 20 | 71.872 | 0 | 89.889 | 4 | 114.902 | 0
40.325 | 0 | 42.107 | 15 | 50.092 | 3 | 55.108 | 4 | 70.09 | 11 | 72.07 | 7 | 90.087 | 0 | 115.1 | 0
40.523 | 0 | 42.305 | 42 | 50.29 | 3 | 55.306 | 3 | 70.288 | 13 | 72.268 | 3 | 90.285 | 0 | 115.298 | 0
40.721 | 0 | 42.503 | 3 | 50.488 | 3 | 55.504 | 0 | 70.486 | 24 | 72.466 | 0 | 90.483 | 0 | 115.496 | 4
40.919 | 0 | 42.701 | 14 | 50.686 | 6 | 55.702 | 0 | 70.684 | 16 | 72.664 | 5 | 90.681 | 0 | 115.694 | 0
41.117 | 0 | 42.899 | 13 | 50.884 | 0 | 55.9 | 0 | 70.882 | 4 | 72.862 | 0 | 90.879 | 0 | 115.892 | 0
41.315 | 0 | 43.097 | 0 | 51.082 | 0 | 56.098 | 0 | 71.08 | 0 | 73.06 | 0 | 91.077 | 0 | 116.09 | 0
41.513 | 0 | 43.295 | 4 | 51.28 | 0 | 56.296 | 0 | 71.278 | 18 | 73.258 | 0 | 91.275 | 0 | 116.288 | 0
41.711 | 0 | 43.493 | 16 | 51.478 | 13 | 56.494 | 0 | 71.476 | 7 | 73.456 | 15 | 91.473 | 0 | 116.486 | 0
41.909 | 0 | 43.691 | 23 | 51.676 | 0 | 56.692 | 0 | 71.674 | 0 | 73.654 | 7 | 91.671 | 0 | 116.684 | 0
42.107 | 0 | 43.889 | 11 | 51.874 | 0 | 56.89 | 0 | 71.872 | 0 | 73.852 | 5 | 91.869 | 0 | 116.882 | 0
42.305 | 0 | 44.087 | 7 | 52.072 | 0 | 57.088 | 0 | 72.07 | 0 | 74.05 | 0 | 92.067 | 0 | 117.08 | 0
42.503 | 0 | 44.285 | 0 | 52.27 | 0 | 57.286 | 0 | 72.268 | 0 | 74.248 | 16 | 92.265 | 0 | 117.278 | 0
42.701 | 0 | 44.483 | 0 | 52.468 | 0 | 57.484 | 0 | 72.466 | 0 | 74.446 | 0 | 92.463 | 0 | 117.476 | 0
42.899 | 0 | 44.681 | 0 | 52.666 | 11 | 57.682 | 0 | 72.664 | 0 | 74.644 | 7 | 92.661 | 0 | 117.674 | 0
43.097 | 0 | 44.879 | 3 | 52.864 | 0 | 57.88 | 0 | 72.862 | 9 | 74.842 | 10 | 92.859 | 0 | 117.872 | 0
43.295 | 0 | 45.077 | 5 | 53.062 | 10 | 58.078 | 9 | 73.06 | 0 | 75.04 | 4 | 93.057 | 0 | 118.07 | 0
43.493 | 0 | 45.275 | 0 | 53.26 | 8 | 58.276 | 0 | 73.258 | 0 | 75.238 | 0 | 93.255 | 0 | 118.268 | 0
43.691 | 0 | 45.473 | 10 | 53.458 | 0 | 58.474 | 4 | 73.456 | 0 | 75.436 | 4 | 93.453 | 9 | 118.466 | 0
43.889 | 0 | 45.671 | 0 | 53.656 | 0 | 58.672 | 0 | 73.654 | 0 | 75.634 | 0 | 93.651 | 0 | 118.664 | 0
44.087 | 0 | 45.869 | 15 | 53.854 | 4 | 58.87 | 0 | 73.852 | 0 | 75.832 | 0 | 93.849 | 0 | 118.862 | 0
44.285 | 0 | 46.067 | 0 | 54.052 | 0 | 59.068 | 0 | 74.05 | 0 | 76.03 | 0 | 94.047 | 0 | 119.06 | 0
44.483 | 0 | 46.265 | 0 | 54.25 | 0 | 59.266 | 0 | 74.248 | 0 | 76.228 | 0 | 94.245 | 0 | 119.258 | 0
44.681 | 0 | 46.463 | 10 | 54.448 | 0 | 59.464 | 0 | 74.446 | 0 | 76.426 | 0 | 94.443 | 0 | 119.456 | 0
44.879 | 0 | 46.661 | 0 | 54.646 | 0 | 59.662 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 0 | 119.654 | 0
45.077 | 0 | 46.859 | 0 | 54.844 | 0 | 59.86 | 0 | 74.842 | 0 | 76.822 | 0 | 94.839 | 4 | 119.852 | 0
45.275 | 0 | 47.056 | 0 | 55.042 | 0 | 60.058 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 0 | 120.05 | 0
45.473 | 0 | 47.254 | 0 | 55.24 | 0 | 60.256 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 0 | 120.248 | 0
45.671 | 0 | 47.452 | 0 | 55.438 | 0 | 60.454 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 0
45.869 | 0 | 47.65 | 0 | 55.636 | 0 | 60.652 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 0
46.067 | 0 | 47.848 | 11 | 55.834 | 0 | 60.85 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 0 | 120.842 | 0
46.265 | 0 | 48.046 | 0 | 56.032 | 0 | 61.048 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 0 | 121.04 | 0
46.463 | 0 | 48.244 | 0 | 56.23 | 0 | 61.246 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 0
46.661 | 0 | 48.442 | 2 | 56.428 | 0 | 61.444 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 0
46.859 | 0 | 48.64 | 2 | 56.626 | 0 | 61.642 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 0
47.056 | 0 | 48.838 | 0 | 56.824 | 0 | 61.84 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 0
47.254 | 0 | 49.036 | 0 | 57.022 | 5 | 62.038 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 0 | 122.03 | 0
47.452 | 0 | 49.234 | 0 | 57.22 | 0 | 62.236 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 5
47.65 | 0 | 49.432 | 0 | 57.418 | 0 | 62.434 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 0
47.848 | 0 | 49.63 | 5 | 57.616 | 0 | 62.632 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 0
48.046 | 0 | 49.828 | 3 | 57.814 | 0 | 62.83 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 0 | 122.822 | 0
48.244 | 0 | 50.026 | 14 | 58.012 | 0 | 63.028 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 0
48.442 | 0 | 50.224 | 5 | 58.21 | 0 | 63.226 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 0
48.64 | 0 | 50.422 | 5 | 58.408 | 0 | 63.424 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 7
48.838 | 0 | 50.62 | 0 | 58.606 | 0 | 63.622 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 0 | 123.614 | 0
49.036 | 0 | 50.818 | 10 | 58.804 | 0 | 63.82 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0
49.234 | 0 | 51.016 | 10 | 59.002 | 0 | 64.018 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 0 | 124.01 | 0
49.432 | 0 | 51.214 | 0 | 59.2 | 0 | 64.216 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 0 | 124.208 | 0
|  | 51.412 | 12 | 59.398 | 0 | 64.414 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 0

## Analyte Polyalanine n10
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 364 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 10386 | 13.002 | 0 | 20.987 | 42 | 26.003 | 0 | 31.019 | 64 | 40.985 | 27 | 42.965 | 3 | 60.982 | 7 | 85.995 | 15 | 111.008 | 17
11.418 | 40156 | 13.2 | 0 | 21.185 | 1193 | 26.201 | 930 | 31.217 | 936 | 41.183 | 664 | 43.163 | 501 | 61.18 | 665 | 86.193 | 620 | 111.206 | 581
11.616 | 63631 | 13.398 | 0 | 21.383 | 2915 | 26.399 | 6949 | 31.415 | 1600 | 41.381 | 4173 | 43.361 | 3480 | 61.378 | 3444 | 86.391 | 2450 | 111.404 | 2176
11.814 | 87152 | 13.596 | 0 | 21.581 | 5746 | 26.597 | 8520 | 31.613 | 2962 | 41.579 | 5312 | 43.559 | 5350 | 61.576 | 4881 | 86.589 | 3404 | 111.602 | 2891
12.012 | 147007 | 13.794 | 0 | 21.779 | 8606 | 26.795 | 9753 | 31.811 | 4795 | 41.777 | 7316 | 43.757 | 6734 | 61.774 | 5601 | 86.787 | 4365 | 111.8 | 3695
12.21 | 513444 | 13.992 | 0 | 21.977 | 9297 | 26.993 | 9814 | 32.009 | 5079 | 41.975 | 7580 | 43.955 | 5804 | 61.972 | 6133 | 86.985 | 4128 | 111.998 | 3440
12.408 | 484432 | 14.19 | 0 | 22.175 | 9243 | 27.191 | 9275 | 32.207 | 5275 | 42.173 | 8127 | 44.153 | 5227 | 62.17 | 6542 | 87.183 | 3962 | 112.196 | 3296
12.606 | 78061 | 14.388 | 0 | 22.373 | 8680 | 27.389 | 8971 | 32.405 | 5598 | 42.371 | 8634 | 44.351 | 5892 | 62.368 | 6862 | 87.381 | 3759 | 112.394 | 3394
12.804 | 2290 | 14.586 | 0 | 22.571 | 9226 | 27.587 | 8253 | 32.603 | 4887 | 42.569 | 8796 | 44.549 | 6318 | 62.566 | 6934 | 87.579 | 3899 | 112.592 | 3453
13.002 | 61 | 14.784 | 0 | 22.769 | 9868 | 27.785 | 7796 | 32.801 | 4755 | 42.767 | 8300 | 44.747 | 6522 | 62.764 | 7255 | 87.777 | 4097 | 112.79 | 3661
13.2 | 14 | 14.982 | 0 | 22.967 | 10474 | 27.983 | 8312 | 32.999 | 5510 | 42.965 | 7840 | 44.945 | 6867 | 62.962 | 7745 | 87.975 | 4381 | 112.988 | 3592
13.398 | 1 | 15.18 | 0 | 23.165 | 10093 | 28.181 | 8962 | 33.197 | 7742 | 43.163 | 7219 | 45.143 | 7811 | 63.16 | 8387 | 88.173 | 3931 | 113.186 | 4141
13.596 | 0 | 15.378 | 0 | 23.363 | 10310 | 28.379 | 10284 | 33.395 | 13790 | 43.361 | 6695 | 45.341 | 8572 | 63.358 | 8230 | 88.371 | 4276 | 113.384 | 4042
13.794 | 0 | 15.576 | 0 | 23.561 | 10527 | 28.577 | 11865 | 33.593 | 26712 | 43.559 | 6341 | 45.539 | 9057 | 63.556 | 7782 | 88.569 | 4172 | 113.582 | 3995
13.992 | 0 | 15.773 | 0 | 23.759 | 11063 | 28.775 | 12780 | 33.791 | 45890 | 43.757 | 5581 | 45.737 | 9171 | 63.754 | 8342 | 88.767 | 4394 | 113.78 | 3831
14.19 | 0 | 15.971 | 0 | 23.957 | 10568 | 28.973 | 12503 | 33.989 | 73216 | 43.955 | 5493 | 45.935 | 10253 | 63.952 | 7919 | 88.965 | 4237 | 113.978 | 4239
14.388 | 0 | 16.169 | 0 | 24.155 | 10625 | 29.171 | 10828 | 34.187 | 104835 | 44.153 | 5385 | 46.133 | 9584 | 64.15 | 7439 | 89.163 | 4297 | 114.176 | 4135
14.586 | 0 | 16.367 | 0 | 24.353 | 11771 | 29.369 | 8365 | 34.385 | 135796 | 44.351 | 5771 | 46.331 | 9657 | 64.348 | 7380 | 89.361 | 4441 | 114.374 | 4678
14.784 | 0 | 16.565 | 0 | 24.551 | 11720 | 29.567 | 7006 | 34.583 | 156976 | 44.549 | 6148 | 46.529 | 8980 | 64.546 | 7456 | 89.559 | 4018 | 114.572 | 4182
14.982 | 0 | 16.763 | 0 | 24.749 | 11629 | 29.765 | 5304 | 34.781 | 159010 | 44.747 | 6867 | 46.727 | 8055 | 64.744 | 6595 | 89.757 | 4395 | 114.77 | 4193
15.18 | 0 | 16.961 | 0 | 24.947 | 10960 | 29.963 | 4474 | 34.979 | 138807 | 44.945 | 7399 | 46.924 | 8342 | 64.942 | 6198 | 89.955 | 4006 | 114.968 | 4485
15.378 | 0 | 17.159 | 9 | 25.145 | 10319 | 30.161 | 4387 | 35.177 | 107850 | 45.143 | 8273 | 47.122 | 7480 | 65.14 | 6132 | 90.153 | 4315 | 115.166 | 4551
15.576 | 0 | 17.357 | 5 | 25.343 | 9685 | 30.359 | 4935 | 35.375 | 72548 | 45.341 | 7951 | 47.32 | 6692 | 65.338 | 5578 | 90.351 | 3922 | 115.364 | 4411
15.773 | 0 | 17.555 | 0 | 25.541 | 9774 | 30.557 | 5064 | 35.573 | 47572 | 45.539 | 9314 | 47.518 | 6771 | 65.536 | 5633 | 90.549 | 3714 | 115.562 | 4583
15.971 | 0 | 17.753 | 49 | 25.739 | 9210 | 30.755 | 4984 | 35.771 | 30460 | 45.737 | 9660 | 47.716 | 6879 | 65.734 | 5226 | 90.747 | 3655 | 115.76 | 4603
16.169 | 0 | 17.951 | 200 | 25.937 | 9433 | 30.953 | 4356 | 35.969 | 18743 | 45.935 | 10321 | 47.914 | 6396 | 65.932 | 5339 | 90.945 | 4024 | 115.958 | 4272
16.367 | 0 | 18.149 | 433 | 26.135 | 9243 | 31.151 | 3954 | 36.167 | 12314 | 46.133 | 9684 | 48.112 | 6514 | 66.13 | 4987 | 91.143 | 3559 | 116.156 | 4438
16.565 | 0 | 18.347 | 610 | 26.333 | 8880 | 31.349 | 3528 | 36.365 | 8596 | 46.331 | 9913 | 48.31 | 6012 | 66.328 | 5199 | 91.341 | 3670 | 116.354 | 4172
16.763 | 0 | 18.545 | 460 | 26.531 | 8943 | 31.547 | 3677 | 36.563 | 6850 | 46.529 | 9364 | 48.508 | 6204 | 66.526 | 5367 | 91.539 | 3218 | 116.552 | 3875
16.961 | 0 | 18.743 | 357 | 26.729 | 8821 | 31.745 | 3960 | 36.761 | 5909 | 46.727 | 8678 | 48.706 | 6396 | 66.724 | 5140 | 91.737 | 3564 | 116.75 | 3914
17.159 | 0 | 18.941 | 572 | 26.927 | 9026 | 31.943 | 4809 | 36.959 | 6076 | 46.924 | 8201 | 48.904 | 5996 | 66.922 | 5136 | 91.935 | 3461 | 116.948 | 4094
17.357 | 0 | 19.139 | 584 | 27.125 | 9664 | 32.141 | 4951 | 37.157 | 7177 | 47.122 | 7713 | 49.102 | 6141 | 67.12 | 5206 | 92.133 | 3522 | 117.146 | 4129
17.555 | 0 | 19.337 | 936 | 27.323 | 9840 | 32.339 | 5522 | 37.355 | 7338 | 47.32 | 6827 | 49.3 | 5789 | 67.318 | 5271 | 92.331 | 3234 | 117.344 | 4018
17.753 | 0 | 19.535 | 2508 | 27.521 | 9512 | 32.537 | 5330 | 37.553 | 7766 | 47.518 | 7226 | 49.498 | 5887 | 67.516 | 5024 | 92.529 | 3377 | 117.542 | 4061
17.951 | 0 | 19.733 | 3870 | 27.719 | 8227 | 32.735 | 5275 | 37.751 | 7646 | 47.716 | 6774 | 49.696 | 6134 | 67.714 | 5354 | 92.727 | 3341 | 117.74 | 3603
18.149 | 0 | 19.931 | 2522 | 27.917 | 8301 | 32.933 | 4831 | 37.949 | 7097 | 47.914 | 6656 | 49.894 | 5737 | 67.912 | 5367 | 92.925 | 3788 | 117.938 | 3705
18.347 | 0 | 20.129 | 2110 | 28.115 | 8923 | 33.131 | 5323 | 38.147 | 6539 | 48.112 | 6151 | 50.092 | 6243 | 68.11 | 5180 | 93.123 | 3866 | 118.136 | 3728
18.545 | 0 | 20.327 | 2602 | 28.313 | 9730 | 33.329 | 7716 | 38.345 | 5985 | 48.31 | 6498 | 50.29 | 6176 | 68.308 | 5171 | 93.321 | 3716 | 118.334 | 3890
18.743 | 0 | 20.525 | 3647 | 28.511 | 11150 | 33.527 | 14821 | 38.543 | 5368 | 48.508 | 6635 | 50.488 | 6092 | 68.506 | 5161 | 93.519 | 3719 | 118.532 | 3572
18.941 | 0 | 20.723 | 3871 | 28.709 | 12601 | 33.725 | 27512 | 38.741 | 5075 | 48.706 | 6405 | 50.686 | 6061 | 68.704 | 5141 | 93.717 | 3776 | 118.73 | 3380
19.139 | 0 | 20.921 | 3927 | 28.907 | 13696 | 33.923 | 48767 | 38.939 | 5533 | 48.904 | 6134 | 50.884 | 5951 | 68.902 | 4371 | 93.915 | 3553 | 118.928 | 3553
19.337 | 0 | 21.119 | 4439 | 29.105 | 13981 | 34.121 | 77680 | 39.137 | 5907 | 49.102 | 5763 | 51.082 | 5561 | 69.1 | 4558 | 94.113 | 3752 | 119.126 | 3403
19.535 | 0 | 21.317 | 6240 | 29.303 | 11497 | 34.319 | 111093 | 39.335 | 6459 | 49.3 | 5798 | 51.28 | 5049 | 69.298 | 4700 | 94.311 | 3910 | 119.324 | 3357
19.733 | 0 | 21.515 | 7940 | 29.501 | 10073 | 34.517 | 144977 | 39.533 | 5912 | 49.498 | 5901 | 51.478 | 5044 | 69.496 | 4457 | 94.509 | 3817 | 119.522 | 3396
19.931 | 0 | 21.713 | 7823 | 29.699 | 7590 | 34.715 | 167053 | 39.731 | 6753 | 49.696 | 5954 | 51.676 | 5029 | 69.694 | 4242 | 94.707 | 4195 | 119.72 | 3338
20.129 | 0 | 21.911 | 8059 | 29.897 | 5717 | 34.913 | 166972 | 39.929 | 6854 | 49.894 | 5962 | 51.874 | 4794 | 69.892 | 4385 | 94.905 | 4043 | 119.918 | 3582
20.327 | 0 | 22.109 | 9022 | 30.095 | 4967 | 35.111 | 144755 | 40.127 | 7342 | 50.092 | 5662 | 52.072 | 4648 | 70.09 | 3902 | 95.103 | 4136 | 120.116 | 3200
20.525 | 0 | 22.307 | 9529 | 30.293 | 4575 | 35.309 | 111289 | 40.325 | 6589 | 50.29 | 5696 | 52.27 | 4726 | 70.288 | 3798 | 95.301 | 4008 | 120.314 | 3547
20.723 | 0 | 22.505 | 9000 | 30.491 | 5436 | 35.507 | 76989 | 40.523 | 6790 | 50.488 | 6026 | 52.468 | 5018 | 70.486 | 3862 | 95.499 | 4221 | 120.512 | 3382
20.921 | 0 | 22.703 | 9036 | 30.689 | 5047 | 35.705 | 49069 | 40.721 | 6296 | 50.686 | 5368 | 52.666 | 5804 | 70.684 | 3677 | 95.697 | 4208 | 120.71 | 3451
21.119 | 0 | 22.901 | 9816 | 30.887 | 5347 | 35.903 | 31226 | 40.919 | 5944 | 50.884 | 5329 | 52.864 | 5799 | 70.882 | 4196 | 95.895 | 4295 | 120.908 | 3581
21.317 | 0 | 23.099 | 10219 | 31.085 | 4423 | 36.101 | 19869 | 41.117 | 5964 | 51.082 | 5357 | 53.062 | 5773 | 71.08 | 3510 | 96.093 | 3838 | 121.106 | 3503
21.515 | 0 | 23.297 | 10406 | 31.283 | 3567 | 36.299 | 12723 | 41.315 | 5260 | 51.28 | 5097 | 53.26 | 6040 | 71.278 | 3641 | 96.291 | 4278 | 121.304 | 3712
21.713 | 0 | 23.495 | 10389 | 31.481 | 2378 | 36.497 | 8551 | 41.513 | 5345 | 51.478 | 4359 | 53.458 | 5947 | 71.476 | 3904 | 96.489 | 3997 | 121.502 | 3779
21.911 | 0 | 23.693 | 10236 | 31.679 | 1430 | 36.695 | 6417 | 41.711 | 4995 | 51.676 | 3966 | 53.656 | 5550 | 71.674 | 3891 | 96.687 | 3849 | 121.7 | 3663
22.109 | 0 | 23.891 | 10826 | 31.877 | 792 | 36.893 | 5368 | 41.909 | 4463 | 51.874 | 3535 | 53.854 | 5137 | 71.872 | 3746 | 96.885 | 4063 | 121.898 | 3625
22.307 | 0 | 24.089 | 11224 | 32.075 | 546 | 37.091 | 5260 | 42.107 | 4280 | 52.072 | 3318 | 54.052 | 4658 | 72.07 | 3380 | 97.083 | 3857 | 122.096 | 3838
22.505 | 0 | 24.287 | 10836 | 32.273 | 479 | 37.289 | 5129 | 42.305 | 4433 | 52.27 | 3301 | 54.25 | 4103 | 72.268 | 3810 | 97.281 | 3858 | 122.294 | 4023
22.703 | 0 | 24.485 | 11067 | 32.471 | 361 | 37.487 | 5521 | 42.503 | 4147 | 52.468 | 3227 | 54.448 | 3587 | 72.466 | 3856 | 97.479 | 3653 | 122.492 | 3738
22.901 | 0 | 24.683 | 11130 | 32.669 | 408 | 37.685 | 5813 | 42.701 | 4453 | 52.666 | 3291 | 54.646 | 3259 | 72.664 | 3620 | 97.677 | 3614 | 122.69 | 4614
23.099 | 0 | 24.881 | 10814 | 32.867 | 347 | 37.883 | 6548 | 42.899 | 3894 | 52.864 | 3262 | 54.844 | 3097 | 72.862 | 3831 | 97.875 | 4104 | 122.888 | 4716
23.297 | 0 | 25.079 | 11087 | 33.065 | 766 | 38.081 | 5726 | 43.097 | 4320 | 53.062 | 3265 | 55.042 | 3344 | 73.06 | 4278 | 98.073 | 3778 | 123.086 | 4951
23.495 | 0 | 25.277 | 10413 | 33.263 | 1608 | 38.279 | 5991 | 43.295 | 3985 | 53.26 | 3484 | 55.24 | 4412 | 73.258 | 4169 | 98.271 | 3384 | 123.284 | 5138
23.693 | 0 | 25.475 | 9938 | 33.461 | 4555 | 38.477 | 4937 | 43.493 | 3799 | 53.458 | 3394 | 55.438 | 5344 | 73.456 | 3946 | 98.469 | 3840 | 123.482 | 5680
23.891 | 0 | 25.673 | 10025 | 33.659 | 11501 | 38.675 | 4692 | 43.691 | 3630 | 53.656 | 3470 | 55.636 | 7477 | 73.654 | 3525 | 98.667 | 3334 | 123.68 | 6199
24.089 | 0 | 25.871 | 9737 | 33.857 | 25455 | 38.873 | 4826 | 43.889 | 3448 | 53.854 | 3203 | 55.834 | 10135 | 73.852 | 3344 | 98.865 | 3861 | 123.878 | 6453
24.287 | 0 | 26.069 | 8901 | 34.055 | 47584 | 39.071 | 5139 | 44.087 | 3308 | 54.052 | 3281 | 56.032 | 13715 | 74.05 | 2660 | 99.063 | 4005 | 124.076 | 6514
24.485 | 0 | 26.267 | 9212 | 34.253 | 78530 | 39.269 | 5306 | 44.285 | 3385 | 54.25 | 3048 | 56.23 | 18437 | 74.248 | 2366 | 99.261 | 3578 | 124.274 | 6942
24.683 | 0 | 26.465 | 8645 | 34.451 | 117006 | 39.467 | 5396 | 44.483 | 3268 | 54.448 | 2994 | 56.428 | 25095 | 74.446 | 2122 | 99.459 | 3721 | 124.472 | 7687
24.881 | 0 | 26.663 | 9259 | 34.649 | 151710 | 39.665 | 4865 | 44.681 | 3310 | 54.646 | 2812 | 56.626 | 32185 | 74.644 | 2003 | 99.657 | 3850 | 124.67 | 8531
25.079 | 0 | 26.861 | 9131 | 34.847 | 176345 | 39.863 | 4984 | 44.879 | 3662 | 54.844 | 2827 | 56.824 | 40484 | 74.842 | 1657 | 99.855 | 4060 | 124.868 | 8864
25.277 | 0 | 27.059 | 9720 | 35.045 | 178742 | 40.061 | 4809 | 45.077 | 3726 | 55.042 | 3488 | 57.022 | 48787 | 75.04 | 1865 | 100.053 | 4810 | 125.066 | 9161
25.475 | 0 | 27.257 | 8872 | 35.243 | 155571 | 40.259 | 4493 | 45.275 | 3961 | 55.24 | 4078 | 57.22 | 60028 | 75.238 | 1584 | 100.251 | 4504 | 125.264 | 9772
25.673 | 0 | 27.455 | 8557 | 35.441 | 118471 | 40.457 | 4197 | 45.473 | 4506 | 55.438 | 4981 | 57.418 | 68112 | 75.436 | 1590 | 100.449 | 5553 | 125.462 | 10412
25.871 | 0 | 27.653 | 8483 | 35.639 | 79975 | 40.655 | 4021 | 45.671 | 4425 | 55.636 | 7132 | 57.616 | 75734 | 75.634 | 1501 | 100.647 | 5836 | 125.66 | 10715
26.069 | 0 | 27.851 | 7814 | 35.837 | 50016 | 40.853 | 3954 | 45.869 | 5007 | 55.834 | 9379 | 57.814 | 80433 | 75.832 | 1616 | 100.845 | 6619 | 125.858 | 11791
26.267 | 0 | 28.049 | 7731 | 36.035 | 28128 | 41.051 | 3601 | 46.067 | 5514 | 56.032 | 12919 | 58.012 | 80039 | 76.03 | 1585 | 101.043 | 7171 | 126.056 | 11696
26.465 | 0 | 28.247 | 7311 | 36.233 | 15597 | 41.249 | 3245 | 46.265 | 5591 | 56.23 | 17249 | 58.21 | 76177 | 76.228 | 1416 | 101.241 | 8026 | 126.254 | 12233
26.663 | 0 | 28.445 | 8960 | 36.431 | 7886 | 41.447 | 2483 | 46.463 | 5433 | 56.428 | 23464 | 58.408 | 68700 | 76.426 | 1799 | 101.439 | 9205 | 126.452 | 12307
26.861 | 0 | 28.643 | 9866 | 36.629 | 3464 | 41.645 | 1570 | 46.661 | 5537 | 56.626 | 30829 | 58.606 | 58010 | 76.624 | 1792 | 101.637 | 9560 | 126.65 | 12667
27.059 | 0 | 28.841 | 10780 | 36.827 | 1254 | 41.843 | 877 | 46.859 | 5095 | 56.824 | 39511 | 58.804 | 47666 | 76.822 | 1957 | 101.835 | 11183 | 126.848 | 12582
27.257 | 0 | 29.039 | 12305 | 37.025 | 538 | 42.041 | 489 | 47.056 | 5065 | 57.022 | 48826 | 59.002 | 37607 | 77.02 | 1681 | 102.033 | 12141 | 127.046 | 12261
27.455 | 0 | 29.237 | 11967 | 37.223 | 376 | 42.239 | 224 | 47.254 | 5122 | 57.22 | 59755 | 59.2 | 28595 | 77.218 | 2367 | 102.231 | 13218 | 127.244 | 12250
27.653 | 0 | 29.435 | 11251 | 37.421 | 356 | 42.437 | 92 | 47.452 | 4550 | 57.418 | 69124 | 59.398 | 22133 | 77.416 | 2844 | 102.429 | 14946 | 127.442 | 11720
27.851 | 0 | 29.633 | 9842 | 37.619 | 418 | 42.635 | 63 | 47.65 | 4146 | 57.616 | 76459 | 59.596 | 16621 | 77.614 | 3285 | 102.627 | 16013 | 127.64 | 11782
28.049 | 0 | 29.831 | 7611 | 37.817 | 384 | 42.833 | 75 | 47.848 | 3601 | 57.814 | 83191 | 59.794 | 12139 | 77.812 | 3968 | 102.825 | 17027 | 127.838 | 10645
28.247 | 0 | 30.029 | 5810 | 38.015 | 413 | 43.031 | 76 | 48.046 | 2800 | 58.012 | 83456 | 59.992 | 9035 | 78.01 | 4376 | 103.023 | 18093 | 128.036 | 9763
28.445 | 0 | 30.227 | 4052 | 38.213 | 268 | 43.229 | 45 | 48.244 | 2456 | 58.21 | 80021 | 60.19 | 5875 | 78.208 | 5923 | 103.221 | 19328 | 128.234 | 8953
28.643 | 0 | 30.425 | 3469 | 38.411 | 302 | 43.427 | 18 | 48.442 | 2020 | 58.408 | 73684 | 60.388 | 4444 | 78.405 | 6902 | 103.419 | 20109 | 128.432 | 7898
28.841 | 0 | 30.623 | 2605 | 38.609 | 375 | 43.625 | 26 | 48.64 | 1914 | 58.606 | 63741 | 60.586 | 2908 | 78.603 | 8464 | 103.617 | 20629 | 128.63 | 6767
29.039 | 0 | 30.821 | 1954 | 38.807 | 290 | 43.823 | 35 | 48.838 | 1804 | 58.804 | 52289 | 60.784 | 1636 | 78.801 | 10163 | 103.815 | 21182 | 128.828 | 6105
29.237 | 0 | 31.019 | 1664 | 39.005 | 296 | 44.021 | 28 | 49.036 | 1419 | 59.002 | 40738 | 60.982 | 925 | 78.999 | 12726 | 104.013 | 20899 | 129.026 | 5241
29.435 | 0 | 31.217 | 1287 | 39.203 | 278 | 44.219 | 16 | 49.234 | 1255 | 59.2 | 31642 | 61.18 | 548 | 79.197 | 15161 | 104.211 | 20972 | 129.224 | 4667
29.633 | 0 | 31.415 | 961 | 39.401 | 205 | 44.417 | 24 | 49.432 | 1152 | 59.398 | 22984 | 61.378 | 259 | 79.395 | 18936 | 104.409 | 20350 | 129.422 | 3930
29.831 | 0 | 31.613 | 957 | 39.599 | 302 | 44.615 | 27 | 49.63 | 873 | 59.596 | 17267 | 61.576 | 177 | 79.593 | 22369 | 104.607 | 18112 | 129.62 | 3523
30.029 | 0 | 31.811 | 666 | 39.797 | 250 | 44.813 | 15 | 49.828 | 679 | 59.794 | 13448 | 61.774 | 140 | 79.791 | 25725 | 104.805 | 17410 | 129.818 | 3035
30.227 | 0 | 32.009 | 712 | 39.995 | 185 | 45.011 | 40 | 50.026 | 682 | 59.992 | 10536 | 61.972 | 150 | 79.989 | 29131 | 105.003 | 15799 | 130.016 | 2743
30.425 | 0 | 32.207 | 541 | 40.193 | 264 | 45.209 | 18 | 50.224 | 585 | 60.19 | 6976 | 62.17 | 163 | 80.187 | 32029 | 105.201 | 13396 | 130.214 | 2354
30.623 | 0 | 32.405 | 498 | 40.391 | 194 | 45.407 | 27 | 50.422 | 591 | 60.388 | 5061 | 62.368 | 130 | 80.385 | 34808 | 105.399 | 12083 | 130.412 | 2483
30.821 | 0 | 32.603 | 548 | 40.589 | 163 | 45.605 | 14 | 50.62 | 454 | 60.586 | 3434 | 62.566 | 123 | 80.583 | 38305 | 105.597 | 10677 | 130.61 | 1922
31.019 | 0 | 32.801 | 407 | 40.787 | 187 | 45.803 | 21 | 50.818 | 494 | 60.784 | 2218 | 62.764 | 155 | 80.781 | 40033 | 105.795 | 8646 | 130.808 | 1778
31.217 | 0 | 32.999 | 495 | 40.985 | 180 | 46.001 | 14 | 51.016 | 294 | 60.982 | 1244 | 62.962 | 166 | 80.979 | 41160 | 105.993 | 7171 | 131.006 | 1752
31.415 | 0 | 33.197 | 575 | 41.183 | 138 | 46.199 | 6 | 51.214 | 370 | 61.18 | 660 | 63.16 | 152 | 81.177 | 40052 | 106.191 | 6207 | 131.204 | 1476
31.613 | 0 | 33.395 | 1255 | 41.381 | 163 | 46.397 | 3 | 51.412 | 235 | 61.378 | 400 | 63.358 | 183 | 81.375 | 39323 | 106.389 | 5189 | 131.402 | 1467
31.811 | 0 | 33.593 | 3176 | 41.579 | 134 | 46.595 | 26 | 51.61 | 152 | 61.576 | 196 | 63.556 | 125 | 81.573 | 36322 | 106.587 | 4801 | 131.6 | 1242
32.009 | 0 | 33.791 | 8878 | 41.777 | 118 | 46.793 | 7 | 51.808 | 93 | 61.774 | 147 | 63.754 | 84 | 81.771 | 33792 | 106.785 | 3992 | 131.798 | 1111
32.207 | 0 | 33.989 | 20323 | 41.975 | 121 | 46.99 | 3 | 52.006 | 65 | 61.972 | 157 | 63.952 | 181 | 81.969 | 29297 | 106.983 | 3182 | 131.996 | 1083
32.405 | 0 | 34.187 | 39772 | 42.173 | 117 | 47.188 | 3 | 52.204 | 51 | 62.17 | 161 | 64.15 | 115 | 82.167 | 25421 | 107.181 | 3300 | 132.194 | 864
32.603 | 0 | 34.385 | 70213 | 42.371 | 62 | 47.386 | 0 | 52.402 | 29 | 62.368 | 185 | 64.348 | 172 | 82.365 | 20462 | 107.379 | 2989 | 132.392 | 637
32.801 | 0 | 34.583 | 107199 | 42.569 | 80 | 47.584 | 13 | 52.6 | 38 | 62.566 | 102 | 64.546 | 114 | 82.563 | 17100 | 107.577 | 2600 | 132.59 | 614
32.999 | 0 | 34.781 | 147644 | 42.767 | 91 | 47.782 | 23 | 52.798 | 26 | 62.764 | 117 | 64.744 | 138 | 82.761 | 13753 | 107.775 | 2148 | 132.788 | 512
33.197 | 0 | 34.979 | 180762 | 42.965 | 78 | 47.98 | 14 | 52.996 | 16 | 62.962 | 170 | 64.942 | 96 | 82.959 | 11389 | 107.973 | 1656 | 132.986 | 309
33.395 | 0 | 35.177 | 193828 | 43.163 | 82 | 48.178 | 11 | 53.194 | 5 | 63.16 | 122 | 65.14 | 185 | 83.157 | 9405 | 108.171 | 1677 | 133.184 | 351
33.593 | 0 | 35.375 | 177592 | 43.361 | 41 | 48.376 | 6 | 53.392 | 0 | 63.358 | 179 | 65.338 | 92 | 83.355 | 8039 | 108.369 | 1254 | 133.382 | 220
33.791 | 0 | 35.573 | 143125 | 43.559 | 34 | 48.574 | 13 | 53.59 | 20 | 63.556 | 146 | 65.536 | 107 | 83.553 | 6707 | 108.567 | 1030 | 133.58 | 185
33.989 | 0 | 35.771 | 102818 | 43.757 | 41 | 48.772 | 25 | 53.788 | 15 | 63.754 | 179 | 65.734 | 118 | 83.751 | 5660 | 108.765 | 756 | 133.778 | 193
34.187 | 0 | 35.969 | 65300 | 43.955 | 37 | 48.97 | 10 | 53.986 | 4 | 63.952 | 154 | 65.932 | 96 | 83.949 | 4425 | 108.963 | 609 | 133.976 | 112
34.385 | 0 | 36.167 | 39773 | 44.153 | 16 | 49.168 | 11 | 54.184 | 4 | 64.15 | 164 | 66.13 | 86 | 84.147 | 3726 | 109.161 | 425 | 134.174 | 54
34.583 | 0 | 36.365 | 21657 | 44.351 | 22 | 49.366 | 23 | 54.382 | 3 | 64.348 | 124 | 66.328 | 85 | 84.345 | 2880 | 109.359 | 410 | 134.372 | 77
34.781 | 0 | 36.563 | 10489 | 44.549 | 33 | 49.564 | 4 | 54.58 | 25 | 64.546 | 119 | 66.526 | 61 | 84.543 | 2363 | 109.556 | 238 | 134.57 | 58
34.979 | 0 | 36.761 | 4799 | 44.747 | 41 | 49.762 | 2 | 54.778 | 0 | 64.744 | 154 | 66.724 | 121 | 84.741 | 1720 | 109.754 | 155 | 134.768 | 28
35.177 | 0 | 36.959 | 1986 | 44.945 | 31 | 49.96 | 0 | 54.976 | 7 | 64.942 | 125 | 66.922 | 127 | 84.939 | 1107 | 109.952 | 129 | 134.966 | 25
35.375 | 0 | 37.157 | 840 | 45.143 | 11 | 50.158 | 10 | 55.174 | 0 | 65.14 | 144 | 67.12 | 85 | 85.137 | 919 | 110.15 | 96 | 135.164 | 44
35.573 | 0 | 37.355 | 533 | 45.341 | 18 | 50.356 | 0 | 55.372 | 7 | 65.338 | 122 | 67.318 | 101 | 85.335 | 567 | 110.348 | 25 | 135.362 | 18
35.771 | 0 | 37.553 | 463 | 45.539 | 15 | 50.554 | 3 | 55.57 | 0 | 65.536 | 140 | 67.516 | 115 | 85.533 | 356 | 110.546 | 43 | 135.56 | 20
35.969 | 0 | 37.751 | 476 | 45.737 | 0 | 50.752 | 4 | 55.768 | 0 | 65.734 | 118 | 67.714 | 44 | 85.731 | 218 | 110.744 | 41 | 135.758 | 42
36.167 | 0 | 37.949 | 318 | 45.935 | 7 | 50.95 | 0 | 55.966 | 5 | 65.932 | 99 | 67.912 | 58 | 85.929 | 160 | 110.942 | 54 | 135.956 | 24
36.365 | 0 | 38.147 | 391 | 46.133 | 0 | 51.148 | 0 | 56.164 | 0 | 66.13 | 106 | 68.11 | 65 | 86.127 | 130 | 111.14 | 23 | 136.154 | 14
36.563 | 0 | 38.345 | 440 | 46.331 | 0 | 51.346 | 12 | 56.362 | 0 | 66.328 | 90 | 68.308 | 98 | 86.325 | 91 | 111.338 | 19 | 136.352 | 25
36.761 | 0 | 38.543 | 443 | 46.529 | 4 | 51.544 | 0 | 56.56 | 0 | 66.526 | 86 | 68.506 | 57 | 86.523 | 40 | 111.536 | 34 | 136.55 | 6
36.959 | 0 | 38.741 | 411 | 46.727 | 5 | 51.742 | 0 | 56.758 | 0 | 66.724 | 104 | 68.704 | 62 | 86.721 | 33 | 111.734 | 34 | 136.748 | 0
37.157 | 0 | 38.939 | 365 | 46.924 | 0 | 51.94 | 0 | 56.956 | 0 | 66.922 | 113 | 68.902 | 36 | 86.919 | 18 | 111.932 | 14 | 136.946 | 0
37.355 | 0 | 39.137 | 330 | 47.122 | 10 | 52.138 | 0 | 57.154 | 0 | 67.12 | 102 | 69.1 | 84 | 87.117 | 16 | 112.13 | 13 | 137.144 | 7
37.553 | 0 | 39.335 | 373 | 47.32 | 0 | 52.336 | 9 | 57.352 | 0 | 67.318 | 103 | 69.298 | 40 | 87.315 | 7 | 112.328 | 3 | 137.342 | 3
37.751 | 0 | 39.533 | 297 | 47.518 | 0 | 52.534 | 0 | 57.55 | 0 | 67.516 | 94 | 69.496 | 60 | 87.513 | 0 | 112.526 | 7 | 137.54 | 0
37.949 | 0 | 39.731 | 223 | 47.716 | 5 | 52.732 | 0 | 57.748 | 0 | 67.714 | 128 | 69.694 | 63 | 87.711 | 8 | 112.724 | 0 | 137.738 | 4
38.147 | 0 | 39.929 | 257 | 47.914 | 2 | 52.93 | 5 | 57.946 | 0 | 67.912 | 58 | 69.892 | 20 | 87.909 | 0 | 112.922 | 0 | 137.936 | 0
38.345 | 0 | 40.127 | 210 | 48.112 | 8 | 53.128 | 0 | 58.144 | 0 | 68.11 | 66 | 70.09 | 19 | 88.107 | 0 | 113.12 | 0 | 138.134 | 12
38.543 | 0 | 40.325 | 140 | 48.31 | 0 | 53.326 | 0 | 58.342 | 0 | 68.308 | 80 | 70.288 | 37 | 88.305 | 0 | 113.318 | 0 | 138.332 | 12
38.741 | 0 | 40.523 | 135 | 48.508 | 0 | 53.524 | 6 | 58.54 | 0 | 68.506 | 46 | 70.486 | 43 | 88.503 | 0 | 113.516 | 0 | 138.53 | 4
38.939 | 0 | 40.721 | 177 | 48.706 | 0 | 53.722 | 0 | 58.738 | 0 | 68.704 | 73 | 70.684 | 51 | 88.701 | 0 | 113.714 | 0 | 138.728 | 0
39.137 | 0 | 40.919 | 179 | 48.904 | 5 | 53.92 | 0 | 58.936 | 5 | 68.902 | 45 | 70.882 | 25 | 88.899 | 3 | 113.912 | 7 | 138.926 | 0
39.335 | 0 | 41.117 | 144 | 49.102 | 4 | 54.118 | 0 | 59.134 | 0 | 69.1 | 56 | 71.08 | 29 | 89.097 | 0 | 114.11 | 0 | 139.124 | 8
39.533 | 0 | 41.315 | 144 | 49.3 | 0 | 54.316 | 0 | 59.332 | 0 | 69.298 | 61 | 71.278 | 34 | 89.295 | 0 | 114.308 | 0 | 139.322 | 3
39.731 | 0 | 41.513 | 67 | 49.498 | 5 | 54.514 | 0 | 59.53 | 0 | 69.496 | 75 | 71.476 | 38 | 89.493 | 0 | 114.506 | 0 | 139.52 | 0
39.929 | 0 | 41.711 | 91 | 49.696 | 0 | 54.712 | 0 | 59.728 | 0 | 69.694 | 44 | 71.674 | 36 | 89.691 | 0 | 114.704 | 0 | 139.718 | 0
40.127 | 0 | 41.909 | 143 | 49.894 | 0 | 54.91 | 0 | 59.926 | 5 | 69.892 | 76 | 71.872 | 46 | 89.889 | 0 | 114.902 | 0 | 139.916 | 0
40.325 | 0 | 42.107 | 96 | 50.092 | 0 | 55.108 | 0 | 60.124 | 0 | 70.09 | 26 | 72.07 | 14 | 90.087 | 0 | 115.1 | 2 | 140.114 | 3
40.523 | 0 | 42.305 | 81 | 50.29 | 0 | 55.306 | 0 | 60.322 | 0 | 70.288 | 43 | 72.268 | 50 | 90.285 | 4 | 115.298 | 0 | 140.312 | 2
40.721 | 0 | 42.503 | 99 | 50.488 | 0 | 55.504 | 0 | 60.52 | 0 | 70.486 | 18 | 72.466 | 47 | 90.483 | 0 | 115.496 | 0 | 140.51 | 0
40.919 | 0 | 42.701 | 85 | 50.686 | 0 | 55.702 | 0 | 60.718 | 0 | 70.684 | 69 | 72.664 | 27 | 90.681 | 0 | 115.694 | 0 | 140.708 | 6
41.117 | 0 | 42.899 | 93 | 50.884 | 0 | 55.9 | 0 | 60.916 | 0 | 70.882 | 50 | 72.862 | 41 | 90.879 | 0 | 115.892 | 0 | 140.906 | 0
41.315 | 0 | 43.097 | 82 | 51.082 | 0 | 56.098 | 0 | 61.114 | 0 | 71.08 | 25 | 73.06 | 16 | 91.077 | 0 | 116.09 | 0 | 141.103 | 0
41.513 | 0 | 43.295 | 59 | 51.28 | 0 | 56.296 | 0 | 61.312 | 0 | 71.278 | 22 | 73.258 | 11 | 91.275 | 0 | 116.288 | 0 | 141.301 | 6
41.711 | 0 | 43.493 | 38 | 51.478 | 0 | 56.494 | 0 | 61.51 | 0 | 71.476 | 40 | 73.456 | 44 | 91.473 | 0 | 116.486 | 0 | 141.499 | 0
41.909 | 0 | 43.691 | 15 | 51.676 | 0 | 56.692 | 0 | 61.708 | 0 | 71.674 | 51 | 73.654 | 25 | 91.671 | 0 | 116.684 | 0 | 141.697 | 4
42.107 | 0 | 43.889 | 28 | 51.874 | 0 | 56.89 | 0 | 61.906 | 0 | 71.872 | 47 | 73.852 | 6 | 91.869 | 0 | 116.882 | 0 | 141.895 | 0
42.305 | 0 | 44.087 | 31 | 52.072 | 0 | 57.088 | 0 | 62.104 | 0 | 72.07 | 8 | 74.05 | 9 | 92.067 | 8 | 117.08 | 0 | 142.093 | 4
42.503 | 0 | 44.285 | 39 | 52.27 | 0 | 57.286 | 0 | 62.302 | 0 | 72.268 | 16 | 74.248 | 18 | 92.265 | 0 | 117.278 | 0 | 142.291 | 0
42.701 | 0 | 44.483 | 53 | 52.468 | 0 | 57.484 | 0 | 62.5 | 0 | 72.466 | 11 | 74.446 | 31 | 92.463 | 0 | 117.476 | 2 | 142.489 | 0
42.899 | 0 | 44.681 | 14 | 52.666 | 0 | 57.682 | 0 | 62.698 | 0 | 72.664 | 0 | 74.644 | 15 | 92.661 | 5 | 117.674 | 0 | 142.687 | 5
43.097 | 0 | 44.879 | 13 | 52.864 | 4 | 57.88 | 0 | 62.896 | 0 | 72.862 | 3 | 74.842 | 13 | 92.859 | 0 | 117.872 | 0 | 142.885 | 0
43.295 | 0 | 45.077 | 14 | 53.062 | 3 | 58.078 | 0 | 63.094 | 0 | 73.06 | 7 | 75.04 | 11 | 93.057 | 0 | 118.07 | 4 | 143.083 | 0
43.493 | 0 | 45.275 | 6 | 53.26 | 0 | 58.276 | 0 | 63.292 | 0 | 73.258 | 7 | 75.238 | 13 | 93.255 | 0 | 118.268 | 0 | 143.281 | 0
43.691 | 0 | 45.473 | 36 | 53.458 | 0 | 58.474 | 0 | 63.49 | 3 | 73.456 | 0 | 75.436 | 4 | 93.453 | 0 | 118.466 | 0 | 143.479 | 0
43.889 | 0 | 45.671 | 8 | 53.656 | 0 | 58.672 | 0 | 63.688 | 0 | 73.654 | 0 | 75.634 | 1 | 93.651 | 0 | 118.664 | 0 | 143.677 | 4
44.087 | 0 | 45.869 | 16 | 53.854 | 0 | 58.87 | 0 | 63.886 | 0 | 73.852 | 0 | 75.832 | 23 | 93.849 | 0 | 118.862 | 0 | 143.875 | 0
44.285 | 0 | 46.067 | 19 | 54.052 | 0 | 59.068 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 5 | 94.047 | 0 | 119.06 | 0 | 144.073 | 0
44.483 | 0 | 46.265 | 7 | 54.25 | 0 | 59.266 | 8 | 64.282 | 0 | 74.248 | 0 | 76.228 | 3 | 94.245 | 0 | 119.258 | 0 | 144.271 | 0
44.681 | 0 | 46.463 | 0 | 54.448 | 1 | 59.464 | 6 | 64.48 | 0 | 74.446 | 0 | 76.426 | 0 | 94.443 | 0 | 119.456 | 0 | 144.469 | 0
44.879 | 0 | 46.661 | 7 | 54.646 | 0 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 0 | 119.654 | 0 | 144.667 | 6
45.077 | 0 | 46.859 | 6 | 54.844 | 0 | 59.86 | 0 | 64.876 | 0 | 74.842 | 0 | 76.822 | 5 | 94.839 | 0 | 119.852 | 0 | 144.865 | 4
45.275 | 0 | 47.056 | 0 | 55.042 | 0 | 60.058 | 0 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 0 | 120.05 | 0 | 145.063 | 0
45.473 | 0 | 47.254 | 9 | 55.24 | 0 | 60.256 | 0 | 65.272 | 0 | 75.238 | 0 | 77.218 | 0 | 95.235 | 8 | 120.248 | 0 | 145.261 | 0
45.671 | 0 | 47.452 | 19 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 0 | 145.459 | 0
45.869 | 0 | 47.65 | 0 | 55.636 | 0 | 60.652 | 0 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 6 | 120.644 | 0 | 145.657 | 0
46.067 | 0 | 47.848 | 10 | 55.834 | 0 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 4 | 120.842 | 0 | 145.855 | 0
46.265 | 0 | 48.046 | 0 | 56.032 | 0 | 61.048 | 0 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 0 | 121.04 | 0 | 146.053 | 0
46.463 | 0 | 48.244 | 16 | 56.23 | 0 | 61.246 | 0 | 66.262 | 0 | 76.228 | 0 | 78.208 | 5 | 96.225 | 0 | 121.238 | 0 | 146.251 | 0
46.661 | 0 | 48.442 | 0 | 56.428 | 0 | 61.444 | 0 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 0 | 146.449 | 5
46.859 | 0 | 48.64 | 0 | 56.626 | 0 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 0 | 146.647 | 0
47.056 | 0 | 48.838 | 0 | 56.824 | 6 | 61.84 | 0 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 0 | 146.845 | 0
47.254 | 0 | 49.036 | 11 | 57.022 | 0 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 5 | 122.03 | 0 | 147.043 | 0
47.452 | 0 | 49.234 | 0 | 57.22 | 0 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 0 | 147.241 | 0
47.65 | 0 | 49.432 | 12 | 57.418 | 0 | 62.434 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 0 | 57.616 | 0 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 6 | 147.637 | 0
48.046 | 0 | 49.828 | 0 | 57.814 | 0 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 0 | 122.822 | 0 | 147.835 | 0
48.244 | 0 | 50.026 | 0 | 58.012 | 0 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 0 | 148.033 | 0
48.442 | 0 | 50.224 | 0 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 0 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 0 | 148.429 | 0
48.838 | 0 | 50.62 | 0 | 58.606 | 0 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 0 | 123.614 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 6 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 0 | 59.002 | 0 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 0 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 0 | 59.2 | 0 | 64.216 | 0 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 0 | 124.208 | 0 | 149.221 | 0
|  | 51.412 | 0 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 6 | 124.406 | 0 | 149.419 | 0

## Analyte Polyalanine n11
0 |  | 2 |  | 10 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 324 | 12.804 | 0 | 20.789 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 6005 | 13.002 | 0 | 20.987 | 5 | 31.019 | 9 | 40.985 | 17 | 42.965 | 0 | 60.982 | 17 | 85.995 | 9 | 111.008 | 24
11.418 | 17602 | 13.2 | 0 | 21.185 | 583 | 31.217 | 308 | 41.183 | 275 | 43.163 | 178 | 61.18 | 285 | 86.193 | 210 | 111.206 | 243
11.616 | 27876 | 13.398 | 0 | 21.383 | 1293 | 31.415 | 1347 | 41.381 | 1457 | 43.361 | 1190 | 61.378 | 1351 | 86.391 | 1084 | 111.404 | 1192
11.814 | 30375 | 13.596 | 0 | 21.581 | 1969 | 31.613 | 6755 | 41.579 | 2100 | 43.559 | 1944 | 61.576 | 2598 | 86.589 | 2012 | 111.602 | 2155
12.012 | 49657 | 13.794 | 0 | 21.779 | 3406 | 31.811 | 5176 | 41.777 | 2862 | 43.757 | 2599 | 61.774 | 11413 | 86.787 | 5470 | 111.8 | 4145
12.21 | 182527 | 13.992 | 4 | 21.977 | 4096 | 32.009 | 5196 | 41.975 | 3594 | 43.955 | 3046 | 61.972 | 19694 | 86.985 | 7978 | 111.998 | 4910
12.408 | 524489 | 14.19 | 4 | 22.175 | 4624 | 32.207 | 5003 | 42.173 | 3603 | 44.153 | 2746 | 62.17 | 21608 | 87.183 | 7770 | 112.196 | 4770
12.606 | 223138 | 14.388 | 0 | 22.373 | 5024 | 32.405 | 4512 | 42.371 | 3608 | 44.351 | 2819 | 62.368 | 26425 | 87.381 | 8579 | 112.394 | 4991
12.804 | 13245 | 14.586 | 3 | 22.571 | 5329 | 32.603 | 3995 | 42.569 | 3607 | 44.549 | 3194 | 62.566 | 31231 | 87.579 | 10539 | 112.592 | 5948
13.002 | 270 | 14.784 | 0 | 22.769 | 5256 | 32.801 | 3849 | 42.767 | 3338 | 44.747 | 3367 | 62.764 | 36898 | 87.777 | 12848 | 112.79 | 7092
13.2 | 6 | 14.982 | 4 | 22.967 | 5174 | 32.999 | 3402 | 42.965 | 3462 | 44.945 | 3349 | 62.962 | 40932 | 87.975 | 14689 | 112.988 | 7197
13.398 | 0 | 15.18 | 0 | 23.165 | 5466 | 33.197 | 3463 | 43.163 | 3269 | 45.143 | 3910 | 63.16 | 43833 | 88.173 | 17073 | 113.186 | 8363
13.596 | 0 | 15.378 | 0 | 23.363 | 5234 | 33.395 | 3450 | 43.361 | 2963 | 45.341 | 4168 | 63.358 | 43858 | 88.371 | 19286 | 113.384 | 9759
13.794 | 0 | 15.576 | 0 | 23.561 | 5305 | 33.593 | 3496 | 43.559 | 2837 | 45.539 | 3967 | 63.556 | 42906 | 88.569 | 20654 | 113.582 | 10367
13.992 | 0 | 15.773 | 5 | 23.759 | 5575 | 33.791 | 3347 | 43.757 | 2854 | 45.737 | 4195 | 63.754 | 38461 | 88.767 | 20922 | 113.78 | 11092
14.19 | 0 | 15.971 | 0 | 23.957 | 5753 | 33.989 | 3508 | 43.955 | 2607 | 45.935 | 4313 | 63.952 | 33288 | 88.965 | 22197 | 113.978 | 12131
14.388 | 0 | 16.169 | 0 | 24.155 | 5753 | 34.187 | 3349 | 44.153 | 2763 | 46.133 | 4041 | 64.15 | 28726 | 89.163 | 21693 | 114.176 | 12651
14.586 | 0 | 16.367 | 0 | 24.353 | 5365 | 34.385 | 3429 | 44.351 | 2432 | 46.331 | 3996 | 64.348 | 22487 | 89.361 | 21666 | 114.374 | 12831
14.784 | 0 | 16.565 | 0 | 24.551 | 5719 | 34.583 | 2937 | 44.549 | 3029 | 46.529 | 3699 | 64.546 | 17644 | 89.559 | 20821 | 114.572 | 13353
14.982 | 0 | 16.763 | 0 | 24.749 | 5831 | 34.781 | 2741 | 44.747 | 3078 | 46.727 | 3587 | 64.744 | 13516 | 89.757 | 19096 | 114.77 | 13305
15.18 | 0 | 16.961 | 0 | 24.947 | 6326 | 34.979 | 2735 | 44.945 | 3445 | 46.924 | 3311 | 64.942 | 10049 | 89.955 | 17364 | 114.968 | 13033
15.378 | 0 | 17.159 | 0 | 25.145 | 5316 | 35.177 | 2310 | 45.143 | 3764 | 47.122 | 3869 | 65.14 | 7400 | 90.153 | 15747 | 115.166 | 12613
15.576 | 0 | 17.357 | 0 | 25.343 | 5569 | 35.375 | 2541 | 45.341 | 3923 | 47.32 | 3650 | 65.338 | 6026 | 90.351 | 13302 | 115.364 | 13166
15.773 | 0 | 17.555 | 0 | 25.541 | 5696 | 35.573 | 2903 | 45.539 | 4249 | 47.518 | 3746 | 65.536 | 4841 | 90.549 | 11688 | 115.562 | 11680
15.971 | 0 | 17.753 | 0 | 25.739 | 5938 | 35.771 | 2889 | 45.737 | 4502 | 47.716 | 3889 | 65.734 | 3988 | 90.747 | 9935 | 115.76 | 11203
16.169 | 0 | 17.951 | 0 | 25.937 | 5802 | 35.969 | 3816 | 45.935 | 4318 | 47.914 | 3721 | 65.932 | 3599 | 90.945 | 8269 | 115.958 | 10612
16.367 | 0 | 18.149 | 5 | 26.135 | 6404 | 36.167 | 5048 | 46.133 | 4082 | 48.112 | 3352 | 66.13 | 3126 | 91.143 | 6918 | 116.156 | 9990
16.565 | 0 | 18.347 | 145 | 26.333 | 5767 | 36.365 | 7927 | 46.331 | 4148 | 48.31 | 3719 | 66.328 | 2920 | 91.341 | 6263 | 116.354 | 8832
16.763 | 0 | 18.545 | 438 | 26.531 | 5693 | 36.563 | 14065 | 46.529 | 3997 | 48.508 | 3584 | 66.526 | 3019 | 91.539 | 5370 | 116.552 | 7628
16.961 | 0 | 18.743 | 554 | 26.729 | 4959 | 36.761 | 27329 | 46.727 | 4186 | 48.706 | 3786 | 66.724 | 2695 | 91.737 | 4947 | 116.75 | 6570
17.159 | 0 | 18.941 | 1213 | 26.927 | 4251 | 36.959 | 46796 | 46.924 | 3613 | 48.904 | 3504 | 66.922 | 2949 | 91.935 | 3953 | 116.948 | 6056
17.357 | 0 | 19.139 | 2384 | 27.125 | 4027 | 37.157 | 70749 | 47.122 | 3782 | 49.102 | 3297 | 67.12 | 2952 | 92.133 | 3587 | 117.146 | 6026
17.555 | 0 | 19.337 | 2046 | 27.323 | 3617 | 37.355 | 94733 | 47.32 | 3874 | 49.3 | 3327 | 67.318 | 2940 | 92.331 | 3332 | 117.344 | 5208
17.753 | 0 | 19.535 | 959 | 27.521 | 3915 | 37.553 | 110442 | 47.518 | 3824 | 49.498 | 3138 | 67.516 | 3118 | 92.529 | 3359 | 117.542 | 4481
17.951 | 0 | 19.733 | 426 | 27.719 | 3769 | 37.751 | 110424 | 47.716 | 4016 | 49.696 | 3516 | 67.714 | 3236 | 92.727 | 2678 | 117.74 | 4202
18.149 | 0 | 19.931 | 438 | 27.917 | 3615 | 37.949 | 97533 | 47.914 | 4080 | 49.894 | 3477 | 67.912 | 3364 | 92.925 | 2564 | 117.938 | 3681
18.347 | 0 | 20.129 | 418 | 28.115 | 3592 | 38.147 | 74909 | 48.112 | 3912 | 50.092 | 3457 | 68.11 | 3727 | 93.123 | 2881 | 118.136 | 3187
18.545 | 0 | 20.327 | 788 | 28.313 | 4057 | 38.345 | 51807 | 48.31 | 3920 | 50.29 | 3389 | 68.308 | 3590 | 93.321 | 2635 | 118.334 | 3201
18.743 | 0 | 20.525 | 1569 | 28.511 | 4223 | 38.543 | 31774 | 48.508 | 3559 | 50.488 | 4174 | 68.506 | 3265 | 93.519 | 2513 | 118.532 | 2581
18.941 | 0 | 20.723 | 2393 | 28.709 | 4834 | 38.741 | 18960 | 48.706 | 3808 | 50.686 | 3646 | 68.704 | 3696 | 93.717 | 2580 | 118.73 | 2832
19.139 | 0 | 20.921 | 2324 | 28.907 | 5304 | 38.939 | 11855 | 48.904 | 3499 | 50.884 | 4012 | 68.902 | 3149 | 93.915 | 2742 | 118.928 | 2460
19.337 | 0 | 21.119 | 2187 | 29.105 | 5394 | 39.137 | 7067 | 49.102 | 3540 | 51.082 | 3900 | 69.1 | 3042 | 94.113 | 2441 | 119.126 | 2310
19.535 | 0 | 21.317 | 2719 | 29.303 | 5179 | 39.335 | 5368 | 49.3 | 3604 | 51.28 | 3990 | 69.298 | 2838 | 94.311 | 2500 | 119.324 | 1914
19.733 | 0 | 21.515 | 3447 | 29.501 | 4166 | 39.533 | 4779 | 49.498 | 3568 | 51.478 | 4019 | 69.496 | 3032 | 94.509 | 2549 | 119.522 | 1853
19.931 | 0 | 21.713 | 3497 | 29.699 | 3400 | 39.731 | 4725 | 49.696 | 3484 | 51.676 | 3802 | 69.694 | 2895 | 94.707 | 2519 | 119.72 | 1771
20.129 | 0 | 21.911 | 3524 | 29.897 | 3271 | 39.929 | 3808 | 49.894 | 3543 | 51.874 | 3768 | 69.892 | 3086 | 94.905 | 2332 | 119.918 | 1802
20.327 | 0 | 22.109 | 3626 | 30.095 | 3030 | 40.127 | 4189 | 50.092 | 3314 | 52.072 | 3622 | 70.09 | 2886 | 95.103 | 2458 | 120.116 | 1783
20.525 | 0 | 22.307 | 4269 | 30.293 | 3166 | 40.325 | 3991 | 50.29 | 3613 | 52.27 | 4303 | 70.288 | 2832 | 95.301 | 2588 | 120.314 | 1978
20.723 | 0 | 22.505 | 4709 | 30.491 | 3520 | 40.523 | 3571 | 50.488 | 3862 | 52.468 | 3755 | 70.486 | 2671 | 95.499 | 2561 | 120.512 | 1744
20.921 | 0 | 22.703 | 4954 | 30.689 | 4046 | 40.721 | 3532 | 50.686 | 3508 | 52.666 | 4106 | 70.684 | 2758 | 95.697 | 2419 | 120.71 | 1813
21.119 | 0 | 22.901 | 5325 | 30.887 | 4602 | 40.919 | 3568 | 50.884 | 4463 | 52.864 | 3904 | 70.882 | 2804 | 95.895 | 2483 | 120.908 | 1603
21.317 | 0 | 23.099 | 5550 | 31.085 | 4631 | 41.117 | 2978 | 51.082 | 4213 | 53.062 | 4232 | 71.08 | 2611 | 96.093 | 2532 | 121.106 | 1749
21.515 | 0 | 23.297 | 5563 | 31.283 | 4716 | 41.315 | 2681 | 51.28 | 3987 | 53.26 | 4100 | 71.278 | 3071 | 96.291 | 2144 | 121.304 | 1524
21.713 | 0 | 23.495 | 5566 | 31.481 | 4609 | 41.513 | 2766 | 51.478 | 4393 | 53.458 | 4145 | 71.476 | 2963 | 96.489 | 2162 | 121.502 | 1524
21.911 | 0 | 23.693 | 5526 | 31.679 | 4413 | 41.711 | 2490 | 51.676 | 3663 | 53.656 | 3868 | 71.674 | 2706 | 96.687 | 2046 | 121.7 | 1518
22.109 | 0 | 23.891 | 5764 | 31.877 | 3947 | 41.909 | 2340 | 51.874 | 3766 | 53.854 | 3406 | 71.872 | 2857 | 96.885 | 2163 | 121.898 | 1510
22.307 | 0 | 24.089 | 5901 | 32.075 | 3288 | 42.107 | 2282 | 52.072 | 3507 | 54.052 | 3338 | 72.07 | 2611 | 97.083 | 2058 | 122.096 | 1402
22.505 | 0 | 24.287 | 5547 | 32.273 | 3090 | 42.305 | 2124 | 52.27 | 3297 | 54.25 | 2617 | 72.268 | 2755 | 97.281 | 2113 | 122.294 | 1547
22.703 | 0 | 24.485 | 5563 | 32.471 | 2625 | 42.503 | 2162 | 52.468 | 3168 | 54.448 | 2523 | 72.466 | 2799 | 97.479 | 1753 | 122.492 | 1518
22.901 | 0 | 24.683 | 5685 | 32.669 | 2133 | 42.701 | 2106 | 52.666 | 2930 | 54.646 | 2185 | 72.664 | 2825 | 97.677 | 1734 | 122.69 | 1344
23.099 | 0 | 24.881 | 5696 | 32.867 | 1513 | 42.899 | 1579 | 52.864 | 3136 | 54.844 | 2167 | 72.862 | 2965 | 97.875 | 1633 | 122.888 | 1283
23.297 | 0 | 25.079 | 5849 | 33.065 | 1368 | 43.097 | 1805 | 53.062 | 3127 | 55.042 | 2200 | 73.06 | 2539 | 98.073 | 1803 | 123.086 | 1256
23.495 | 0 | 25.277 | 6049 | 33.263 | 1099 | 43.295 | 1863 | 53.26 | 2787 | 55.24 | 1845 | 73.258 | 2454 | 98.271 | 1933 | 123.284 | 1137
23.693 | 0 | 25.475 | 6042 | 33.461 | 658 | 43.493 | 1857 | 53.458 | 2899 | 55.438 | 1687 | 73.456 | 2699 | 98.469 | 1818 | 123.482 | 1164
23.891 | 0 | 25.673 | 5819 | 33.659 | 585 | 43.691 | 1785 | 53.656 | 2833 | 55.636 | 1568 | 73.654 | 2753 | 98.667 | 1724 | 123.68 | 1138
24.089 | 0 | 25.871 | 5459 | 33.857 | 381 | 43.889 | 1713 | 53.854 | 2737 | 55.834 | 1533 | 73.852 | 2409 | 98.865 | 1481 | 123.878 | 1129
24.287 | 0 | 26.069 | 5555 | 34.055 | 350 | 44.087 | 1549 | 54.052 | 2223 | 56.032 | 1697 | 74.05 | 2213 | 99.063 | 1558 | 124.076 | 982
24.485 | 0 | 26.267 | 5274 | 34.253 | 139 | 44.285 | 1514 | 54.25 | 2239 | 56.23 | 1643 | 74.248 | 1914 | 99.261 | 1485 | 124.274 | 1030
24.683 | 0 | 26.465 | 4607 | 34.451 | 176 | 44.483 | 1732 | 54.448 | 2077 | 56.428 | 1711 | 74.446 | 1619 | 99.459 | 1341 | 124.472 | 991
24.881 | 0 | 26.663 | 4702 | 34.649 | 97 | 44.681 | 1720 | 54.646 | 2003 | 56.626 | 1663 | 74.644 | 1486 | 99.657 | 1242 | 124.67 | 976
25.079 | 0 | 26.861 | 4268 | 34.847 | 103 | 44.879 | 1525 | 54.844 | 2056 | 56.824 | 1387 | 74.842 | 1253 | 99.855 | 1319 | 124.868 | 875
25.277 | 0 | 27.059 | 4264 | 35.045 | 96 | 45.077 | 1639 | 55.042 | 2029 | 57.022 | 1623 | 75.04 | 1231 | 100.053 | 1253 | 125.066 | 936
25.475 | 0 | 27.257 | 4295 | 35.243 | 84 | 45.275 | 1824 | 55.24 | 1993 | 57.22 | 1297 | 75.238 | 1351 | 100.251 | 1052 | 125.264 | 791
25.673 | 0 | 27.455 | 4028 | 35.441 | 61 | 45.473 | 1693 | 55.438 | 2001 | 57.418 | 1517 | 75.436 | 1251 | 100.449 | 1248 | 125.462 | 743
25.871 | 0 | 27.653 | 3679 | 35.639 | 55 | 45.671 | 1592 | 55.636 | 1813 | 57.616 | 1411 | 75.634 | 1155 | 100.647 | 1108 | 125.66 | 904
26.069 | 0 | 27.851 | 4027 | 35.837 | 82 | 45.869 | 1788 | 55.834 | 1715 | 57.814 | 1280 | 75.832 | 970 | 100.845 | 1300 | 125.858 | 917
26.267 | 0 | 28.049 | 3885 | 36.035 | 286 | 46.067 | 1840 | 56.032 | 1813 | 58.012 | 1274 | 76.03 | 1045 | 101.043 | 913 | 126.056 | 693
26.465 | 0 | 28.247 | 3577 | 36.233 | 549 | 46.265 | 1820 | 56.23 | 1551 | 58.21 | 1352 | 76.228 | 891 | 101.241 | 1040 | 126.254 | 664
26.663 | 0 | 28.445 | 4207 | 36.431 | 2177 | 46.463 | 1922 | 56.428 | 1175 | 58.408 | 1206 | 76.426 | 947 | 101.439 | 1031 | 126.452 | 553
26.861 | 0 | 28.643 | 4085 | 36.629 | 5251 | 46.661 | 2225 | 56.626 | 982 | 58.606 | 1248 | 76.624 | 859 | 101.637 | 864 | 126.65 | 629
27.059 | 0 | 28.841 | 4361 | 36.827 | 12992 | 46.859 | 2192 | 56.824 | 690 | 58.804 | 1154 | 76.822 | 988 | 101.835 | 710 | 126.848 | 826
27.257 | 0 | 29.039 | 4854 | 37.025 | 27794 | 47.056 | 2093 | 57.022 | 374 | 59.002 | 1074 | 77.02 | 864 | 102.033 | 416 | 127.046 | 458
27.455 | 0 | 29.237 | 5210 | 37.223 | 51314 | 47.254 | 1866 | 57.22 | 218 | 59.2 | 1018 | 77.218 | 867 | 102.231 | 490 | 127.244 | 532
27.653 | 0 | 29.435 | 4838 | 37.421 | 80966 | 47.452 | 1564 | 57.418 | 129 | 59.398 | 826 | 77.416 | 804 | 102.429 | 534 | 127.442 | 636
27.851 | 0 | 29.633 | 4273 | 37.619 | 108657 | 47.65 | 1593 | 57.616 | 119 | 59.596 | 707 | 77.614 | 886 | 102.627 | 399 | 127.64 | 490
28.049 | 0 | 29.831 | 3892 | 37.817 | 126276 | 47.848 | 1580 | 57.814 | 56 | 59.794 | 640 | 77.812 | 758 | 102.825 | 413 | 127.838 | 572
28.247 | 0 | 30.029 | 3277 | 38.015 | 127032 | 48.046 | 1396 | 58.012 | 73 | 59.992 | 338 | 78.01 | 684 | 103.023 | 332 | 128.036 | 659
28.445 | 0 | 30.227 | 2396 | 38.213 | 112396 | 48.244 | 955 | 58.21 | 65 | 60.19 | 246 | 78.208 | 652 | 103.221 | 351 | 128.234 | 573
28.643 | 0 | 30.425 | 2202 | 38.411 | 87005 | 48.442 | 962 | 58.408 | 78 | 60.388 | 133 | 78.405 | 411 | 103.419 | 336 | 128.432 | 617
28.841 | 0 | 30.623 | 2406 | 38.609 | 60288 | 48.64 | 1005 | 58.606 | 38 | 60.586 | 180 | 78.603 | 281 | 103.617 | 367 | 128.63 | 444
29.039 | 0 | 30.821 | 2621 | 38.807 | 37252 | 48.838 | 948 | 58.804 | 22 | 60.784 | 252 | 78.801 | 172 | 103.815 | 381 | 128.828 | 417
29.237 | 0 | 31.019 | 2746 | 39.005 | 20448 | 49.036 | 994 | 59.002 | 38 | 60.982 | 369 | 78.999 | 132 | 104.013 | 294 | 129.026 | 437
29.435 | 0 | 31.217 | 3210 | 39.203 | 11112 | 49.234 | 957 | 59.2 | 21 | 61.18 | 679 | 79.197 | 111 | 104.211 | 294 | 129.224 | 385
29.633 | 0 | 31.415 | 3632 | 39.401 | 5520 | 49.432 | 1000 | 59.398 | 30 | 61.378 | 1165 | 79.395 | 99 | 104.409 | 309 | 129.422 | 540
29.831 | 0 | 31.613 | 3818 | 39.599 | 2496 | 49.63 | 948 | 59.596 | 26 | 61.576 | 2234 | 79.593 | 44 | 104.607 | 346 | 129.62 | 294
30.029 | 0 | 31.811 | 4623 | 39.797 | 1328 | 49.828 | 865 | 59.794 | 56 | 61.774 | 3379 | 79.791 | 50 | 104.805 | 337 | 129.818 | 409
30.227 | 0 | 32.009 | 4590 | 39.995 | 625 | 50.026 | 1174 | 59.992 | 45 | 61.972 | 5974 | 79.989 | 79 | 105.003 | 258 | 130.016 | 364
30.425 | 0 | 32.207 | 4445 | 40.193 | 281 | 50.224 | 1052 | 60.19 | 36 | 62.17 | 9424 | 80.187 | 30 | 105.201 | 308 | 130.214 | 444
30.623 | 0 | 32.405 | 3991 | 40.391 | 200 | 50.422 | 1082 | 60.388 | 71 | 62.368 | 14298 | 80.385 | 32 | 105.399 | 234 | 130.412 | 318
30.821 | 0 | 32.603 | 3349 | 40.589 | 164 | 50.62 | 1164 | 60.586 | 119 | 62.566 | 20474 | 80.583 | 34 | 105.597 | 161 | 130.61 | 449
31.019 | 0 | 32.801 | 2656 | 40.787 | 165 | 50.818 | 1195 | 60.784 | 221 | 62.764 | 28815 | 80.781 | 86 | 105.795 | 75 | 130.808 | 453
31.217 | 0 | 32.999 | 2279 | 40.985 | 130 | 51.016 | 1405 | 60.982 | 371 | 62.962 | 36053 | 80.979 | 45 | 105.993 | 46 | 131.006 | 556
31.415 | 0 | 33.197 | 1944 | 41.183 | 147 | 51.214 | 1021 | 61.18 | 543 | 63.16 | 44258 | 81.177 | 38 | 106.191 | 37 | 131.204 | 585
31.613 | 0 | 33.395 | 1384 | 41.381 | 121 | 51.412 | 873 | 61.378 | 1095 | 63.358 | 50027 | 81.375 | 19 | 106.389 | 43 | 131.402 | 479
31.811 | 0 | 33.593 | 1067 | 41.579 | 180 | 51.61 | 900 | 61.576 | 1826 | 63.556 | 54443 | 81.573 | 11 | 106.587 | 31 | 131.6 | 407
32.009 | 0 | 33.791 | 632 | 41.777 | 168 | 51.808 | 511 | 61.774 | 3127 | 63.754 | 57197 | 81.771 | 67 | 106.785 | 37 | 131.798 | 346
32.207 | 0 | 33.989 | 610 | 41.975 | 104 | 52.006 | 372 | 61.972 | 5193 | 63.952 | 57249 | 81.969 | 44 | 106.983 | 24 | 131.996 | 100
32.405 | 0 | 34.187 | 364 | 42.173 | 197 | 52.204 | 137 | 62.17 | 8284 | 64.15 | 53459 | 82.167 | 12 | 107.181 | 35 | 132.194 | 87
32.603 | 0 | 34.385 | 371 | 42.371 | 146 | 52.402 | 81 | 62.368 | 12630 | 64.348 | 46480 | 82.365 | 12 | 107.379 | 22 | 132.392 | 39
32.801 | 0 | 34.583 | 254 | 42.569 | 82 | 52.6 | 45 | 62.566 | 19213 | 64.546 | 38407 | 82.563 | 13 | 107.577 | 21 | 132.59 | 41
32.999 | 0 | 34.781 | 211 | 42.767 | 131 | 52.798 | 29 | 62.764 | 26264 | 64.744 | 31515 | 82.761 | 35 | 107.775 | 41 | 132.788 | 20
33.197 | 0 | 34.979 | 155 | 42.965 | 107 | 52.996 | 10 | 62.962 | 35634 | 64.942 | 24778 | 82.959 | 28 | 107.973 | 54 | 132.986 | 30
33.395 | 0 | 35.177 | 102 | 43.163 | 94 | 53.194 | 15 | 63.16 | 43515 | 65.14 | 18916 | 83.157 | 30 | 108.171 | 4 | 133.184 | 18
33.593 | 0 | 35.375 | 125 | 43.361 | 102 | 53.392 | 0 | 63.358 | 51023 | 65.338 | 13086 | 83.355 | 11 | 108.369 | 26 | 133.382 | 25
33.791 | 0 | 35.573 | 109 | 43.559 | 97 | 53.59 | 7 | 63.556 | 56589 | 65.536 | 9301 | 83.553 | 26 | 108.567 | 35 | 133.58 | 21
33.989 | 0 | 35.771 | 113 | 43.757 | 84 | 53.788 | 4 | 63.754 | 58833 | 65.734 | 6434 | 83.751 | 16 | 108.765 | 44 | 133.778 | 27
34.187 | 0 | 35.969 | 76 | 43.955 | 92 | 53.986 | 0 | 63.952 | 58713 | 65.932 | 4378 | 83.949 | 18 | 108.963 | 33 | 133.976 | 26
34.385 | 0 | 36.167 | 240 | 44.153 | 125 | 54.184 | 0 | 64.15 | 55795 | 66.13 | 2971 | 84.147 | 12 | 109.161 | 21 | 134.174 | 58
34.583 | 0 | 36.365 | 543 | 44.351 | 29 | 54.382 | 0 | 64.348 | 50911 | 66.328 | 1825 | 84.345 | 14 | 109.359 | 56 | 134.372 | 37
34.781 | 0 | 36.563 | 1586 | 44.549 | 45 | 54.58 | 0 | 64.546 | 43737 | 66.526 | 1407 | 84.543 | 30 | 109.556 | 43 | 134.57 | 32
34.979 | 0 | 36.761 | 4186 | 44.747 | 57 | 54.778 | 6 | 64.744 | 35980 | 66.724 | 1069 | 84.741 | 25 | 109.754 | 29 | 134.768 | 90
35.177 | 0 | 36.959 | 10857 | 44.945 | 71 | 54.976 | 29 | 64.942 | 28394 | 66.922 | 856 | 84.939 | 38 | 109.952 | 60 | 134.966 | 101
35.375 | 0 | 37.157 | 23870 | 45.143 | 27 | 55.174 | 9 | 65.14 | 21860 | 67.12 | 604 | 85.137 | 69 | 110.15 | 121 | 135.164 | 90
35.573 | 0 | 37.355 | 46706 | 45.341 | 39 | 55.372 | 6 | 65.338 | 15686 | 67.318 | 367 | 85.335 | 82 | 110.348 | 84 | 135.362 | 93
35.771 | 0 | 37.553 | 76058 | 45.539 | 17 | 55.57 | 0 | 65.536 | 11193 | 67.516 | 263 | 85.533 | 129 | 110.546 | 116 | 135.56 | 131
35.969 | 0 | 37.751 | 108754 | 45.737 | 38 | 55.768 | 8 | 65.734 | 7584 | 67.714 | 188 | 85.731 | 109 | 110.744 | 115 | 135.758 | 111
36.167 | 0 | 37.949 | 133710 | 45.935 | 37 | 55.966 | 5 | 65.932 | 5221 | 67.912 | 110 | 85.929 | 186 | 110.942 | 224 | 135.956 | 164
36.365 | 0 | 38.147 | 142167 | 46.133 | 37 | 56.164 | 10 | 66.13 | 3521 | 68.11 | 67 | 86.127 | 441 | 111.14 | 286 | 136.154 | 240
36.563 | 0 | 38.345 | 135138 | 46.331 | 69 | 56.362 | 3 | 66.328 | 2297 | 68.308 | 63 | 86.325 | 387 | 111.338 | 246 | 136.352 | 211
36.761 | 0 | 38.543 | 109610 | 46.529 | 34 | 56.56 | 12 | 66.526 | 1728 | 68.506 | 68 | 86.523 | 470 | 111.536 | 411 | 136.55 | 245
36.959 | 0 | 38.741 | 77388 | 46.727 | 21 | 56.758 | 0 | 66.724 | 1299 | 68.704 | 54 | 86.721 | 688 | 111.734 | 433 | 136.748 | 355
37.157 | 0 | 38.939 | 51338 | 46.924 | 31 | 56.956 | 2 | 66.922 | 814 | 68.902 | 90 | 86.919 | 665 | 111.932 | 421 | 136.946 | 204
37.355 | 0 | 39.137 | 29519 | 47.122 | 30 | 57.154 | 11 | 67.12 | 661 | 69.1 | 51 | 87.117 | 832 | 112.13 | 456 | 137.144 | 338
37.553 | 0 | 39.335 | 14885 | 47.32 | 22 | 57.352 | 0 | 67.318 | 471 | 69.298 | 47 | 87.315 | 480 | 112.328 | 327 | 137.342 | 172
37.751 | 0 | 39.533 | 7685 | 47.518 | 13 | 57.55 | 4 | 67.516 | 252 | 69.496 | 51 | 87.513 | 372 | 112.526 | 221 | 137.54 | 80
37.949 | 0 | 39.731 | 3929 | 47.716 | 6 | 57.748 | 8 | 67.714 | 241 | 69.694 | 53 | 87.711 | 299 | 112.724 | 179 | 137.738 | 59
38.147 | 0 | 39.929 | 1871 | 47.914 | 9 | 57.946 | 0 | 67.912 | 93 | 69.892 | 64 | 87.909 | 130 | 112.922 | 70 | 137.936 | 26
38.345 | 0 | 40.127 | 1014 | 48.112 | 4 | 58.144 | 6 | 68.11 | 139 | 70.09 | 95 | 88.107 | 69 | 113.12 | 19 | 138.134 | 18
38.543 | 0 | 40.325 | 486 | 48.31 | 15 | 58.342 | 0 | 68.308 | 66 | 70.288 | 55 | 88.305 | 38 | 113.318 | 53 | 138.332 | 24
38.741 | 0 | 40.523 | 241 | 48.508 | 17 | 58.54 | 0 | 68.506 | 45 | 70.486 | 54 | 88.503 | 23 | 113.516 | 21 | 138.53 | 9
38.939 | 0 | 40.721 | 172 | 48.706 | 0 | 58.738 | 0 | 68.704 | 109 | 70.684 | 25 | 88.701 | 0 | 113.714 | 18 | 138.728 | 0
39.137 | 0 | 40.919 | 162 | 48.904 | 6 | 58.936 | 0 | 68.902 | 46 | 70.882 | 27 | 88.899 | 7 | 113.912 | 0 | 138.926 | 18
39.335 | 0 | 41.117 | 175 | 49.102 | 14 | 59.134 | 0 | 69.1 | 49 | 71.08 | 52 | 89.097 | 0 | 114.11 | 5 | 139.124 | 7
39.533 | 0 | 41.315 | 133 | 49.3 | 0 | 59.332 | 3 | 69.298 | 36 | 71.278 | 21 | 89.295 | 0 | 114.308 | 3 | 139.322 | 9
39.731 | 0 | 41.513 | 147 | 49.498 | 0 | 59.53 | 0 | 69.496 | 39 | 71.476 | 14 | 89.493 | 0 | 114.506 | 0 | 139.52 | 3
39.929 | 0 | 41.711 | 113 | 49.696 | 0 | 59.728 | 0 | 69.694 | 79 | 71.674 | 33 | 89.691 | 0 | 114.704 | 0 | 139.718 | 3
40.127 | 0 | 41.909 | 131 | 49.894 | 0 | 59.926 | 0 | 69.892 | 49 | 71.872 | 65 | 89.889 | 4 | 114.902 | 0 | 139.916 | 10
40.325 | 0 | 42.107 | 194 | 50.092 | 17 | 60.124 | 0 | 70.09 | 54 | 72.07 | 47 | 90.087 | 0 | 115.1 | 4 | 140.114 | 11
40.523 | 0 | 42.305 | 120 | 50.29 | 7 | 60.322 | 5 | 70.288 | 82 | 72.268 | 47 | 90.285 | 0 | 115.298 | 0 | 140.312 | 8
40.721 | 0 | 42.503 | 205 | 50.488 | 0 | 60.52 | 0 | 70.486 | 48 | 72.466 | 35 | 90.483 | 6 | 115.496 | 0 | 140.51 | 25
40.919 | 0 | 42.701 | 124 | 50.686 | 0 | 60.718 | 0 | 70.684 | 40 | 72.664 | 26 | 90.681 | 0 | 115.694 | 3 | 140.708 | 12
41.117 | 0 | 42.899 | 139 | 50.884 | 0 | 60.916 | 0 | 70.882 | 43 | 72.862 | 44 | 90.879 | 0 | 115.892 | 3 | 140.906 | 0
41.315 | 0 | 43.097 | 140 | 51.082 | 0 | 61.114 | 0 | 71.08 | 49 | 73.06 | 32 | 91.077 | 0 | 116.09 | 10 | 141.103 | 5
41.513 | 0 | 43.295 | 138 | 51.28 | 2 | 61.312 | 0 | 71.278 | 50 | 73.258 | 28 | 91.275 | 0 | 116.288 | 0 | 141.301 | 3
41.711 | 0 | 43.493 | 91 | 51.478 | 0 | 61.51 | 0 | 71.476 | 33 | 73.456 | 56 | 91.473 | 0 | 116.486 | 0 | 141.499 | 3
41.909 | 0 | 43.691 | 115 | 51.676 | 0 | 61.708 | 0 | 71.674 | 26 | 73.654 | 12 | 91.671 | 0 | 116.684 | 0 | 141.697 | 6
42.107 | 0 | 43.889 | 55 | 51.874 | 0 | 61.906 | 0 | 71.872 | 41 | 73.852 | 41 | 91.869 | 0 | 116.882 | 3 | 141.895 | 4
42.305 | 0 | 44.087 | 94 | 52.072 | 0 | 62.104 | 0 | 72.07 | 5 | 74.05 | 34 | 92.067 | 0 | 117.08 | 0 | 142.093 | 5
42.503 | 0 | 44.285 | 40 | 52.27 | 0 | 62.302 | 0 | 72.268 | 46 | 74.248 | 43 | 92.265 | 0 | 117.278 | 3 | 142.291 | 0
42.701 | 0 | 44.483 | 91 | 52.468 | 0 | 62.5 | 0 | 72.466 | 10 | 74.446 | 33 | 92.463 | 0 | 117.476 | 3 | 142.489 | 0
42.899 | 0 | 44.681 | 50 | 52.666 | 2 | 62.698 | 0 | 72.664 | 1 | 74.644 | 30 | 92.661 | 0 | 117.674 | 0 | 142.687 | 0
43.097 | 0 | 44.879 | 60 | 52.864 | 0 | 62.896 | 0 | 72.862 | 0 | 74.842 | 57 | 92.859 | 0 | 117.872 | 3 | 142.885 | 6
43.295 | 0 | 45.077 | 24 | 53.062 | 4 | 63.094 | 0 | 73.06 | 0 | 75.04 | 46 | 93.057 | 0 | 118.07 | 0 | 143.083 | 5
43.493 | 0 | 45.275 | 58 | 53.26 | 0 | 63.292 | 0 | 73.258 | 0 | 75.238 | 24 | 93.255 | 0 | 118.268 | 5 | 143.281 | 0
43.691 | 0 | 45.473 | 36 | 53.458 | 0 | 63.49 | 0 | 73.456 | 0 | 75.436 | 18 | 93.453 | 0 | 118.466 | 0 | 143.479 | 0
43.889 | 0 | 45.671 | 43 | 53.656 | 0 | 63.688 | 0 | 73.654 | 5 | 75.634 | 4 | 93.651 | 0 | 118.664 | 0 | 143.677 | 4
44.087 | 0 | 45.869 | 37 | 53.854 | 0 | 63.886 | 0 | 73.852 | 0 | 75.832 | 12 | 93.849 | 0 | 118.862 | 0 | 143.875 | 0
44.285 | 0 | 46.067 | 50 | 54.052 | 0 | 64.084 | 0 | 74.05 | 0 | 76.03 | 49 | 94.047 | 0 | 119.06 | 5 | 144.073 | 0
44.483 | 0 | 46.265 | 48 | 54.25 | 5 | 64.282 | 0 | 74.248 | 0 | 76.228 | 13 | 94.245 | 0 | 119.258 | 8 | 144.271 | 0
44.681 | 0 | 46.463 | 36 | 54.448 | 0 | 64.48 | 0 | 74.446 | 0 | 76.426 | 15 | 94.443 | 0 | 119.456 | 0 | 144.469 | 10
44.879 | 0 | 46.661 | 40 | 54.646 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 0 | 94.641 | 0 | 119.654 | 0 | 144.667 | 13
45.077 | 0 | 46.859 | 30 | 54.844 | 0 | 64.876 | 0 | 74.842 | 0 | 76.822 | 0 | 94.839 | 5 | 119.852 | 0 | 144.865 | 0
45.275 | 0 | 47.056 | 6 | 55.042 | 0 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 0 | 120.05 | 0 | 145.063 | 0
45.473 | 0 | 47.254 | 24 | 55.24 | 0 | 65.272 | 0 | 75.238 | 0 | 77.218 | 3 | 95.235 | 3 | 120.248 | 0 | 145.261 | 0
45.671 | 0 | 47.452 | 30 | 55.438 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 5 | 145.459 | 0
45.869 | 0 | 47.65 | 9 | 55.636 | 0 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 0 | 145.657 | 0
46.067 | 0 | 47.848 | 9 | 55.834 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 0 | 120.842 | 0 | 145.855 | 0
46.265 | 0 | 48.046 | 41 | 56.032 | 0 | 66.064 | 0 | 76.03 | 0 | 78.01 | 0 | 96.027 | 0 | 121.04 | 0 | 146.053 | 0
46.463 | 0 | 48.244 | 20 | 56.23 | 0 | 66.262 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 0 | 146.251 | 0
46.661 | 0 | 48.442 | 4 | 56.428 | 0 | 66.46 | 0 | 76.426 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 5 | 146.449 | 0
46.859 | 0 | 48.64 | 0 | 56.626 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 0 | 146.647 | 0
47.056 | 0 | 48.838 | 8 | 56.824 | 3 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 0 | 146.845 | 0
47.254 | 0 | 49.036 | 3 | 57.022 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 4 | 122.03 | 0 | 147.043 | 0
47.452 | 0 | 49.234 | 4 | 57.22 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 4 | 147.241 | 0
47.65 | 0 | 49.432 | 0 | 57.418 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 0 | 57.616 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 0 | 147.637 | 0
48.046 | 0 | 49.828 | 2 | 57.814 | 2 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 0 | 122.822 | 0 | 147.835 | 0
48.244 | 0 | 50.026 | 22 | 58.012 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 4 | 148.033 | 0
48.442 | 0 | 50.224 | 4 | 58.21 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 0 | 58.408 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 0 | 148.429 | 0
48.838 | 0 | 50.62 | 4 | 58.606 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 0 | 123.614 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 5 | 58.804 | 0 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 0 | 59.002 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 0 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 0 | 59.2 | 0 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 5 | 124.208 | 0 | 149.221 | 0
|  | 51.412 | 0 | 59.398 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 0 | 149.419 | 0

## Analyte Polyalanine n12
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 92 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 2203 | 13.002 | 0 | 20.987 | 4 | 26.003 | 0 | 31.019 | 4 | 42.965 | 6 | 60.982 | 3 | 85.995 | 4 | 111.008 | 0
11.418 | 7909 | 13.2 | 0 | 21.185 | 233 | 26.201 | 113 | 31.217 | 148 | 43.163 | 338 | 61.18 | 86 | 86.193 | 96 | 111.206 | 127
11.616 | 12207 | 13.398 | 0 | 21.383 | 820 | 26.399 | 971 | 31.415 | 820 | 43.361 | 715 | 61.378 | 477 | 86.391 | 349 | 111.404 | 483
11.814 | 12301 | 13.596 | 0 | 21.581 | 1264 | 26.597 | 1375 | 31.613 | 2632 | 43.559 | 1019 | 61.576 | 769 | 86.589 | 727 | 111.602 | 829
12.012 | 16911 | 13.794 | 0 | 21.779 | 1518 | 26.795 | 1510 | 31.811 | 2172 | 43.757 | 1320 | 61.774 | 1002 | 86.787 | 903 | 111.8 | 807
12.21 | 42284 | 13.992 | 0 | 21.977 | 1893 | 26.993 | 2667 | 32.009 | 2116 | 43.955 | 2143 | 61.972 | 1221 | 86.985 | 1141 | 111.998 | 877
12.408 | 266153 | 14.19 | 0 | 22.175 | 1946 | 27.191 | 2570 | 32.207 | 2051 | 44.153 | 2183 | 62.17 | 1293 | 87.183 | 994 | 112.196 | 780
12.606 | 339011 | 14.388 | 5 | 22.373 | 2355 | 27.389 | 2305 | 32.405 | 1956 | 44.351 | 1842 | 62.368 | 1297 | 87.381 | 904 | 112.394 | 1014
12.804 | 52956 | 14.586 | 0 | 22.571 | 2628 | 27.587 | 2808 | 32.603 | 1644 | 44.549 | 1779 | 62.566 | 1237 | 87.579 | 1051 | 112.592 | 890
13.002 | 1628 | 14.784 | 0 | 22.769 | 2463 | 27.785 | 2751 | 32.801 | 1424 | 44.747 | 1789 | 62.764 | 1467 | 87.777 | 903 | 112.79 | 1028
13.2 | 53 | 14.982 | 0 | 22.967 | 2436 | 27.983 | 2781 | 32.999 | 1287 | 44.945 | 1697 | 62.962 | 1430 | 87.975 | 981 | 112.988 | 834
13.398 | 0 | 15.18 | 0 | 23.165 | 2653 | 28.181 | 2176 | 33.197 | 1495 | 45.143 | 1585 | 63.16 | 1274 | 88.173 | 1182 | 113.186 | 1021
13.596 | 0 | 15.378 | 0 | 23.363 | 2981 | 28.379 | 2145 | 33.395 | 1601 | 45.341 | 1752 | 63.358 | 1454 | 88.371 | 1123 | 113.384 | 1012
13.794 | 0 | 15.576 | 0 | 23.561 | 2822 | 28.577 | 1518 | 33.593 | 1751 | 45.539 | 1782 | 63.556 | 1396 | 88.569 | 1083 | 113.582 | 871
13.992 | 0 | 15.773 | 0 | 23.759 | 3121 | 28.775 | 1361 | 33.791 | 1924 | 45.737 | 1757 | 63.754 | 1503 | 88.767 | 1132 | 113.78 | 811
14.19 | 0 | 15.971 | 0 | 23.957 | 2767 | 28.973 | 1157 | 33.989 | 2250 | 45.935 | 1524 | 63.952 | 1304 | 88.965 | 1275 | 113.978 | 1047
14.388 | 0 | 16.169 | 0 | 24.155 | 2934 | 29.171 | 1322 | 34.187 | 2262 | 46.133 | 1534 | 64.15 | 1372 | 89.163 | 1089 | 114.176 | 1054
14.586 | 0 | 16.367 | 0 | 24.353 | 2991 | 29.369 | 1339 | 34.385 | 2026 | 46.331 | 1448 | 64.348 | 1307 | 89.361 | 1340 | 114.374 | 826
14.784 | 0 | 16.565 | 0 | 24.551 | 2983 | 29.567 | 975 | 34.583 | 2274 | 46.529 | 1409 | 64.546 | 1244 | 89.559 | 1375 | 114.572 | 1060
14.982 | 0 | 16.763 | 0 | 24.749 | 2847 | 29.765 | 1247 | 34.781 | 1942 | 46.727 | 1316 | 64.744 | 1316 | 89.757 | 1398 | 114.77 | 1038
15.18 | 0 | 16.961 | 0 | 24.947 | 2939 | 29.963 | 1236 | 34.979 | 2327 | 46.924 | 1153 | 64.942 | 1258 | 89.955 | 1232 | 114.968 | 1016
15.378 | 0 | 17.159 | 0 | 25.145 | 2811 | 30.161 | 1432 | 35.177 | 2248 | 47.122 | 1223 | 65.14 | 1104 | 90.153 | 1371 | 115.166 | 980
15.576 | 0 | 17.357 | 0 | 25.343 | 2671 | 30.359 | 1628 | 35.375 | 2275 | 47.32 | 1287 | 65.338 | 1087 | 90.351 | 1361 | 115.364 | 983
15.773 | 0 | 17.555 | 0 | 25.541 | 2811 | 30.557 | 1767 | 35.573 | 2419 | 47.518 | 1284 | 65.536 | 1083 | 90.549 | 1391 | 115.562 | 981
15.971 | 0 | 17.753 | 0 | 25.739 | 2451 | 30.755 | 1887 | 35.771 | 2946 | 47.716 | 1454 | 65.734 | 1202 | 90.747 | 1272 | 115.76 | 979
16.169 | 0 | 17.951 | 0 | 25.937 | 2613 | 30.953 | 1898 | 35.969 | 3083 | 47.914 | 1514 | 65.932 | 1285 | 90.945 | 1413 | 115.958 | 1079
16.367 | 0 | 18.149 | 0 | 26.135 | 2444 | 31.151 | 2086 | 36.167 | 2954 | 48.112 | 1704 | 66.13 | 1451 | 91.143 | 1480 | 116.156 | 961
16.565 | 0 | 18.347 | 0 | 26.333 | 2539 | 31.349 | 2002 | 36.365 | 2578 | 48.31 | 1604 | 66.328 | 1721 | 91.341 | 1435 | 116.354 | 988
16.763 | 0 | 18.545 | 0 | 26.531 | 2292 | 31.547 | 2125 | 36.563 | 2164 | 48.508 | 1851 | 66.526 | 2269 | 91.539 | 1292 | 116.552 | 932
16.961 | 0 | 18.743 | 9 | 26.729 | 2325 | 31.745 | 1756 | 36.761 | 1733 | 48.706 | 1608 | 66.724 | 3016 | 91.737 | 1416 | 116.75 | 964
17.159 | 0 | 18.941 | 10 | 26.927 | 2283 | 31.943 | 1793 | 36.959 | 1436 | 48.904 | 1763 | 66.922 | 3448 | 91.935 | 1185 | 116.948 | 882
17.357 | 0 | 19.139 | 194 | 27.125 | 2245 | 32.141 | 1870 | 37.157 | 1481 | 49.102 | 1476 | 67.12 | 4652 | 92.133 | 1333 | 117.146 | 850
17.555 | 0 | 19.337 | 609 | 27.323 | 2439 | 32.339 | 1882 | 37.355 | 1590 | 49.3 | 1646 | 67.318 | 5277 | 92.331 | 1372 | 117.344 | 956
17.753 | 0 | 19.535 | 777 | 27.521 | 2610 | 32.537 | 1964 | 37.553 | 1526 | 49.498 | 1690 | 67.516 | 7405 | 92.529 | 1238 | 117.542 | 943
17.951 | 0 | 19.733 | 869 | 27.719 | 2737 | 32.735 | 1733 | 37.751 | 1703 | 49.696 | 1512 | 67.714 | 8988 | 92.727 | 1270 | 117.74 | 865
18.149 | 0 | 19.931 | 1565 | 27.917 | 3133 | 32.933 | 1664 | 37.949 | 1712 | 49.894 | 1603 | 67.912 | 10558 | 92.925 | 1214 | 117.938 | 865
18.347 | 0 | 20.129 | 2088 | 28.115 | 2744 | 33.131 | 1545 | 38.147 | 1896 | 50.092 | 1679 | 68.11 | 13310 | 93.123 | 1157 | 118.136 | 767
18.545 | 0 | 20.327 | 1434 | 28.313 | 2568 | 33.329 | 1583 | 38.345 | 1576 | 50.29 | 1863 | 68.308 | 14664 | 93.321 | 1274 | 118.334 | 767
18.743 | 0 | 20.525 | 629 | 28.511 | 1997 | 33.527 | 1534 | 38.543 | 1761 | 50.488 | 2057 | 68.506 | 18056 | 93.519 | 1322 | 118.532 | 804
18.941 | 0 | 20.723 | 373 | 28.709 | 1638 | 33.725 | 1668 | 38.741 | 2051 | 50.686 | 1987 | 68.704 | 19730 | 93.717 | 1315 | 118.73 | 820
19.139 | 0 | 20.921 | 319 | 28.907 | 1361 | 33.923 | 2036 | 38.939 | 2643 | 50.884 | 2305 | 68.902 | 21670 | 93.915 | 1609 | 118.928 | 850
19.337 | 0 | 21.119 | 284 | 29.105 | 1431 | 34.121 | 2548 | 39.137 | 4121 | 51.082 | 2259 | 69.1 | 23676 | 94.113 | 1596 | 119.126 | 813
19.535 | 0 | 21.317 | 485 | 29.303 | 1644 | 34.319 | 2300 | 39.335 | 7590 | 51.28 | 2478 | 69.298 | 23478 | 94.311 | 1648 | 119.324 | 817
19.733 | 0 | 21.515 | 1155 | 29.501 | 1184 | 34.517 | 1941 | 39.533 | 11620 | 51.478 | 2402 | 69.496 | 23593 | 94.509 | 1591 | 119.522 | 792
19.931 | 0 | 21.713 | 1767 | 29.699 | 1190 | 34.715 | 2112 | 39.731 | 18901 | 51.676 | 2206 | 69.694 | 22466 | 94.707 | 2154 | 119.72 | 961
20.129 | 0 | 21.911 | 2185 | 29.897 | 1245 | 34.913 | 2208 | 39.929 | 27030 | 51.874 | 2135 | 69.892 | 21108 | 94.905 | 2447 | 119.918 | 885
20.327 | 0 | 22.109 | 1799 | 30.095 | 1437 | 35.111 | 1803 | 40.127 | 38330 | 52.072 | 1687 | 70.09 | 19638 | 95.103 | 2623 | 120.116 | 810
20.525 | 0 | 22.307 | 2061 | 30.293 | 1440 | 35.309 | 2160 | 40.325 | 48399 | 52.27 | 1678 | 70.288 | 17081 | 95.301 | 2974 | 120.314 | 750
20.723 | 0 | 22.505 | 2163 | 30.491 | 1664 | 35.507 | 1751 | 40.523 | 55476 | 52.468 | 1669 | 70.486 | 14126 | 95.499 | 3260 | 120.512 | 960
20.921 | 0 | 22.703 | 2682 | 30.689 | 1886 | 35.705 | 1828 | 40.721 | 59080 | 52.666 | 1790 | 70.684 | 11865 | 95.697 | 4153 | 120.71 | 960
21.119 | 0 | 22.901 | 2560 | 30.887 | 2042 | 35.903 | 1794 | 40.919 | 55698 | 52.864 | 1891 | 70.882 | 10079 | 95.895 | 4762 | 120.908 | 755
21.317 | 0 | 23.099 | 2386 | 31.085 | 2084 | 36.101 | 2049 | 41.117 | 49091 | 53.062 | 1828 | 71.08 | 8391 | 96.093 | 5267 | 121.106 | 845
21.515 | 0 | 23.297 | 2898 | 31.283 | 2222 | 36.299 | 2046 | 41.315 | 40653 | 53.26 | 2020 | 71.278 | 6979 | 96.291 | 5847 | 121.304 | 932
21.713 | 0 | 23.495 | 2987 | 31.481 | 2110 | 36.497 | 1832 | 41.513 | 30329 | 53.458 | 1954 | 71.476 | 5387 | 96.489 | 7043 | 121.502 | 1081
21.911 | 0 | 23.693 | 2963 | 31.679 | 1686 | 36.695 | 1592 | 41.711 | 22051 | 53.656 | 1797 | 71.674 | 4764 | 96.687 | 8149 | 121.7 | 1175
22.109 | 0 | 23.891 | 2903 | 31.877 | 1699 | 36.893 | 1489 | 41.909 | 14341 | 53.854 | 1712 | 71.872 | 4230 | 96.885 | 8770 | 121.898 | 1218
22.307 | 0 | 24.089 | 3051 | 32.075 | 1291 | 37.091 | 1474 | 42.107 | 9655 | 54.052 | 1565 | 72.07 | 3096 | 97.083 | 9853 | 122.096 | 1335
22.505 | 0 | 24.287 | 2941 | 32.273 | 1020 | 37.289 | 1303 | 42.305 | 6210 | 54.25 | 1480 | 72.268 | 2829 | 97.281 | 10023 | 122.294 | 1357
22.703 | 0 | 24.485 | 2982 | 32.471 | 1041 | 37.487 | 1461 | 42.503 | 3658 | 54.448 | 1277 | 72.466 | 2568 | 97.479 | 10751 | 122.492 | 1569
22.901 | 0 | 24.683 | 3046 | 32.669 | 774 | 37.685 | 1712 | 42.701 | 2688 | 54.646 | 1274 | 72.664 | 2143 | 97.677 | 11190 | 122.69 | 1828
23.099 | 0 | 24.881 | 2660 | 32.867 | 698 | 37.883 | 1682 | 42.899 | 1840 | 54.844 | 1041 | 72.862 | 1810 | 97.875 | 11437 | 122.888 | 1481
23.297 | 0 | 25.079 | 2974 | 33.065 | 660 | 38.081 | 1588 | 43.097 | 1335 | 55.042 | 1003 | 73.06 | 1633 | 98.073 | 11226 | 123.086 | 1811
23.495 | 0 | 25.277 | 3128 | 33.263 | 623 | 38.279 | 1408 | 43.295 | 1258 | 55.24 | 896 | 73.258 | 1462 | 98.271 | 11398 | 123.284 | 1952
23.693 | 0 | 25.475 | 2747 | 33.461 | 588 | 38.477 | 1641 | 43.493 | 984 | 55.438 | 927 | 73.456 | 1482 | 98.469 | 11070 | 123.482 | 2289
23.891 | 0 | 25.673 | 2914 | 33.659 | 950 | 38.675 | 1775 | 43.691 | 939 | 55.636 | 1196 | 73.654 | 1220 | 98.667 | 11130 | 123.68 | 2301
24.089 | 0 | 25.871 | 2704 | 33.857 | 916 | 38.873 | 1954 | 43.889 | 883 | 55.834 | 1030 | 73.852 | 1115 | 98.865 | 10260 | 123.878 | 2727
24.287 | 0 | 26.069 | 2646 | 34.055 | 942 | 39.071 | 2871 | 44.087 | 914 | 56.032 | 1031 | 74.05 | 933 | 99.063 | 9772 | 124.076 | 2967
24.485 | 0 | 26.267 | 2657 | 34.253 | 938 | 39.269 | 4682 | 44.285 | 967 | 56.23 | 993 | 74.248 | 947 | 99.261 | 8944 | 124.274 | 3152
24.683 | 0 | 26.465 | 2289 | 34.451 | 971 | 39.467 | 7590 | 44.483 | 975 | 56.428 | 1232 | 74.446 | 759 | 99.459 | 8126 | 124.472 | 3703
24.881 | 0 | 26.663 | 2546 | 34.649 | 919 | 39.665 | 12919 | 44.681 | 747 | 56.626 | 1288 | 74.644 | 688 | 99.657 | 7490 | 124.67 | 3651
25.079 | 0 | 26.861 | 2298 | 34.847 | 893 | 39.863 | 19992 | 44.879 | 934 | 56.824 | 1079 | 74.842 | 664 | 99.855 | 6590 | 124.868 | 4223
25.277 | 0 | 27.059 | 2099 | 35.045 | 883 | 40.061 | 29992 | 45.077 | 786 | 57.022 | 1055 | 75.04 | 657 | 100.053 | 5963 | 125.066 | 4417
25.475 | 0 | 27.257 | 2024 | 35.243 | 814 | 40.259 | 41400 | 45.275 | 804 | 57.22 | 1197 | 75.238 | 518 | 100.251 | 4955 | 125.264 | 4840
25.673 | 0 | 27.455 | 1906 | 35.441 | 713 | 40.457 | 51606 | 45.473 | 725 | 57.418 | 976 | 75.436 | 595 | 100.449 | 4144 | 125.462 | 5330
25.871 | 0 | 27.653 | 2151 | 35.639 | 658 | 40.655 | 59912 | 45.671 | 623 | 57.616 | 1122 | 75.634 | 576 | 100.647 | 3686 | 125.66 | 5896
26.069 | 0 | 27.851 | 1901 | 35.837 | 501 | 40.853 | 62616 | 45.869 | 633 | 57.814 | 1066 | 75.832 | 582 | 100.845 | 3343 | 125.858 | 5897
26.267 | 0 | 28.049 | 1793 | 36.035 | 697 | 41.051 | 61122 | 46.067 | 664 | 58.012 | 1100 | 76.03 | 633 | 101.043 | 2959 | 126.056 | 6496
26.465 | 0 | 28.247 | 1702 | 36.233 | 291 | 41.249 | 52553 | 46.265 | 605 | 58.21 | 1143 | 76.228 | 478 | 101.241 | 2484 | 126.254 | 6745
26.663 | 0 | 28.445 | 1711 | 36.431 | 222 | 41.447 | 44132 | 46.463 | 775 | 58.408 | 1209 | 76.426 | 522 | 101.439 | 2160 | 126.452 | 6322
26.861 | 0 | 28.643 | 1608 | 36.629 | 354 | 41.645 | 32813 | 46.661 | 602 | 58.606 | 898 | 76.624 | 478 | 101.637 | 1933 | 126.65 | 6528
27.059 | 0 | 28.841 | 1386 | 36.827 | 194 | 41.843 | 22974 | 46.859 | 555 | 58.804 | 777 | 76.822 | 636 | 101.835 | 1505 | 126.848 | 7085
27.257 | 0 | 29.039 | 1534 | 37.025 | 205 | 42.041 | 15118 | 47.056 | 589 | 59.002 | 970 | 77.02 | 394 | 102.033 | 1322 | 127.046 | 6538
27.455 | 0 | 29.237 | 1446 | 37.223 | 182 | 42.239 | 8936 | 47.254 | 607 | 59.2 | 881 | 77.218 | 425 | 102.231 | 989 | 127.244 | 6627
27.653 | 0 | 29.435 | 1501 | 37.421 | 147 | 42.437 | 5561 | 47.452 | 645 | 59.398 | 702 | 77.416 | 577 | 102.429 | 992 | 127.442 | 6324
27.851 | 0 | 29.633 | 1628 | 37.619 | 116 | 42.635 | 3171 | 47.65 | 571 | 59.596 | 656 | 77.614 | 521 | 102.627 | 764 | 127.64 | 6080
28.049 | 0 | 29.831 | 1501 | 37.817 | 97 | 42.833 | 1826 | 47.848 | 533 | 59.794 | 589 | 77.812 | 532 | 102.825 | 821 | 127.838 | 6506
28.247 | 0 | 30.029 | 1480 | 38.015 | 100 | 43.031 | 902 | 48.046 | 518 | 59.992 | 501 | 78.01 | 417 | 103.023 | 779 | 128.036 | 5650
28.445 | 0 | 30.227 | 1454 | 38.213 | 95 | 43.229 | 597 | 48.244 | 563 | 60.19 | 202 | 78.208 | 451 | 103.221 | 646 | 128.234 | 5521
28.643 | 0 | 30.425 | 1488 | 38.411 | 118 | 43.427 | 377 | 48.442 | 491 | 60.388 | 188 | 78.405 | 503 | 103.419 | 512 | 128.432 | 5049
28.841 | 0 | 30.623 | 1721 | 38.609 | 180 | 43.625 | 255 | 48.64 | 493 | 60.586 | 104 | 78.603 | 396 | 103.617 | 584 | 128.63 | 4754
29.039 | 0 | 30.821 | 1936 | 38.807 | 292 | 43.823 | 171 | 48.838 | 599 | 60.784 | 98 | 78.801 | 289 | 103.815 | 501 | 128.828 | 4443
29.237 | 0 | 31.019 | 2004 | 39.005 | 915 | 44.021 | 154 | 49.036 | 577 | 60.982 | 115 | 78.999 | 327 | 104.013 | 585 | 129.026 | 3905
29.435 | 0 | 31.217 | 2257 | 39.203 | 1476 | 44.219 | 133 | 49.234 | 703 | 61.18 | 160 | 79.197 | 274 | 104.211 | 501 | 129.224 | 3418
29.633 | 0 | 31.415 | 2322 | 39.401 | 3644 | 44.417 | 125 | 49.432 | 697 | 61.378 | 169 | 79.395 | 273 | 104.409 | 440 | 129.422 | 3227
29.831 | 0 | 31.613 | 2210 | 39.599 | 6920 | 44.615 | 98 | 49.63 | 728 | 61.576 | 91 | 79.593 | 335 | 104.607 | 469 | 129.62 | 2871
30.029 | 0 | 31.811 | 2257 | 39.797 | 11972 | 44.813 | 88 | 49.828 | 499 | 61.774 | 97 | 79.791 | 287 | 104.805 | 300 | 129.818 | 2627
30.227 | 0 | 32.009 | 1935 | 39.995 | 20125 | 45.011 | 53 | 50.026 | 599 | 61.972 | 135 | 79.989 | 414 | 105.003 | 409 | 130.016 | 2323
30.425 | 0 | 32.207 | 1608 | 40.193 | 30663 | 45.209 | 70 | 50.224 | 625 | 62.17 | 74 | 80.187 | 289 | 105.201 | 418 | 130.214 | 2044
30.623 | 0 | 32.405 | 1533 | 40.391 | 44836 | 45.407 | 57 | 50.422 | 826 | 62.368 | 66 | 80.385 | 312 | 105.399 | 420 | 130.412 | 1804
30.821 | 0 | 32.603 | 1144 | 40.589 | 57244 | 45.605 | 45 | 50.62 | 709 | 62.566 | 106 | 80.583 | 263 | 105.597 | 318 | 130.61 | 1713
31.019 | 0 | 32.801 | 1014 | 40.787 | 65268 | 45.803 | 49 | 50.818 | 776 | 62.764 | 11 | 80.781 | 278 | 105.795 | 291 | 130.808 | 1441
31.217 | 0 | 32.999 | 784 | 40.985 | 69330 | 46.001 | 41 | 51.016 | 677 | 62.962 | 19 | 80.979 | 284 | 105.993 | 212 | 131.006 | 1322
31.415 | 0 | 33.197 | 822 | 41.183 | 67886 | 46.199 | 34 | 51.214 | 660 | 63.16 | 11 | 81.177 | 254 | 106.191 | 219 | 131.204 | 1192
31.613 | 0 | 33.395 | 675 | 41.381 | 59902 | 46.397 | 31 | 51.412 | 573 | 63.358 | 29 | 81.375 | 249 | 106.389 | 210 | 131.402 | 1108
31.811 | 0 | 33.593 | 812 | 41.579 | 49751 | 46.595 | 49 | 51.61 | 351 | 63.556 | 27 | 81.573 | 333 | 106.587 | 162 | 131.6 | 901
32.009 | 0 | 33.791 | 852 | 41.777 | 37851 | 46.793 | 42 | 51.808 | 268 | 63.754 | 22 | 81.771 | 203 | 106.785 | 206 | 131.798 | 812
32.207 | 0 | 33.989 | 1069 | 41.975 | 26406 | 46.99 | 58 | 52.006 | 140 | 63.952 | 32 | 81.969 | 278 | 106.983 | 171 | 131.996 | 710
32.405 | 0 | 34.187 | 1215 | 42.173 | 16895 | 47.188 | 45 | 52.204 | 87 | 64.15 | 8 | 82.167 | 200 | 107.181 | 137 | 132.194 | 579
32.603 | 0 | 34.385 | 1230 | 42.371 | 10959 | 47.386 | 38 | 52.402 | 19 | 64.348 | 27 | 82.365 | 263 | 107.379 | 111 | 132.392 | 574
32.801 | 0 | 34.583 | 1294 | 42.569 | 6445 | 47.584 | 15 | 52.6 | 46 | 64.546 | 29 | 82.563 | 301 | 107.577 | 155 | 132.59 | 482
32.999 | 0 | 34.781 | 1432 | 42.767 | 3687 | 47.782 | 57 | 52.798 | 6 | 64.744 | 35 | 82.761 | 314 | 107.775 | 154 | 132.788 | 380
33.197 | 0 | 34.979 | 946 | 42.965 | 2163 | 47.98 | 19 | 52.996 | 0 | 64.942 | 25 | 82.959 | 268 | 107.973 | 129 | 132.986 | 328
33.395 | 0 | 35.177 | 922 | 43.163 | 1301 | 48.178 | 47 | 53.194 | 37 | 65.14 | 29 | 83.157 | 201 | 108.171 | 175 | 133.184 | 198
33.593 | 0 | 35.375 | 1146 | 43.361 | 709 | 48.376 | 21 | 53.392 | 0 | 65.338 | 50 | 83.355 | 211 | 108.369 | 124 | 133.382 | 267
33.791 | 0 | 35.573 | 812 | 43.559 | 347 | 48.574 | 38 | 53.59 | 8 | 65.536 | 93 | 83.553 | 233 | 108.567 | 116 | 133.58 | 219
33.989 | 0 | 35.771 | 865 | 43.757 | 262 | 48.772 | 22 | 53.788 | 12 | 65.734 | 45 | 83.751 | 110 | 108.765 | 122 | 133.778 | 208
34.187 | 0 | 35.969 | 775 | 43.955 | 165 | 48.97 | 20 | 53.986 | 0 | 65.932 | 174 | 83.949 | 180 | 108.963 | 121 | 133.976 | 137
34.385 | 0 | 36.167 | 625 | 44.153 | 128 | 49.168 | 51 | 54.184 | 6 | 66.13 | 222 | 84.147 | 118 | 109.161 | 72 | 134.174 | 134
34.583 | 0 | 36.365 | 542 | 44.351 | 134 | 49.366 | 24 | 54.382 | 0 | 66.328 | 317 | 84.345 | 87 | 109.359 | 61 | 134.372 | 193
34.781 | 0 | 36.563 | 401 | 44.549 | 133 | 49.564 | 16 | 54.58 | 3 | 66.526 | 565 | 84.543 | 102 | 109.556 | 112 | 134.57 | 162
34.979 | 0 | 36.761 | 226 | 44.747 | 75 | 49.762 | 43 | 54.778 | 0 | 66.724 | 765 | 84.741 | 211 | 109.754 | 118 | 134.768 | 147
35.177 | 0 | 36.959 | 329 | 44.945 | 65 | 49.96 | 22 | 54.976 | 0 | 66.922 | 1115 | 84.939 | 167 | 109.952 | 55 | 134.966 | 69
35.375 | 0 | 37.157 | 254 | 45.143 | 59 | 50.158 | 33 | 55.174 | 0 | 67.12 | 1454 | 85.137 | 70 | 110.15 | 38 | 135.164 | 80
35.573 | 0 | 37.355 | 196 | 45.341 | 83 | 50.356 | 15 | 55.372 | 0 | 67.318 | 2337 | 85.335 | 142 | 110.348 | 38 | 135.362 | 31
35.771 | 0 | 37.553 | 136 | 45.539 | 73 | 50.554 | 15 | 55.57 | 0 | 67.516 | 3353 | 85.533 | 111 | 110.546 | 40 | 135.56 | 119
35.969 | 0 | 37.751 | 139 | 45.737 | 60 | 50.752 | 14 | 55.768 | 0 | 67.714 | 4656 | 85.731 | 109 | 110.744 | 88 | 135.758 | 85
36.167 | 0 | 37.949 | 69 | 45.935 | 134 | 50.95 | 8 | 55.966 | 2 | 67.912 | 6295 | 85.929 | 97 | 110.942 | 35 | 135.956 | 57
36.365 | 0 | 38.147 | 90 | 46.133 | 62 | 51.148 | 9 | 56.164 | 0 | 68.11 | 8057 | 86.127 | 56 | 111.14 | 41 | 136.154 | 93
36.563 | 0 | 38.345 | 120 | 46.331 | 77 | 51.346 | 10 | 56.362 | 0 | 68.308 | 10602 | 86.325 | 57 | 111.338 | 23 | 136.352 | 26
36.761 | 0 | 38.543 | 121 | 46.529 | 40 | 51.544 | 26 | 56.56 | 0 | 68.506 | 13376 | 86.523 | 42 | 111.536 | 41 | 136.55 | 32
36.959 | 0 | 38.741 | 183 | 46.727 | 64 | 51.742 | 0 | 56.758 | 3 | 68.704 | 16379 | 86.721 | 93 | 111.734 | 25 | 136.748 | 12
37.157 | 0 | 38.939 | 305 | 46.924 | 73 | 51.94 | 10 | 56.956 | 0 | 68.902 | 19669 | 86.919 | 37 | 111.932 | 42 | 136.946 | 32
37.355 | 0 | 39.137 | 705 | 47.122 | 52 | 52.138 | 0 | 57.154 | 0 | 69.1 | 23297 | 87.117 | 18 | 112.13 | 16 | 137.144 | 34
37.553 | 0 | 39.335 | 1404 | 47.32 | 41 | 52.336 | 3 | 57.352 | 0 | 69.298 | 26458 | 87.315 | 0 | 112.328 | 3 | 137.342 | 11
37.751 | 0 | 39.533 | 3099 | 47.518 | 56 | 52.534 | 5 | 57.55 | 0 | 69.496 | 28003 | 87.513 | 10 | 112.526 | 5 | 137.54 | 27
37.949 | 0 | 39.731 | 6311 | 47.716 | 30 | 52.732 | 7 | 57.748 | 0 | 69.694 | 29681 | 87.711 | 4 | 112.724 | 5 | 137.738 | 14
38.147 | 0 | 39.929 | 11141 | 47.914 | 39 | 52.93 | 5 | 57.946 | 0 | 69.892 | 30469 | 87.909 | 5 | 112.922 | 5 | 137.936 | 8
38.345 | 0 | 40.127 | 19724 | 48.112 | 42 | 53.128 | 0 | 58.144 | 0 | 70.09 | 29622 | 88.107 | 8 | 113.12 | 10 | 138.134 | 9
38.543 | 0 | 40.325 | 31011 | 48.31 | 51 | 53.326 | 0 | 58.342 | 0 | 70.288 | 28244 | 88.305 | 19 | 113.318 | 16 | 138.332 | 5
38.741 | 0 | 40.523 | 44136 | 48.508 | 34 | 53.524 | 0 | 58.54 | 0 | 70.486 | 26299 | 88.503 | 0 | 113.516 | 9 | 138.53 | 4
38.939 | 0 | 40.721 | 60098 | 48.706 | 23 | 53.722 | 3 | 58.738 | 0 | 70.684 | 23431 | 88.701 | 0 | 113.714 | 9 | 138.728 | 16
39.137 | 0 | 40.919 | 71285 | 48.904 | 26 | 53.92 | 6 | 58.936 | 0 | 70.882 | 20317 | 88.899 | 0 | 113.912 | 6 | 138.926 | 9
39.335 | 0 | 41.117 | 77733 | 49.102 | 10 | 54.118 | 12 | 59.134 | 0 | 71.08 | 17545 | 89.097 | 2 | 114.11 | 9 | 139.124 | 8
39.533 | 0 | 41.315 | 79937 | 49.3 | 25 | 54.316 | 0 | 59.332 | 6 | 71.278 | 14492 | 89.295 | 9 | 114.308 | 11 | 139.322 | 10
39.731 | 0 | 41.513 | 72138 | 49.498 | 27 | 54.514 | 0 | 59.53 | 0 | 71.476 | 11353 | 89.493 | 5 | 114.506 | 17 | 139.52 | 10
39.929 | 0 | 41.711 | 61377 | 49.696 | 10 | 54.712 | 11 | 59.728 | 0 | 71.674 | 9412 | 89.691 | 9 | 114.704 | 3 | 139.718 | 9
40.127 | 0 | 41.909 | 47480 | 49.894 | 11 | 54.91 | 3 | 59.926 | 2 | 71.872 | 7724 | 89.889 | 7 | 114.902 | 15 | 139.916 | 0
40.325 | 0 | 42.107 | 33813 | 50.092 | 16 | 55.108 | 5 | 60.124 | 6 | 72.07 | 6123 | 90.087 | 6 | 115.1 | 0 | 140.114 | 12
40.523 | 0 | 42.305 | 22998 | 50.29 | 18 | 55.306 | 0 | 60.322 | 0 | 72.268 | 4681 | 90.285 | 10 | 115.298 | 5 | 140.312 | 21
40.721 | 0 | 42.503 | 14755 | 50.488 | 12 | 55.504 | 0 | 60.52 | 0 | 72.466 | 3413 | 90.483 | 0 | 115.496 | 0 | 140.51 | 9
40.919 | 0 | 42.701 | 9156 | 50.686 | 8 | 55.702 | 5 | 60.718 | 0 | 72.664 | 2576 | 90.681 | 0 | 115.694 | 7 | 140.708 | 0
41.117 | 0 | 42.899 | 5156 | 50.884 | 20 | 55.9 | 0 | 60.916 | 3 | 72.862 | 1913 | 90.879 | 0 | 115.892 | 7 | 140.906 | 0
41.315 | 0 | 43.097 | 3331 | 51.082 | 5 | 56.098 | 0 | 61.114 | 0 | 73.06 | 1491 | 91.077 | 4 | 116.09 | 12 | 141.103 | 4
41.513 | 0 | 43.295 | 1767 | 51.28 | 0 | 56.296 | 0 | 61.312 | 0 | 73.258 | 1214 | 91.275 | 0 | 116.288 | 3 | 141.301 | 11
41.711 | 0 | 43.493 | 996 | 51.478 | 10 | 56.494 | 0 | 61.51 | 0 | 73.456 | 839 | 91.473 | 0 | 116.486 | 9 | 141.499 | 3
41.909 | 0 | 43.691 | 576 | 51.676 | 0 | 56.692 | 0 | 61.708 | 0 | 73.654 | 641 | 91.671 | 0 | 116.684 | 0 | 141.697 | 0
42.107 | 0 | 43.889 | 257 | 51.874 | 9 | 56.89 | 0 | 61.906 | 0 | 73.852 | 549 | 91.869 | 3 | 116.882 | 4 | 141.895 | 5
42.305 | 0 | 44.087 | 255 | 52.072 | 13 | 57.088 | 0 | 62.104 | 0 | 74.05 | 359 | 92.067 | 4 | 117.08 | 17 | 142.093 | 3
42.503 | 0 | 44.285 | 141 | 52.27 | 5 | 57.286 | 0 | 62.302 | 0 | 74.248 | 191 | 92.265 | 0 | 117.278 | 21 | 142.291 | 11
42.701 | 0 | 44.483 | 128 | 52.468 | 6 | 57.484 | 0 | 62.5 | 0 | 74.446 | 174 | 92.463 | 0 | 117.476 | 10 | 142.489 | 25
42.899 | 0 | 44.681 | 117 | 52.666 | 0 | 57.682 | 0 | 62.698 | 0 | 74.644 | 163 | 92.661 | 0 | 117.674 | 19 | 142.687 | 15
43.097 | 0 | 44.879 | 66 | 52.864 | 0 | 57.88 | 0 | 62.896 | 0 | 74.842 | 102 | 92.859 | 0 | 117.872 | 0 | 142.885 | 0
43.295 | 0 | 45.077 | 80 | 53.062 | 0 | 58.078 | 6 | 63.094 | 0 | 75.04 | 88 | 93.057 | 8 | 118.07 | 8 | 143.083 | 7
43.493 | 0 | 45.275 | 107 | 53.26 | 0 | 58.276 | 0 | 63.292 | 0 | 75.238 | 45 | 93.255 | 0 | 118.268 | 0 | 143.281 | 3
43.691 | 0 | 45.473 | 67 | 53.458 | 0 | 58.474 | 0 | 63.49 | 0 | 75.436 | 46 | 93.453 | 0 | 118.466 | 13 | 143.479 | 3
43.889 | 0 | 45.671 | 25 | 53.656 | 0 | 58.672 | 0 | 63.688 | 0 | 75.634 | 22 | 93.651 | 0 | 118.664 | 5 | 143.677 | 0
44.087 | 0 | 45.869 | 54 | 53.854 | 8 | 58.87 | 0 | 63.886 | 0 | 75.832 | 7 | 93.849 | 0 | 118.862 | 7 | 143.875 | 2
44.285 | 0 | 46.067 | 77 | 54.052 | 0 | 59.068 | 9 | 64.084 | 4 | 76.03 | 21 | 94.047 | 0 | 119.06 | 3 | 144.073 | 12
44.483 | 0 | 46.265 | 79 | 54.25 | 6 | 59.266 | 6 | 64.282 | 0 | 76.228 | 10 | 94.245 | 0 | 119.258 | 13 | 144.271 | 0
44.681 | 0 | 46.463 | 97 | 54.448 | 6 | 59.464 | 0 | 64.48 | 0 | 76.426 | 12 | 94.443 | 0 | 119.456 | 5 | 144.469 | 4
44.879 | 0 | 46.661 | 44 | 54.646 | 0 | 59.662 | 0 | 64.678 | 0 | 76.624 | 0 | 94.641 | 0 | 119.654 | 12 | 144.667 | 3
45.077 | 0 | 46.859 | 85 | 54.844 | 0 | 59.86 | 2 | 64.876 | 0 | 76.822 | 9 | 94.839 | 0 | 119.852 | 0 | 144.865 | 3
45.275 | 0 | 47.056 | 42 | 55.042 | 0 | 60.058 | 0 | 65.074 | 0 | 77.02 | 6 | 95.037 | 0 | 120.05 | 6 | 145.063 | 2
45.473 | 0 | 47.254 | 30 | 55.24 | 0 | 60.256 | 0 | 65.272 | 0 | 77.218 | 0 | 95.235 | 5 | 120.248 | 0 | 145.261 | 0
45.671 | 0 | 47.452 | 30 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 77.416 | 0 | 95.433 | 0 | 120.446 | 0 | 145.459 | 0
45.869 | 0 | 47.65 | 95 | 55.636 | 0 | 60.652 | 0 | 65.668 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 6 | 145.657 | 7
46.067 | 0 | 47.848 | 36 | 55.834 | 0 | 60.85 | 3 | 65.866 | 0 | 77.812 | 0 | 95.829 | 0 | 120.842 | 0 | 145.855 | 0
46.265 | 0 | 48.046 | 13 | 56.032 | 4 | 61.048 | 6 | 66.064 | 0 | 78.01 | 0 | 96.027 | 0 | 121.04 | 9 | 146.053 | 0
46.463 | 0 | 48.244 | 59 | 56.23 | 0 | 61.246 | 0 | 66.262 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 3 | 146.251 | 0
46.661 | 0 | 48.442 | 40 | 56.428 | 4 | 61.444 | 0 | 66.46 | 0 | 78.405 | 0 | 96.423 | 0 | 121.436 | 0 | 146.449 | 0
46.859 | 0 | 48.64 | 61 | 56.626 | 5 | 61.642 | 0 | 66.658 | 0 | 78.603 | 4 | 96.621 | 0 | 121.634 | 12 | 146.647 | 0
47.056 | 0 | 48.838 | 27 | 56.824 | 5 | 61.84 | 0 | 66.856 | 0 | 78.801 | 0 | 96.819 | 2 | 121.832 | 0 | 146.845 | 0
47.254 | 0 | 49.036 | 84 | 57.022 | 7 | 62.038 | 0 | 67.054 | 0 | 78.999 | 0 | 97.017 | 0 | 122.03 | 0 | 147.043 | 0
47.452 | 0 | 49.234 | 11 | 57.22 | 0 | 62.236 | 0 | 67.252 | 0 | 79.197 | 0 | 97.215 | 0 | 122.228 | 2 | 147.241 | 0
47.65 | 0 | 49.432 | 12 | 57.418 | 0 | 62.434 | 0 | 67.45 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 0 | 147.439 | 0
47.848 | 0 | 49.63 | 56 | 57.616 | 4 | 62.632 | 0 | 67.648 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 0 | 147.637 | 0
48.046 | 0 | 49.828 | 42 | 57.814 | 9 | 62.83 | 6 | 67.846 | 0 | 79.791 | 0 | 97.809 | 0 | 122.822 | 5 | 147.835 | 0
48.244 | 0 | 50.026 | 35 | 58.012 | 0 | 63.028 | 0 | 68.044 | 0 | 79.989 | 0 | 98.007 | 0 | 123.02 | 6 | 148.033 | 0
48.442 | 0 | 50.224 | 20 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 80.187 | 0 | 98.205 | 0 | 123.218 | 7 | 148.231 | 0
48.64 | 0 | 50.422 | 25 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 3 | 148.429 | 0
48.838 | 0 | 50.62 | 23 | 58.606 | 0 | 63.622 | 0 | 68.638 | 0 | 80.583 | 0 | 98.601 | 0 | 123.614 | 11 | 148.627 | 0
49.036 | 0 | 50.818 | 13 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 0 | 148.825 | 0
49.234 | 0 | 51.016 | 41 | 59.002 | 0 | 64.018 | 0 | 69.034 | 0 | 80.979 | 0 | 98.997 | 0 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 6 | 59.2 | 0 | 64.216 | 0 | 69.232 | 0 | 81.177 | 0 | 99.195 | 6 | 124.208 | 10 | 149.221 | 0
|  | 51.412 | 2 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 12 | 149.419 | 0

## Analyte Polyalanine n13
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 4 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 344 | 13.002 | 0 | 20.987 | 0 | 26.003 | 0 | 31.019 | 0 | 40.985 | 0 | 60.982 | 0 | 85.995 | 0 | 111.008 | 0
11.418 | 2692 | 13.2 | 0 | 21.185 | 9 | 26.201 | 48 | 31.217 | 15 | 41.183 | 20 | 61.18 | 31 | 86.193 | 36 | 111.206 | 39
11.616 | 5771 | 13.398 | 0 | 21.383 | 51 | 26.399 | 525 | 31.415 | 68 | 41.381 | 236 | 61.378 | 254 | 86.391 | 116 | 111.404 | 157
11.814 | 5690 | 13.596 | 0 | 21.581 | 98 | 26.597 | 816 | 31.613 | 432 | 41.579 | 408 | 61.576 | 293 | 86.589 | 203 | 111.602 | 280
12.012 | 6516 | 13.794 | 0 | 21.779 | 84 | 26.795 | 738 | 31.811 | 489 | 41.777 | 537 | 61.774 | 509 | 86.787 | 451 | 111.8 | 536
12.21 | 12403 | 13.992 | 0 | 21.977 | 270 | 26.993 | 975 | 32.009 | 577 | 41.975 | 2824 | 61.972 | 615 | 86.985 | 493 | 111.998 | 717
12.408 | 76052 | 14.19 | 0 | 22.175 | 491 | 27.191 | 1157 | 32.207 | 510 | 42.173 | 5719 | 62.17 | 577 | 87.183 | 543 | 112.196 | 999
12.606 | 279852 | 14.388 | 0 | 22.373 | 982 | 27.389 | 894 | 32.405 | 529 | 42.371 | 8113 | 62.368 | 526 | 87.381 | 511 | 112.394 | 603
12.804 | 122060 | 14.586 | 0 | 22.571 | 1266 | 27.587 | 843 | 32.603 | 688 | 42.569 | 11718 | 62.566 | 524 | 87.579 | 418 | 112.592 | 590
13.002 | 7701 | 14.784 | 0 | 22.769 | 1397 | 27.785 | 814 | 32.801 | 617 | 42.767 | 17344 | 62.764 | 635 | 87.777 | 467 | 112.79 | 602
13.2 | 110 | 14.982 | 0 | 22.967 | 1314 | 27.983 | 930 | 32.999 | 754 | 42.965 | 23474 | 62.962 | 759 | 87.975 | 502 | 112.988 | 612
13.398 | 23 | 15.18 | 5 | 23.165 | 1255 | 28.181 | 788 | 33.197 | 900 | 43.163 | 28752 | 63.16 | 591 | 88.173 | 421 | 113.186 | 577
13.596 | 0 | 15.378 | 0 | 23.363 | 1439 | 28.379 | 738 | 33.395 | 1253 | 43.361 | 31608 | 63.358 | 840 | 88.371 | 451 | 113.384 | 475
13.794 | 0 | 15.576 | 0 | 23.561 | 1757 | 28.577 | 779 | 33.593 | 1221 | 43.559 | 33198 | 63.556 | 763 | 88.569 | 451 | 113.582 | 576
13.992 | 0 | 15.773 | 0 | 23.759 | 1602 | 28.775 | 851 | 33.791 | 1386 | 43.757 | 31928 | 63.754 | 812 | 88.767 | 485 | 113.78 | 698
14.19 | 0 | 15.971 | 0 | 23.957 | 1548 | 28.973 | 860 | 33.989 | 1486 | 43.955 | 27815 | 63.952 | 877 | 88.965 | 624 | 113.978 | 455
14.388 | 0 | 16.169 | 0 | 24.155 | 1527 | 29.171 | 902 | 34.187 | 1221 | 44.153 | 22684 | 64.15 | 756 | 89.163 | 673 | 114.176 | 523
14.586 | 0 | 16.367 | 0 | 24.353 | 1594 | 29.369 | 831 | 34.385 | 1250 | 44.351 | 17897 | 64.348 | 848 | 89.361 | 614 | 114.374 | 490
14.784 | 0 | 16.565 | 0 | 24.551 | 1648 | 29.567 | 892 | 34.583 | 1180 | 44.549 | 13263 | 64.546 | 798 | 89.559 | 534 | 114.572 | 509
14.982 | 0 | 16.763 | 0 | 24.749 | 1473 | 29.765 | 951 | 34.781 | 998 | 44.747 | 8867 | 64.744 | 851 | 89.757 | 693 | 114.77 | 530
15.18 | 0 | 16.961 | 0 | 24.947 | 1440 | 29.963 | 787 | 34.979 | 1012 | 44.945 | 6639 | 64.942 | 605 | 89.955 | 622 | 114.968 | 486
15.378 | 0 | 17.159 | 0 | 25.145 | 1372 | 30.161 | 645 | 35.177 | 847 | 45.143 | 4583 | 65.14 | 678 | 90.153 | 592 | 115.166 | 430
15.576 | 0 | 17.357 | 0 | 25.343 | 1534 | 30.359 | 575 | 35.375 | 893 | 45.341 | 3121 | 65.338 | 633 | 90.351 | 632 | 115.364 | 532
15.773 | 0 | 17.555 | 0 | 25.541 | 1278 | 30.557 | 423 | 35.573 | 843 | 45.539 | 2533 | 65.536 | 596 | 90.549 | 808 | 115.562 | 487
15.971 | 0 | 17.753 | 0 | 25.739 | 1215 | 30.755 | 410 | 35.771 | 1159 | 45.737 | 1822 | 65.734 | 609 | 90.747 | 739 | 115.76 | 535
16.169 | 0 | 17.951 | 0 | 25.937 | 1271 | 30.953 | 381 | 35.969 | 1139 | 45.935 | 1321 | 65.932 | 617 | 90.945 | 825 | 115.958 | 504
16.367 | 0 | 18.149 | 0 | 26.135 | 1170 | 31.151 | 447 | 36.167 | 1090 | 46.133 | 1121 | 66.13 | 543 | 91.143 | 729 | 116.156 | 564
16.565 | 0 | 18.347 | 0 | 26.333 | 1304 | 31.349 | 439 | 36.365 | 1065 | 46.331 | 1041 | 66.328 | 568 | 91.341 | 581 | 116.354 | 417
16.763 | 0 | 18.545 | 0 | 26.531 | 1235 | 31.547 | 389 | 36.563 | 995 | 46.529 | 874 | 66.526 | 650 | 91.539 | 789 | 116.552 | 443
16.961 | 0 | 18.743 | 0 | 26.729 | 1224 | 31.745 | 526 | 36.761 | 837 | 46.727 | 1103 | 66.724 | 802 | 91.737 | 596 | 116.75 | 549
17.159 | 0 | 18.941 | 0 | 26.927 | 1105 | 31.943 | 620 | 36.959 | 939 | 46.924 | 786 | 66.922 | 555 | 91.935 | 558 | 116.948 | 526
17.357 | 0 | 19.139 | 0 | 27.125 | 997 | 32.141 | 546 | 37.157 | 960 | 47.122 | 1037 | 67.12 | 639 | 92.133 | 579 | 117.146 | 501
17.555 | 0 | 19.337 | 0 | 27.323 | 993 | 32.339 | 722 | 37.355 | 1096 | 47.32 | 1042 | 67.318 | 643 | 92.331 | 557 | 117.344 | 469
17.753 | 0 | 19.535 | 0 | 27.521 | 916 | 32.537 | 700 | 37.553 | 1056 | 47.518 | 1032 | 67.516 | 689 | 92.529 | 554 | 117.542 | 583
17.951 | 0 | 19.733 | 6 | 27.719 | 988 | 32.735 | 733 | 37.751 | 1378 | 47.716 | 940 | 67.714 | 626 | 92.727 | 550 | 117.74 | 552
18.149 | 0 | 19.931 | 46 | 27.917 | 816 | 32.933 | 691 | 37.949 | 1364 | 47.914 | 1052 | 67.912 | 745 | 92.925 | 497 | 117.938 | 516
18.347 | 0 | 20.129 | 204 | 28.115 | 773 | 33.131 | 810 | 38.147 | 1332 | 48.112 | 908 | 68.11 | 705 | 93.123 | 528 | 118.136 | 574
18.545 | 0 | 20.327 | 321 | 28.313 | 806 | 33.329 | 935 | 38.345 | 1104 | 48.31 | 885 | 68.308 | 745 | 93.321 | 380 | 118.334 | 426
18.743 | 0 | 20.525 | 336 | 28.511 | 886 | 33.527 | 1135 | 38.543 | 1315 | 48.508 | 904 | 68.506 | 819 | 93.519 | 565 | 118.532 | 456
18.941 | 0 | 20.723 | 462 | 28.709 | 828 | 33.725 | 1337 | 38.741 | 1102 | 48.706 | 809 | 68.704 | 677 | 93.717 | 434 | 118.73 | 402
19.139 | 0 | 20.921 | 770 | 28.907 | 850 | 33.923 | 1564 | 38.939 | 1073 | 48.904 | 691 | 68.902 | 735 | 93.915 | 514 | 118.928 | 395
19.337 | 0 | 21.119 | 624 | 29.105 | 903 | 34.121 | 1417 | 39.137 | 970 | 49.102 | 688 | 69.1 | 591 | 94.113 | 546 | 119.126 | 476
19.535 | 0 | 21.317 | 556 | 29.303 | 1020 | 34.319 | 1488 | 39.335 | 978 | 49.3 | 662 | 69.298 | 612 | 94.311 | 439 | 119.324 | 310
19.733 | 0 | 21.515 | 327 | 29.501 | 793 | 34.517 | 1332 | 39.533 | 870 | 49.498 | 534 | 69.496 | 406 | 94.509 | 464 | 119.522 | 549
19.931 | 0 | 21.713 | 177 | 29.699 | 839 | 34.715 | 1112 | 39.731 | 835 | 49.696 | 624 | 69.694 | 513 | 94.707 | 434 | 119.72 | 345
20.129 | 0 | 21.911 | 226 | 29.897 | 995 | 34.913 | 1159 | 39.929 | 721 | 49.894 | 527 | 69.892 | 523 | 94.905 | 455 | 119.918 | 520
20.327 | 0 | 22.109 | 287 | 30.095 | 842 | 35.111 | 849 | 40.127 | 787 | 50.092 | 661 | 70.09 | 521 | 95.103 | 538 | 120.116 | 456
20.525 | 0 | 22.307 | 404 | 30.293 | 719 | 35.309 | 946 | 40.325 | 676 | 50.29 | 599 | 70.288 | 398 | 95.301 | 564 | 120.314 | 597
20.723 | 0 | 22.505 | 749 | 30.491 | 538 | 35.507 | 1047 | 40.523 | 711 | 50.488 | 626 | 70.486 | 456 | 95.499 | 576 | 120.512 | 586
20.921 | 0 | 22.703 | 1099 | 30.689 | 583 | 35.705 | 925 | 40.721 | 553 | 50.686 | 712 | 70.684 | 544 | 95.697 | 561 | 120.71 | 528
21.119 | 0 | 22.901 | 1306 | 30.887 | 499 | 35.903 | 929 | 40.919 | 745 | 50.884 | 627 | 70.882 | 607 | 95.895 | 462 | 120.908 | 409
21.317 | 0 | 23.099 | 1320 | 31.085 | 460 | 36.101 | 1199 | 41.117 | 588 | 51.082 | 736 | 71.08 | 621 | 96.093 | 464 | 121.106 | 540
21.515 | 0 | 23.297 | 1580 | 31.283 | 498 | 36.299 | 887 | 41.315 | 647 | 51.28 | 786 | 71.278 | 574 | 96.291 | 520 | 121.304 | 424
21.713 | 0 | 23.495 | 1678 | 31.481 | 516 | 36.497 | 1136 | 41.513 | 841 | 51.478 | 729 | 71.476 | 628 | 96.489 | 524 | 121.502 | 443
21.911 | 0 | 23.693 | 1801 | 31.679 | 442 | 36.695 | 1119 | 41.711 | 1034 | 51.676 | 655 | 71.674 | 813 | 96.687 | 571 | 121.7 | 509
22.109 | 0 | 23.891 | 1718 | 31.877 | 494 | 36.893 | 1033 | 41.909 | 1313 | 51.874 | 759 | 71.872 | 1130 | 96.885 | 543 | 121.898 | 495
22.307 | 0 | 24.089 | 1430 | 32.075 | 519 | 37.091 | 856 | 42.107 | 2193 | 52.072 | 646 | 72.07 | 835 | 97.083 | 623 | 122.096 | 471
22.505 | 0 | 24.287 | 1668 | 32.273 | 386 | 37.289 | 808 | 42.305 | 3820 | 52.27 | 613 | 72.268 | 1184 | 97.281 | 709 | 122.294 | 471
22.703 | 0 | 24.485 | 1582 | 32.471 | 438 | 37.487 | 984 | 42.503 | 6237 | 52.468 | 591 | 72.466 | 1204 | 97.479 | 550 | 122.492 | 490
22.901 | 0 | 24.683 | 1678 | 32.669 | 627 | 37.685 | 1025 | 42.701 | 10484 | 52.666 | 637 | 72.664 | 1577 | 97.677 | 614 | 122.69 | 434
23.099 | 0 | 24.881 | 1575 | 32.867 | 609 | 37.883 | 1073 | 42.899 | 15861 | 52.864 | 618 | 72.862 | 1864 | 97.875 | 519 | 122.888 | 574
23.297 | 0 | 25.079 | 1447 | 33.065 | 668 | 38.081 | 1163 | 43.097 | 21780 | 53.062 | 462 | 73.06 | 2535 | 98.073 | 576 | 123.086 | 473
23.495 | 0 | 25.277 | 1518 | 33.263 | 690 | 38.279 | 923 | 43.295 | 27678 | 53.26 | 634 | 73.258 | 3108 | 98.271 | 561 | 123.284 | 356
23.693 | 0 | 25.475 | 1454 | 33.461 | 672 | 38.477 | 960 | 43.493 | 33711 | 53.458 | 625 | 73.456 | 4103 | 98.469 | 484 | 123.482 | 281
23.891 | 0 | 25.673 | 1429 | 33.659 | 930 | 38.675 | 898 | 43.691 | 35467 | 53.656 | 654 | 73.654 | 5221 | 98.667 | 442 | 123.68 | 504
24.089 | 0 | 25.871 | 1412 | 33.857 | 837 | 38.873 | 783 | 43.889 | 36263 | 53.854 | 733 | 73.852 | 6325 | 98.865 | 449 | 123.878 | 335
24.287 | 0 | 26.069 | 1207 | 34.055 | 731 | 39.071 | 745 | 44.087 | 34089 | 54.052 | 775 | 74.05 | 7700 | 99.063 | 407 | 124.076 | 267
24.485 | 0 | 26.267 | 1341 | 34.253 | 659 | 39.269 | 715 | 44.285 | 28615 | 54.25 | 651 | 74.248 | 9561 | 99.261 | 407 | 124.274 | 344
24.683 | 0 | 26.465 | 1182 | 34.451 | 619 | 39.467 | 825 | 44.483 | 23229 | 54.448 | 676 | 74.446 | 10787 | 99.459 | 400 | 124.472 | 298
24.881 | 0 | 26.663 | 1060 | 34.649 | 502 | 39.665 | 813 | 44.681 | 17207 | 54.646 | 587 | 74.644 | 12540 | 99.657 | 322 | 124.67 | 350
25.079 | 0 | 26.861 | 1161 | 34.847 | 428 | 39.863 | 738 | 44.879 | 12383 | 54.844 | 653 | 74.842 | 13399 | 99.855 | 452 | 124.868 | 316
25.277 | 0 | 27.059 | 1251 | 35.045 | 383 | 40.061 | 783 | 45.077 | 8532 | 55.042 | 550 | 75.04 | 14478 | 100.053 | 455 | 125.066 | 287
25.475 | 0 | 27.257 | 1133 | 35.243 | 351 | 40.259 | 706 | 45.275 | 5320 | 55.24 | 641 | 75.238 | 14923 | 100.251 | 326 | 125.264 | 324
25.673 | 0 | 27.455 | 1115 | 35.441 | 234 | 40.457 | 721 | 45.473 | 3726 | 55.438 | 601 | 75.436 | 14996 | 100.449 | 452 | 125.462 | 255
25.871 | 0 | 27.653 | 1059 | 35.639 | 284 | 40.655 | 728 | 45.671 | 2396 | 55.636 | 640 | 75.634 | 14981 | 100.647 | 247 | 125.66 | 320
26.069 | 0 | 27.851 | 1041 | 35.837 | 211 | 40.853 | 642 | 45.869 | 1713 | 55.834 | 646 | 75.832 | 14147 | 100.845 | 398 | 125.858 | 264
26.267 | 0 | 28.049 | 919 | 36.035 | 258 | 41.051 | 574 | 46.067 | 1067 | 56.032 | 728 | 76.03 | 13558 | 101.043 | 257 | 126.056 | 205
26.465 | 0 | 28.247 | 918 | 36.233 | 244 | 41.249 | 587 | 46.265 | 838 | 56.23 | 657 | 76.228 | 11560 | 101.241 | 413 | 126.254 | 194
26.663 | 0 | 28.445 | 865 | 36.431 | 161 | 41.447 | 568 | 46.463 | 659 | 56.428 | 554 | 76.426 | 10527 | 101.439 | 359 | 126.452 | 268
26.861 | 0 | 28.643 | 877 | 36.629 | 260 | 41.645 | 521 | 46.661 | 455 | 56.626 | 541 | 76.624 | 9803 | 101.637 | 403 | 126.65 | 194
27.059 | 0 | 28.841 | 983 | 36.827 | 374 | 41.843 | 611 | 46.859 | 414 | 56.824 | 447 | 76.822 | 7758 | 101.835 | 377 | 126.848 | 159
27.257 | 0 | 29.039 | 793 | 37.025 | 242 | 42.041 | 1222 | 47.056 | 330 | 57.022 | 389 | 77.02 | 6481 | 102.033 | 309 | 127.046 | 210
27.455 | 0 | 29.237 | 715 | 37.223 | 355 | 42.239 | 1899 | 47.254 | 342 | 57.22 | 339 | 77.218 | 5261 | 102.231 | 404 | 127.244 | 136
27.653 | 0 | 29.435 | 635 | 37.421 | 310 | 42.437 | 3548 | 47.452 | 242 | 57.418 | 285 | 77.416 | 4454 | 102.429 | 396 | 127.442 | 186
27.851 | 0 | 29.633 | 645 | 37.619 | 399 | 42.635 | 7099 | 47.65 | 344 | 57.616 | 174 | 77.614 | 3869 | 102.627 | 352 | 127.64 | 140
28.049 | 0 | 29.831 | 445 | 37.817 | 262 | 42.833 | 10813 | 47.848 | 289 | 57.814 | 233 | 77.812 | 3096 | 102.825 | 436 | 127.838 | 152
28.247 | 0 | 30.029 | 473 | 38.015 | 406 | 43.031 | 16857 | 48.046 | 262 | 58.012 | 249 | 78.01 | 2487 | 103.023 | 450 | 128.036 | 150
28.445 | 0 | 30.227 | 509 | 38.213 | 284 | 43.229 | 22893 | 48.244 | 220 | 58.21 | 128 | 78.208 | 1955 | 103.221 | 531 | 128.234 | 161
28.643 | 0 | 30.425 | 438 | 38.411 | 242 | 43.427 | 29828 | 48.442 | 260 | 58.408 | 103 | 78.405 | 1671 | 103.419 | 780 | 128.432 | 164
28.841 | 0 | 30.623 | 368 | 38.609 | 333 | 43.625 | 36470 | 48.64 | 203 | 58.606 | 155 | 78.603 | 1271 | 103.617 | 892 | 128.63 | 92
29.039 | 0 | 30.821 | 370 | 38.807 | 227 | 43.823 | 38645 | 48.838 | 262 | 58.804 | 89 | 78.801 | 1093 | 103.815 | 1097 | 128.828 | 140
29.237 | 0 | 31.019 | 390 | 39.005 | 289 | 44.021 | 38871 | 49.036 | 202 | 59.002 | 126 | 78.999 | 741 | 104.013 | 1246 | 129.026 | 113
29.435 | 0 | 31.217 | 528 | 39.203 | 231 | 44.219 | 36309 | 49.234 | 257 | 59.2 | 44 | 79.197 | 743 | 104.211 | 1469 | 129.224 | 50
29.633 | 0 | 31.415 | 545 | 39.401 | 78 | 44.417 | 30844 | 49.432 | 142 | 59.398 | 70 | 79.395 | 654 | 104.409 | 1921 | 129.422 | 69
29.831 | 0 | 31.613 | 490 | 39.599 | 142 | 44.615 | 24738 | 49.63 | 128 | 59.596 | 89 | 79.593 | 562 | 104.607 | 2369 | 129.62 | 145
30.029 | 0 | 31.811 | 486 | 39.797 | 57 | 44.813 | 18666 | 49.828 | 176 | 59.794 | 110 | 79.791 | 403 | 104.805 | 2732 | 129.818 | 71
30.227 | 0 | 32.009 | 632 | 39.995 | 42 | 45.011 | 12738 | 50.026 | 178 | 59.992 | 117 | 79.989 | 332 | 105.003 | 3045 | 130.016 | 102
30.425 | 0 | 32.207 | 784 | 40.193 | 93 | 45.209 | 8757 | 50.224 | 183 | 60.19 | 56 | 80.187 | 213 | 105.201 | 3629 | 130.214 | 53
30.623 | 0 | 32.405 | 655 | 40.391 | 70 | 45.407 | 6074 | 50.422 | 240 | 60.388 | 43 | 80.385 | 277 | 105.399 | 4390 | 130.412 | 51
30.821 | 0 | 32.603 | 592 | 40.589 | 39 | 45.605 | 3492 | 50.62 | 107 | 60.586 | 53 | 80.583 | 161 | 105.597 | 4813 | 130.61 | 61
31.019 | 0 | 32.801 | 691 | 40.787 | 68 | 45.803 | 2458 | 50.818 | 193 | 60.784 | 73 | 80.781 | 165 | 105.795 | 5840 | 130.808 | 120
31.217 | 0 | 32.999 | 724 | 40.985 | 96 | 46.001 | 1605 | 51.016 | 154 | 60.982 | 113 | 80.979 | 110 | 105.993 | 5945 | 131.006 | 99
31.415 | 0 | 33.197 | 666 | 41.183 | 116 | 46.199 | 1017 | 51.214 | 120 | 61.18 | 90 | 81.177 | 117 | 106.191 | 6651 | 131.204 | 92
31.613 | 0 | 33.395 | 915 | 41.381 | 90 | 46.397 | 388 | 51.412 | 144 | 61.378 | 88 | 81.375 | 111 | 106.389 | 7084 | 131.402 | 71
31.811 | 0 | 33.593 | 828 | 41.579 | 100 | 46.595 | 300 | 51.61 | 120 | 61.576 | 108 | 81.573 | 78 | 106.587 | 7595 | 131.6 | 83
32.009 | 0 | 33.791 | 1115 | 41.777 | 222 | 46.793 | 297 | 51.808 | 66 | 61.774 | 93 | 81.771 | 112 | 106.785 | 8068 | 131.798 | 95
32.207 | 0 | 33.989 | 1064 | 41.975 | 666 | 46.99 | 83 | 52.006 | 58 | 61.972 | 238 | 81.969 | 70 | 106.983 | 8324 | 131.996 | 111
32.405 | 0 | 34.187 | 978 | 42.173 | 961 | 47.188 | 115 | 52.204 | 16 | 62.17 | 185 | 82.167 | 64 | 107.181 | 7685 | 132.194 | 166
32.603 | 0 | 34.385 | 1052 | 42.371 | 2042 | 47.386 | 77 | 52.402 | 31 | 62.368 | 142 | 82.365 | 58 | 107.379 | 7697 | 132.392 | 81
32.801 | 0 | 34.583 | 872 | 42.569 | 3856 | 47.584 | 101 | 52.6 | 12 | 62.566 | 104 | 82.563 | 68 | 107.577 | 7810 | 132.59 | 104
32.999 | 0 | 34.781 | 756 | 42.767 | 6760 | 47.782 | 63 | 52.798 | 11 | 62.764 | 207 | 82.761 | 114 | 107.775 | 7800 | 132.788 | 140
33.197 | 0 | 34.979 | 686 | 42.965 | 11588 | 47.98 | 61 | 52.996 | 25 | 62.962 | 94 | 82.959 | 121 | 107.973 | 7033 | 132.986 | 125
33.395 | 0 | 35.177 | 491 | 43.163 | 17008 | 48.178 | 54 | 53.194 | 20 | 63.16 | 190 | 83.157 | 41 | 108.171 | 6813 | 133.184 | 153
33.593 | 0 | 35.375 | 445 | 43.361 | 25097 | 48.376 | 66 | 53.392 | 8 | 63.358 | 87 | 83.355 | 67 | 108.369 | 6631 | 133.382 | 220
33.791 | 0 | 35.573 | 428 | 43.559 | 32619 | 48.574 | 38 | 53.59 | 21 | 63.556 | 158 | 83.553 | 100 | 108.567 | 5555 | 133.58 | 282
33.989 | 0 | 35.771 | 303 | 43.757 | 38637 | 48.772 | 88 | 53.788 | 5 | 63.754 | 131 | 83.751 | 63 | 108.765 | 5379 | 133.778 | 304
34.187 | 0 | 35.969 | 247 | 43.955 | 43116 | 48.97 | 37 | 53.986 | 3 | 63.952 | 123 | 83.949 | 99 | 108.963 | 4773 | 133.976 | 315
34.385 | 0 | 36.167 | 308 | 44.153 | 43384 | 49.168 | 31 | 54.184 | 4 | 64.15 | 128 | 84.147 | 152 | 109.161 | 3866 | 134.174 | 402
34.583 | 0 | 36.365 | 327 | 44.351 | 41844 | 49.366 | 45 | 54.382 | 6 | 64.348 | 65 | 84.345 | 90 | 109.359 | 3608 | 134.372 | 427
34.781 | 0 | 36.563 | 356 | 44.549 | 35345 | 49.564 | 31 | 54.58 | 4 | 64.546 | 73 | 84.543 | 127 | 109.556 | 2971 | 134.57 | 601
34.979 | 0 | 36.761 | 360 | 44.747 | 28380 | 49.762 | 38 | 54.778 | 8 | 64.744 | 88 | 84.741 | 45 | 109.754 | 2583 | 134.768 | 572
35.177 | 0 | 36.959 | 407 | 44.945 | 22182 | 49.96 | 39 | 54.976 | 0 | 64.942 | 242 | 84.939 | 70 | 109.952 | 2287 | 134.966 | 838
35.375 | 0 | 37.157 | 406 | 45.143 | 15416 | 50.158 | 42 | 55.174 | 0 | 65.14 | 168 | 85.137 | 54 | 110.15 | 1902 | 135.164 | 813
35.573 | 0 | 37.355 | 339 | 45.341 | 10491 | 50.356 | 47 | 55.372 | 0 | 65.338 | 131 | 85.335 | 105 | 110.348 | 1701 | 135.362 | 898
35.771 | 0 | 37.553 | 379 | 45.539 | 6886 | 50.554 | 29 | 55.57 | 0 | 65.536 | 82 | 85.533 | 36 | 110.546 | 1414 | 135.56 | 1249
35.969 | 0 | 37.751 | 480 | 45.737 | 4628 | 50.752 | 21 | 55.768 | 0 | 65.734 | 45 | 85.731 | 45 | 110.744 | 1179 | 135.758 | 1474
36.167 | 0 | 37.949 | 363 | 45.935 | 2893 | 50.95 | 43 | 55.966 | 5 | 65.932 | 141 | 85.929 | 109 | 110.942 | 1185 | 135.956 | 1484
36.365 | 0 | 38.147 | 350 | 46.133 | 1777 | 51.148 | 15 | 56.164 | 0 | 66.13 | 118 | 86.127 | 59 | 111.14 | 965 | 136.154 | 1842
36.563 | 0 | 38.345 | 360 | 46.331 | 1022 | 51.346 | 5 | 56.362 | 0 | 66.328 | 65 | 86.325 | 59 | 111.338 | 766 | 136.352 | 2052
36.761 | 0 | 38.543 | 334 | 46.529 | 642 | 51.544 | 27 | 56.56 | 0 | 66.526 | 103 | 86.523 | 108 | 111.536 | 702 | 136.55 | 2365
36.959 | 0 | 38.741 | 383 | 46.727 | 497 | 51.742 | 8 | 56.758 | 0 | 66.724 | 63 | 86.721 | 34 | 111.734 | 735 | 136.748 | 2620
37.157 | 0 | 38.939 | 390 | 46.924 | 194 | 51.94 | 19 | 56.956 | 0 | 66.922 | 78 | 86.919 | 24 | 111.932 | 490 | 136.946 | 2690
37.355 | 0 | 39.137 | 343 | 47.122 | 151 | 52.138 | 13 | 57.154 | 0 | 67.12 | 20 | 87.117 | 21 | 112.13 | 372 | 137.144 | 3254
37.553 | 0 | 39.335 | 308 | 47.32 | 94 | 52.336 | 17 | 57.352 | 0 | 67.318 | 38 | 87.315 | 50 | 112.328 | 353 | 137.342 | 3519
37.751 | 0 | 39.533 | 263 | 47.518 | 85 | 52.534 | 12 | 57.55 | 0 | 67.516 | 34 | 87.513 | 20 | 112.526 | 296 | 137.54 | 3883
37.949 | 0 | 39.731 | 205 | 47.716 | 63 | 52.732 | 0 | 57.748 | 0 | 67.714 | 26 | 87.711 | 7 | 112.724 | 236 | 137.738 | 3891
38.147 | 0 | 39.929 | 122 | 47.914 | 89 | 52.93 | 0 | 57.946 | 3 | 67.912 | 15 | 87.909 | 6 | 112.922 | 199 | 137.936 | 4175
38.345 | 0 | 40.127 | 153 | 48.112 | 101 | 53.128 | 15 | 58.144 | 0 | 68.11 | 51 | 88.107 | 31 | 113.12 | 132 | 138.134 | 4747
38.543 | 0 | 40.325 | 59 | 48.31 | 86 | 53.326 | 37 | 58.342 | 0 | 68.308 | 10 | 88.305 | 19 | 113.318 | 92 | 138.332 | 4895
38.741 | 0 | 40.523 | 55 | 48.508 | 71 | 53.524 | 19 | 58.54 | 0 | 68.506 | 19 | 88.503 | 12 | 113.516 | 78 | 138.53 | 4660
38.939 | 0 | 40.721 | 150 | 48.706 | 58 | 53.722 | 0 | 58.738 | 0 | 68.704 | 26 | 88.701 | 12 | 113.714 | 139 | 138.728 | 4876
39.137 | 0 | 40.919 | 44 | 48.904 | 78 | 53.92 | 12 | 58.936 | 0 | 68.902 | 23 | 88.899 | 0 | 113.912 | 105 | 138.926 | 4560
39.335 | 0 | 41.117 | 57 | 49.102 | 31 | 54.118 | 4 | 59.134 | 0 | 69.1 | 31 | 89.097 | 6 | 114.11 | 39 | 139.124 | 4746
39.533 | 0 | 41.315 | 64 | 49.3 | 35 | 54.316 | 3 | 59.332 | 0 | 69.298 | 29 | 89.295 | 20 | 114.308 | 58 | 139.322 | 4851
39.731 | 0 | 41.513 | 139 | 49.498 | 44 | 54.514 | 0 | 59.53 | 0 | 69.496 | 22 | 89.493 | 3 | 114.506 | 30 | 139.52 | 4613
39.929 | 0 | 41.711 | 176 | 49.696 | 29 | 54.712 | 2 | 59.728 | 0 | 69.694 | 0 | 89.691 | 14 | 114.704 | 47 | 139.718 | 4360
40.127 | 0 | 41.909 | 168 | 49.894 | 42 | 54.91 | 0 | 59.926 | 0 | 69.892 | 40 | 89.889 | 0 | 114.902 | 23 | 139.916 | 4360
40.325 | 0 | 42.107 | 544 | 50.092 | 34 | 55.108 | 4 | 60.124 | 0 | 70.09 | 6 | 90.087 | 6 | 115.1 | 27 | 140.114 | 4183
40.523 | 0 | 42.305 | 839 | 50.29 | 8 | 55.306 | 0 | 60.322 | 0 | 70.288 | 15 | 90.285 | 6 | 115.298 | 23 | 140.312 | 3762
40.721 | 0 | 42.503 | 1695 | 50.488 | 44 | 55.504 | 5 | 60.52 | 0 | 70.486 | 19 | 90.483 | 0 | 115.496 | 25 | 140.51 | 3912
40.919 | 0 | 42.701 | 3768 | 50.686 | 19 | 55.702 | 0 | 60.718 | 0 | 70.684 | 25 | 90.681 | 3 | 115.694 | 16 | 140.708 | 3413
41.117 | 0 | 42.899 | 6691 | 50.884 | 34 | 55.9 | 0 | 60.916 | 0 | 70.882 | 28 | 90.879 | 3 | 115.892 | 18 | 140.906 | 3299
41.315 | 0 | 43.097 | 11211 | 51.082 | 42 | 56.098 | 10 | 61.114 | 0 | 71.08 | 0 | 91.077 | 5 | 116.09 | 11 | 141.103 | 2924
41.513 | 0 | 43.295 | 17446 | 51.28 | 26 | 56.296 | 0 | 61.312 | 9 | 71.278 | 19 | 91.275 | 4 | 116.288 | 9 | 141.301 | 2732
41.711 | 0 | 43.493 | 26064 | 51.478 | 34 | 56.494 | 5 | 61.51 | 0 | 71.476 | 45 | 91.473 | 2 | 116.486 | 16 | 141.499 | 2510
41.909 | 0 | 43.691 | 35158 | 51.676 | 53 | 56.692 | 2 | 61.708 | 0 | 71.674 | 49 | 91.671 | 8 | 116.684 | 22 | 141.697 | 2208
42.107 | 0 | 43.889 | 43010 | 51.874 | 13 | 56.89 | 8 | 61.906 | 0 | 71.872 | 18 | 91.869 | 28 | 116.882 | 10 | 141.895 | 1915
42.305 | 0 | 44.087 | 48802 | 52.072 | 25 | 57.088 | 0 | 62.104 | 4 | 72.07 | 51 | 92.067 | 0 | 117.08 | 0 | 142.093 | 1631
42.503 | 0 | 44.285 | 51223 | 52.27 | 11 | 57.286 | 0 | 62.302 | 0 | 72.268 | 36 | 92.265 | 3 | 117.278 | 28 | 142.291 | 1503
42.701 | 0 | 44.483 | 49769 | 52.468 | 14 | 57.484 | 0 | 62.5 | 0 | 72.466 | 23 | 92.463 | 0 | 117.476 | 21 | 142.489 | 1334
42.899 | 0 | 44.681 | 43135 | 52.666 | 7 | 57.682 | 0 | 62.698 | 0 | 72.664 | 53 | 92.661 | 8 | 117.674 | 20 | 142.687 | 1136
43.097 | 0 | 44.879 | 36562 | 52.864 | 21 | 57.88 | 2 | 62.896 | 6 | 72.862 | 20 | 92.859 | 16 | 117.872 | 22 | 142.885 | 1105
43.295 | 0 | 45.077 | 27840 | 53.062 | 27 | 58.078 | 0 | 63.094 | 0 | 73.06 | 40 | 93.057 | 2 | 118.07 | 11 | 143.083 | 890
43.493 | 0 | 45.275 | 20830 | 53.26 | 16 | 58.276 | 0 | 63.292 | 0 | 73.258 | 30 | 93.255 | 17 | 118.268 | 11 | 143.281 | 960
43.691 | 0 | 45.473 | 13822 | 53.458 | 29 | 58.474 | 3 | 63.49 | 0 | 73.456 | 7 | 93.453 | 9 | 118.466 | 5 | 143.479 | 739
43.889 | 0 | 45.671 | 9344 | 53.656 | 4 | 58.672 | 0 | 63.688 | 0 | 73.654 | 23 | 93.651 | 6 | 118.664 | 21 | 143.677 | 742
44.087 | 0 | 45.869 | 6123 | 53.854 | 14 | 58.87 | 0 | 63.886 | 0 | 73.852 | 3 | 93.849 | 6 | 118.862 | 2 | 143.875 | 494
44.285 | 0 | 46.067 | 3690 | 54.052 | 34 | 59.068 | 0 | 64.084 | 0 | 74.05 | 4 | 94.047 | 15 | 119.06 | 13 | 144.073 | 634
44.483 | 0 | 46.265 | 2423 | 54.25 | 22 | 59.266 | 0 | 64.282 | 0 | 74.248 | 0 | 94.245 | 6 | 119.258 | 7 | 144.271 | 479
44.681 | 0 | 46.463 | 1504 | 54.448 | 6 | 59.464 | 0 | 64.48 | 0 | 74.446 | 0 | 94.443 | 8 | 119.456 | 0 | 144.469 | 369
44.879 | 0 | 46.661 | 797 | 54.646 | 5 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 94.641 | 3 | 119.654 | 17 | 144.667 | 285
45.077 | 0 | 46.859 | 483 | 54.844 | 3 | 59.86 | 0 | 64.876 | 0 | 74.842 | 0 | 94.839 | 8 | 119.852 | 24 | 144.865 | 306
45.275 | 0 | 47.056 | 347 | 55.042 | 0 | 60.058 | 3 | 65.074 | 0 | 75.04 | 0 | 95.037 | 5 | 120.05 | 37 | 145.063 | 281
45.473 | 0 | 47.254 | 169 | 55.24 | 11 | 60.256 | 0 | 65.272 | 0 | 75.238 | 0 | 95.235 | 6 | 120.248 | 6 | 145.261 | 199
45.671 | 0 | 47.452 | 149 | 55.438 | 0 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 95.433 | 13 | 120.446 | 22 | 145.459 | 165
45.869 | 0 | 47.65 | 170 | 55.636 | 0 | 60.652 | 0 | 65.668 | 0 | 75.634 | 0 | 95.631 | 0 | 120.644 | 7 | 145.657 | 85
46.067 | 0 | 47.848 | 97 | 55.834 | 0 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 95.829 | 5 | 120.842 | 0 | 145.855 | 41
46.265 | 0 | 48.046 | 98 | 56.032 | 2 | 61.048 | 0 | 66.064 | 0 | 76.03 | 0 | 96.027 | 0 | 121.04 | 10 | 146.053 | 34
46.463 | 0 | 48.244 | 101 | 56.23 | 0 | 61.246 | 0 | 66.262 | 0 | 76.228 | 0 | 96.225 | 0 | 121.238 | 2 | 146.251 | 15
46.661 | 0 | 48.442 | 82 | 56.428 | 12 | 61.444 | 4 | 66.46 | 0 | 76.426 | 0 | 96.423 | 0 | 121.436 | 7 | 146.449 | 4
46.859 | 0 | 48.64 | 55 | 56.626 | 0 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 96.621 | 0 | 121.634 | 6 | 146.647 | 3
47.056 | 0 | 48.838 | 37 | 56.824 | 0 | 61.84 | 3 | 66.856 | 0 | 76.822 | 0 | 96.819 | 0 | 121.832 | 19 | 146.845 | 3
47.254 | 0 | 49.036 | 75 | 57.022 | 0 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 97.017 | 0 | 122.03 | 12 | 147.043 | 0
47.452 | 0 | 49.234 | 32 | 57.22 | 0 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 97.215 | 5 | 122.228 | 9 | 147.241 | 0
47.65 | 0 | 49.432 | 47 | 57.418 | 0 | 62.434 | 0 | 67.45 | 4 | 77.416 | 0 | 97.413 | 0 | 122.426 | 3 | 147.439 | 0
47.848 | 0 | 49.63 | 43 | 57.616 | 0 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 97.611 | 0 | 122.624 | 8 | 147.637 | 0
48.046 | 0 | 49.828 | 33 | 57.814 | 0 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 97.809 | 4 | 122.822 | 6 | 147.835 | 0
48.244 | 0 | 50.026 | 38 | 58.012 | 0 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 98.007 | 0 | 123.02 | 3 | 148.033 | 0
48.442 | 0 | 50.224 | 9 | 58.21 | 3 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 98.205 | 0 | 123.218 | 0 | 148.231 | 0
48.64 | 0 | 50.422 | 18 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 98.403 | 0 | 123.416 | 12 | 148.429 | 0
48.838 | 0 | 50.62 | 36 | 58.606 | 4 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 98.601 | 0 | 123.614 | 0 | 148.627 | 0
49.036 | 0 | 50.818 | 40 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 78.801 | 0 | 98.799 | 0 | 123.812 | 6 | 148.825 | 0
49.234 | 0 | 51.016 | 35 | 59.002 | 7 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 98.997 | 8 | 124.01 | 0 | 149.023 | 0
49.432 | 0 | 51.214 | 19 | 59.2 | 0 | 64.216 | 0 | 69.232 | 0 | 79.197 | 0 | 99.195 | 0 | 124.208 | 6 | 149.221 | 0
|  | 51.412 | 28 | 59.398 | 7 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 99.393 | 0 | 124.406 | 3 | 149.419 | 0

## Analyte Polyalanine n14
0 |  | 2 |  | 10 |  | 15 |  | 20 |  | 30 |  | 32 |  | 50 |  | 75 |  | 100
Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity | Arrival time | Intensity
10.23 | 0 | 12.012 | 0 | 19.997 | 0 | 25.013 | 0 | 30.029 | 0 | 39.995 | 0 | 41.975 | 0 | 59.992 | 0 | 85.005 | 0 | 110.018 | 0
10.428 | 0 | 12.21 | 0 | 20.195 | 0 | 25.211 | 0 | 30.227 | 0 | 40.193 | 0 | 42.173 | 0 | 60.19 | 0 | 85.203 | 0 | 110.216 | 0
10.626 | 0 | 12.408 | 0 | 20.393 | 0 | 25.409 | 0 | 30.425 | 0 | 40.391 | 0 | 42.371 | 0 | 60.388 | 0 | 85.401 | 0 | 110.414 | 0
10.824 | 0 | 12.606 | 0 | 20.591 | 0 | 25.607 | 0 | 30.623 | 0 | 40.589 | 0 | 42.569 | 0 | 60.586 | 0 | 85.599 | 0 | 110.612 | 0
11.022 | 0 | 12.804 | 0 | 20.789 | 0 | 25.805 | 0 | 30.821 | 0 | 40.787 | 0 | 42.767 | 0 | 60.784 | 0 | 85.797 | 0 | 110.81 | 0
11.22 | 29 | 13.002 | 0 | 20.987 | 0 | 26.003 | 0 | 31.019 | 0 | 40.985 | 0 | 42.965 | 0 | 60.982 | 0 | 85.995 | 0 | 111.008 | 0
11.418 | 532 | 13.2 | 0 | 21.185 | 32 | 26.201 | 6 | 31.217 | 12 | 41.183 | 4 | 43.163 | 0 | 61.18 | 31 | 86.193 | 14 | 111.206 | 5
11.616 | 2431 | 13.398 | 0 | 21.383 | 68 | 26.399 | 186 | 31.415 | 6 | 41.381 | 131 | 43.361 | 62 | 61.378 | 130 | 86.391 | 143 | 111.404 | 123
11.814 | 3494 | 13.596 | 0 | 21.581 | 18 | 26.597 | 372 | 31.613 | 78 | 41.579 | 296 | 43.559 | 146 | 61.576 | 239 | 86.589 | 167 | 111.602 | 228
12.012 | 2633 | 13.794 | 0 | 21.779 | 40 | 26.795 | 411 | 31.811 | 68 | 41.777 | 206 | 43.757 | 140 | 61.774 | 368 | 86.787 | 424 | 111.8 | 336
12.21 | 4995 | 13.992 | 0 | 21.977 | 138 | 26.993 | 365 | 32.009 | 130 | 41.975 | 351 | 43.955 | 270 | 61.972 | 435 | 86.985 | 300 | 111.998 | 1080
12.408 | 16317 | 14.19 | 0 | 22.175 | 223 | 27.191 | 579 | 32.207 | 233 | 42.173 | 416 | 44.153 | 613 | 62.17 | 379 | 87.183 | 300 | 112.196 | 2012
12.606 | 135461 | 14.388 | 5 | 22.373 | 76 | 27.389 | 515 | 32.405 | 155 | 42.371 | 316 | 44.351 | 838 | 62.368 | 379 | 87.381 | 267 | 112.394 | 1677
12.804 | 155904 | 14.586 | 0 | 22.571 | 130 | 27.587 | 377 | 32.603 | 143 | 42.569 | 196 | 44.549 | 1155 | 62.566 | 309 | 87.579 | 174 | 112.592 | 1582
13.002 | 23915 | 14.784 | 0 | 22.769 | 196 | 27.785 | 437 | 32.801 | 152 | 42.767 | 345 | 44.747 | 1803 | 62.764 | 341 | 87.777 | 255 | 112.79 | 1785
13.2 | 668 | 14.982 | 0 | 22.967 | 229 | 27.983 | 325 | 32.999 | 126 | 42.965 | 245 | 44.945 | 3180 | 62.962 | 368 | 87.975 | 299 | 112.988 | 1870
13.398 | 34 | 15.18 | 0 | 23.165 | 455 | 28.181 | 327 | 33.197 | 146 | 43.163 | 335 | 45.143 | 4190 | 63.16 | 338 | 88.173 | 271 | 113.186 | 2263
13.596 | 0 | 15.378 | 0 | 23.363 | 663 | 28.379 | 393 | 33.395 | 242 | 43.361 | 247 | 45.341 | 6531 | 63.358 | 298 | 88.371 | 348 | 113.384 | 2092
13.794 | 0 | 15.576 | 0 | 23.561 | 880 | 28.577 | 355 | 33.593 | 148 | 43.559 | 253 | 45.539 | 9051 | 63.556 | 253 | 88.569 | 251 | 113.582 | 2733
13.992 | 0 | 15.773 | 0 | 23.759 | 956 | 28.775 | 228 | 33.791 | 260 | 43.757 | 302 | 45.737 | 11844 | 63.754 | 336 | 88.767 | 236 | 113.78 | 2884
14.19 | 0 | 15.971 | 0 | 23.957 | 954 | 28.973 | 185 | 33.989 | 189 | 43.955 | 366 | 45.935 | 15067 | 63.952 | 499 | 88.965 | 255 | 113.978 | 3102
14.388 | 0 | 16.169 | 0 | 24.155 | 857 | 29.171 | 300 | 34.187 | 183 | 44.153 | 433 | 46.133 | 16852 | 64.15 | 373 | 89.163 | 241 | 114.176 | 3542
14.586 | 0 | 16.367 | 0 | 24.353 | 967 | 29.369 | 269 | 34.385 | 290 | 44.351 | 625 | 46.331 | 18508 | 64.348 | 273 | 89.361 | 277 | 114.374 | 3608
14.784 | 0 | 16.565 | 0 | 24.551 | 1070 | 29.567 | 281 | 34.583 | 340 | 44.549 | 1045 | 46.529 | 18671 | 64.546 | 254 | 89.559 | 264 | 114.572 | 3625
14.982 | 0 | 16.763 | 0 | 24.749 | 905 | 29.765 | 318 | 34.781 | 282 | 44.747 | 1646 | 46.727 | 17021 | 64.744 | 264 | 89.757 | 254 | 114.77 | 3843
15.18 | 0 | 16.961 | 0 | 24.947 | 822 | 29.963 | 231 | 34.979 | 431 | 44.945 | 2780 | 46.924 | 14792 | 64.942 | 257 | 89.955 | 313 | 114.968 | 3902
15.378 | 0 | 17.159 | 0 | 25.145 | 823 | 30.161 | 195 | 35.177 | 436 | 45.143 | 3831 | 47.122 | 13061 | 65.14 | 250 | 90.153 | 230 | 115.166 | 4214
15.576 | 0 | 17.357 | 0 | 25.343 | 784 | 30.359 | 183 | 35.375 | 721 | 45.341 | 6168 | 47.32 | 10258 | 65.338 | 277 | 90.351 | 226 | 115.364 | 3854
15.773 | 0 | 17.555 | 0 | 25.541 | 810 | 30.557 | 189 | 35.573 | 725 | 45.539 | 8888 | 47.518 | 7543 | 65.536 | 229 | 90.549 | 238 | 115.562 | 3902
15.971 | 0 | 17.753 | 0 | 25.739 | 753 | 30.755 | 221 | 35.771 | 753 | 45.737 | 11873 | 47.716 | 5684 | 65.734 | 279 | 90.747 | 244 | 115.76 | 4012
16.169 | 0 | 17.951 | 0 | 25.937 | 644 | 30.953 | 165 | 35.969 | 835 | 45.935 | 14308 | 47.914 | 4055 | 65.932 | 242 | 90.945 | 226 | 115.958 | 4199
16.367 | 0 | 18.149 | 0 | 26.135 | 624 | 31.151 | 183 | 36.167 | 765 | 46.133 | 16912 | 48.112 | 2811 | 66.13 | 215 | 91.143 | 272 | 116.156 | 3454
16.565 | 0 | 18.347 | 0 | 26.333 | 601 | 31.349 | 238 | 36.365 | 846 | 46.331 | 18232 | 48.31 | 1805 | 66.328 | 215 | 91.341 | 287 | 116.354 | 3699
16.763 | 0 | 18.545 | 0 | 26.531 | 670 | 31.547 | 236 | 36.563 | 829 | 46.529 | 18602 | 48.508 | 1347 | 66.526 | 244 | 91.539 | 221 | 116.552 | 3242
16.961 | 0 | 18.743 | 5 | 26.729 | 556 | 31.745 | 297 | 36.761 | 705 | 46.727 | 18013 | 48.706 | 1066 | 66.724 | 309 | 91.737 | 210 | 116.75 | 3017
17.159 | 0 | 18.941 | 3 | 26.927 | 546 | 31.943 | 167 | 36.959 | 627 | 46.924 | 16159 | 48.904 | 761 | 66.922 | 300 | 91.935 | 225 | 116.948 | 3058
17.357 | 0 | 19.139 | 0 | 27.125 | 483 | 32.141 | 188 | 37.157 | 604 | 47.122 | 13544 | 49.102 | 648 | 67.12 | 168 | 92.133 | 264 | 117.146 | 2454
17.555 | 0 | 19.337 | 0 | 27.323 | 640 | 32.339 | 136 | 37.355 | 650 | 47.32 | 10955 | 49.3 | 466 | 67.318 | 202 | 92.331 | 181 | 117.344 | 2126
17.753 | 0 | 19.535 | 0 | 27.521 | 498 | 32.537 | 153 | 37.553 | 552 | 47.518 | 8590 | 49.498 | 503 | 67.516 | 364 | 92.529 | 238 | 117.542 | 1926
17.951 | 0 | 19.733 | 0 | 27.719 | 464 | 32.735 | 112 | 37.751 | 650 | 47.716 | 6222 | 49.696 | 422 | 67.714 | 227 | 92.727 | 216 | 117.74 | 1845
18.149 | 0 | 19.931 | 0 | 27.917 | 462 | 32.933 | 113 | 37.949 | 564 | 47.914 | 4308 | 49.894 | 463 | 67.912 | 255 | 92.925 | 192 | 117.938 | 1643
18.347 | 0 | 20.129 | 0 | 28.115 | 410 | 33.131 | 213 | 38.147 | 520 | 48.112 | 2977 | 50.092 | 447 | 68.11 | 249 | 93.123 | 183 | 118.136 | 1413
18.545 | 0 | 20.327 | 0 | 28.313 | 511 | 33.329 | 215 | 38.345 | 527 | 48.31 | 2092 | 50.29 | 365 | 68.308 | 263 | 93.321 | 230 | 118.334 | 1332
18.743 | 0 | 20.525 | 7 | 28.511 | 483 | 33.527 | 126 | 38.543 | 532 | 48.508 | 1622 | 50.488 | 532 | 68.506 | 278 | 93.519 | 147 | 118.532 | 1191
18.941 | 0 | 20.723 | 7 | 28.709 | 373 | 33.725 | 253 | 38.741 | 448 | 48.706 | 1215 | 50.686 | 431 | 68.704 | 323 | 93.717 | 251 | 118.73 | 965
19.139 | 0 | 20.921 | 21 | 28.907 | 315 | 33.923 | 202 | 38.939 | 458 | 48.904 | 819 | 50.884 | 375 | 68.902 | 287 | 93.915 | 255 | 118.928 | 935
19.337 | 0 | 21.119 | 37 | 29.105 | 231 | 34.121 | 193 | 39.137 | 387 | 49.102 | 662 | 51.082 | 597 | 69.1 | 193 | 94.113 | 162 | 119.126 | 918
19.535 | 0 | 21.317 | 75 | 29.303 | 346 | 34.319 | 131 | 39.335 | 388 | 49.3 | 539 | 51.28 | 537 | 69.298 | 274 | 94.311 | 195 | 119.324 | 692
19.733 | 0 | 21.515 | 78 | 29.501 | 341 | 34.517 | 257 | 39.533 | 503 | 49.498 | 463 | 51.478 | 469 | 69.496 | 232 | 94.509 | 198 | 119.522 | 580
19.931 | 0 | 21.713 | 75 | 29.699 | 285 | 34.715 | 285 | 39.731 | 405 | 49.696 | 448 | 51.676 | 427 | 69.694 | 217 | 94.707 | 219 | 119.72 | 471
20.129 | 0 | 21.911 | 266 | 29.897 | 283 | 34.913 | 432 | 39.929 | 531 | 49.894 | 441 | 51.874 | 422 | 69.892 | 309 | 94.905 | 210 | 119.918 | 481
20.327 | 0 | 22.109 | 252 | 30.095 | 334 | 35.111 | 455 | 40.127 | 490 | 50.092 | 366 | 52.072 | 446 | 70.09 | 326 | 95.103 | 264 | 120.116 | 455
20.525 | 0 | 22.307 | 250 | 30.293 | 296 | 35.309 | 469 | 40.325 | 539 | 50.29 | 397 | 52.27 | 375 | 70.288 | 307 | 95.301 | 221 | 120.314 | 416
20.723 | 0 | 22.505 | 140 | 30.491 | 186 | 35.507 | 713 | 40.523 | 411 | 50.488 | 449 | 52.468 | 246 | 70.486 | 305 | 95.499 | 267 | 120.512 | 321
20.921 | 0 | 22.703 | 90 | 30.689 | 229 | 35.705 | 854 | 40.721 | 428 | 50.686 | 428 | 52.666 | 293 | 70.684 | 368 | 95.697 | 254 | 120.71 | 399
21.119 | 0 | 22.901 | 189 | 30.887 | 208 | 35.903 | 815 | 40.919 | 398 | 50.884 | 486 | 52.864 | 264 | 70.882 | 339 | 95.895 | 212 | 120.908 | 282
21.317 | 0 | 23.099 | 204 | 31.085 | 196 | 36.101 | 897 | 41.117 | 463 | 51.082 | 454 | 53.062 | 258 | 71.08 | 330 | 96.093 | 308 | 121.106 | 335
21.515 | 0 | 23.297 | 340 | 31.283 | 185 | 36.299 | 966 | 41.315 | 473 | 51.28 | 577 | 53.26 | 301 | 71.278 | 370 | 96.291 | 287 | 121.304 | 220
21.713 | 0 | 23.495 | 506 | 31.481 | 205 | 36.497 | 960 | 41.513 | 486 | 51.478 | 426 | 53.458 | 274 | 71.476 | 312 | 96.489 | 323 | 121.502 | 316
21.911 | 0 | 23.693 | 858 | 31.679 | 330 | 36.695 | 720 | 41.711 | 403 | 51.676 | 336 | 53.656 | 332 | 71.674 | 313 | 96.687 | 315 | 121.7 | 260
22.109 | 0 | 23.891 | 983 | 31.877 | 193 | 36.893 | 775 | 41.909 | 370 | 51.874 | 450 | 53.854 | 357 | 71.872 | 377 | 96.885 | 337 | 121.898 | 190
22.307 | 0 | 24.089 | 872 | 32.075 | 198 | 37.091 | 614 | 42.107 | 395 | 52.072 | 386 | 54.052 | 328 | 72.07 | 366 | 97.083 | 336 | 122.096 | 243
22.505 | 0 | 24.287 | 1000 | 32.273 | 156 | 37.289 | 498 | 42.305 | 345 | 52.27 | 278 | 54.25 | 333 | 72.268 | 336 | 97.281 | 220 | 122.294 | 288
22.703 | 0 | 24.485 | 998 | 32.471 | 122 | 37.487 | 506 | 42.503 | 276 | 52.468 | 338 | 54.448 | 333 | 72.466 | 348 | 97.479 | 315 | 122.492 | 270
22.901 | 0 | 24.683 | 1047 | 32.669 | 109 | 37.685 | 512 | 42.701 | 247 | 52.666 | 294 | 54.646 | 353 | 72.664 | 368 | 97.677 | 287 | 122.69 | 314
23.099 | 0 | 24.881 | 1177 | 32.867 | 171 | 37.883 | 527 | 42.899 | 289 | 52.864 | 338 | 54.844 | 278 | 72.862 | 355 | 97.875 | 243 | 122.888 | 279
23.297 | 0 | 25.079 | 1010 | 33.065 | 137 | 38.081 | 521 | 43.097 | 277 | 53.062 | 304 | 55.042 | 247 | 73.06 | 352 | 98.073 | 251 | 123.086 | 233
23.495 | 0 | 25.277 | 845 | 33.263 | 179 | 38.279 | 569 | 43.295 | 250 | 53.26 | 325 | 55.24 | 317 | 73.258 | 459 | 98.271 | 229 | 123.284 | 135
23.693 | 0 | 25.475 | 908 | 33.461 | 198 | 38.477 | 532 | 43.493 | 259 | 53.458 | 364 | 55.438 | 237 | 73.456 | 292 | 98.469 | 254 | 123.482 | 284
23.891 | 0 | 25.673 | 979 | 33.659 | 142 | 38.675 | 450 | 43.691 | 285 | 53.656 | 318 | 55.636 | 254 | 73.654 | 413 | 98.667 | 290 | 123.68 | 244
24.089 | 0 | 25.871 | 780 | 33.857 | 151 | 38.873 | 558 | 43.889 | 215 | 53.854 | 287 | 55.834 | 220 | 73.852 | 309 | 98.865 | 186 | 123.878 | 176
24.287 | 0 | 26.069 | 819 | 34.055 | 244 | 39.071 | 449 | 44.087 | 267 | 54.052 | 356 | 56.032 | 248 | 74.05 | 204 | 99.063 | 175 | 124.076 | 170
24.485 | 0 | 26.267 | 683 | 34.253 | 170 | 39.269 | 390 | 44.285 | 463 | 54.25 | 282 | 56.23 | 246 | 74.248 | 240 | 99.261 | 133 | 124.274 | 122
24.683 | 0 | 26.465 | 618 | 34.451 | 209 | 39.467 | 394 | 44.483 | 623 | 54.448 | 301 | 56.428 | 340 | 74.446 | 215 | 99.459 | 212 | 124.472 | 114
24.881 | 0 | 26.663 | 559 | 34.649 | 196 | 39.665 | 458 | 44.681 | 925 | 54.646 | 278 | 56.626 | 223 | 74.644 | 153 | 99.657 | 143 | 124.67 | 232
25.079 | 0 | 26.861 | 587 | 34.847 | 153 | 39.863 | 478 | 44.879 | 1224 | 54.844 | 240 | 56.824 | 276 | 74.842 | 130 | 99.855 | 139 | 124.868 | 121
25.277 | 0 | 27.059 | 549 | 35.045 | 144 | 40.061 | 460 | 45.077 | 2222 | 55.042 | 296 | 57.022 | 210 | 75.04 | 116 | 100.053 | 174 | 125.066 | 118
25.475 | 0 | 27.257 | 525 | 35.243 | 200 | 40.259 | 438 | 45.275 | 3279 | 55.24 | 277 | 57.22 | 275 | 75.238 | 180 | 100.251 | 248 | 125.264 | 203
25.673 | 0 | 27.455 | 631 | 35.441 | 186 | 40.457 | 520 | 45.473 | 5653 | 55.438 | 300 | 57.418 | 307 | 75.436 | 133 | 100.449 | 128 | 125.462 | 211
25.871 | 0 | 27.653 | 497 | 35.639 | 234 | 40.655 | 295 | 45.671 | 7737 | 55.636 | 236 | 57.616 | 213 | 75.634 | 149 | 100.647 | 112 | 125.66 | 131
26.069 | 0 | 27.851 | 581 | 35.837 | 215 | 40.853 | 420 | 45.869 | 10517 | 55.834 | 286 | 57.814 | 191 | 75.832 | 157 | 100.845 | 192 | 125.858 | 132
26.267 | 0 | 28.049 | 471 | 36.035 | 344 | 41.051 | 331 | 46.067 | 13952 | 56.032 | 310 | 58.012 | 208 | 76.03 | 155 | 101.043 | 187 | 126.056 | 162
26.465 | 0 | 28.247 | 547 | 36.233 | 288 | 41.249 | 329 | 46.265 | 16788 | 56.23 | 282 | 58.21 | 288 | 76.228 | 136 | 101.241 | 119 | 126.254 | 143
26.663 | 0 | 28.445 | 457 | 36.431 | 370 | 41.447 | 256 | 46.463 | 19385 | 56.428 | 229 | 58.408 | 282 | 76.426 | 233 | 101.439 | 124 | 126.452 | 128
26.861 | 0 | 28.643 | 472 | 36.629 | 398 | 41.645 | 251 | 46.661 | 21193 | 56.626 | 125 | 58.606 | 313 | 76.624 | 99 | 101.637 | 124 | 126.65 | 35
27.059 | 0 | 28.841 | 385 | 36.827 | 320 | 41.843 | 281 | 46.859 | 20219 | 56.824 | 106 | 58.804 | 284 | 76.822 | 217 | 101.835 | 184 | 126.848 | 137
27.257 | 0 | 29.039 | 447 | 37.025 | 323 | 42.041 | 164 | 47.056 | 19300 | 57.022 | 120 | 59.002 | 201 | 77.02 | 252 | 102.033 | 105 | 127.046 | 147
27.455 | 0 | 29.237 | 360 | 37.223 | 213 | 42.239 | 80 | 47.254 | 16772 | 57.22 | 113 | 59.2 | 200 | 77.218 | 260 | 102.231 | 73 | 127.244 | 114
27.653 | 0 | 29.435 | 353 | 37.421 | 285 | 42.437 | 66 | 47.452 | 13909 | 57.418 | 59 | 59.398 | 246 | 77.416 | 304 | 102.429 | 151 | 127.442 | 65
27.851 | 0 | 29.633 | 324 | 37.619 | 149 | 42.635 | 111 | 47.65 | 10622 | 57.616 | 105 | 59.596 | 274 | 77.614 | 355 | 102.627 | 120 | 127.64 | 128
28.049 | 0 | 29.831 | 308 | 37.817 | 116 | 42.833 | 56 | 47.848 | 8283 | 57.814 | 45 | 59.794 | 180 | 77.812 | 459 | 102.825 | 108 | 127.838 | 115
28.247 | 0 | 30.029 | 267 | 38.015 | 123 | 43.031 | 16 | 48.046 | 5767 | 58.012 | 57 | 59.992 | 148 | 78.01 | 716 | 103.023 | 66 | 128.036 | 105
28.445 | 0 | 30.227 | 273 | 38.213 | 106 | 43.229 | 72 | 48.244 | 3835 | 58.21 | 78 | 60.19 | 213 | 78.208 | 771 | 103.221 | 85 | 128.234 | 69
28.643 | 0 | 30.425 | 238 | 38.411 | 134 | 43.427 | 34 | 48.442 | 2665 | 58.408 | 94 | 60.388 | 119 | 78.405 | 1141 | 103.419 | 62 | 128.432 | 93
28.841 | 0 | 30.623 | 321 | 38.609 | 86 | 43.625 | 33 | 48.64 | 1834 | 58.606 | 133 | 60.586 | 68 | 78.603 | 1389 | 103.617 | 73 | 128.63 | 77
29.039 | 0 | 30.821 | 225 | 38.807 | 59 | 43.823 | 87 | 48.838 | 1126 | 58.804 | 78 | 60.784 | 248 | 78.801 | 1607 | 103.815 | 71 | 128.828 | 90
29.237 | 0 | 31.019 | 174 | 39.005 | 127 | 44.021 | 52 | 49.036 | 685 | 59.002 | 77 | 60.982 | 196 | 78.999 | 2103 | 104.013 | 90 | 129.026 | 72
29.435 | 0 | 31.217 | 166 | 39.203 | 70 | 44.219 | 104 | 49.234 | 464 | 59.2 | 111 | 61.18 | 150 | 79.197 | 2579 | 104.211 | 49 | 129.224 | 84
29.633 | 0 | 31.415 | 189 | 39.401 | 86 | 44.417 | 175 | 49.432 | 337 | 59.398 | 84 | 61.378 | 145 | 79.395 | 3029 | 104.409 | 55 | 129.422 | 48
29.831 | 0 | 31.613 | 194 | 39.599 | 77 | 44.615 | 330 | 49.63 | 277 | 59.596 | 139 | 61.576 | 188 | 79.593 | 3772 | 104.607 | 99 | 129.62 | 79
30.029 | 0 | 31.811 | 140 | 39.797 | 105 | 44.813 | 682 | 49.828 | 157 | 59.794 | 103 | 61.774 | 109 | 79.791 | 4606 | 104.805 | 99 | 129.818 | 66
30.227 | 0 | 32.009 | 141 | 39.995 | 24 | 45.011 | 1148 | 50.026 | 129 | 59.992 | 144 | 61.972 | 91 | 79.989 | 6078 | 105.003 | 48 | 130.016 | 72
30.425 | 0 | 32.207 | 131 | 40.193 | 152 | 45.209 | 2234 | 50.224 | 119 | 60.19 | 49 | 62.17 | 40 | 80.187 | 6256 | 105.201 | 87 | 130.214 | 53
30.623 | 0 | 32.405 | 135 | 40.391 | 135 | 45.407 | 3561 | 50.422 | 124 | 60.388 | 167 | 62.368 | 141 | 80.385 | 7147 | 105.399 | 92 | 130.412 | 61
30.821 | 0 | 32.603 | 157 | 40.589 | 126 | 45.605 | 5453 | 50.62 | 133 | 60.586 | 148 | 62.566 | 141 | 80.583 | 8010 | 105.597 | 81 | 130.61 | 71
31.019 | 0 | 32.801 | 109 | 40.787 | 92 | 45.803 | 7911 | 50.818 | 78 | 60.784 | 83 | 62.764 | 42 | 80.781 | 8574 | 105.795 | 68 | 130.808 | 95
31.217 | 0 | 32.999 | 120 | 40.985 | 157 | 46.001 | 11447 | 51.016 | 144 | 60.982 | 80 | 62.962 | 111 | 80.979 | 8886 | 105.993 | 54 | 131.006 | 78
31.415 | 0 | 33.197 | 239 | 41.183 | 72 | 46.199 | 15130 | 51.214 | 77 | 61.18 | 129 | 63.16 | 94 | 81.177 | 8848 | 106.191 | 45 | 131.204 | 104
31.613 | 0 | 33.395 | 211 | 41.381 | 92 | 46.397 | 17218 | 51.412 | 73 | 61.378 | 175 | 63.358 | 65 | 81.375 | 8830 | 106.389 | 48 | 131.402 | 72
31.811 | 0 | 33.593 | 187 | 41.579 | 157 | 46.595 | 20721 | 51.61 | 52 | 61.576 | 106 | 63.556 | 86 | 81.573 | 9142 | 106.587 | 21 | 131.6 | 61
32.009 | 0 | 33.791 | 136 | 41.777 | 43 | 46.793 | 21499 | 51.808 | 60 | 61.774 | 149 | 63.754 | 36 | 81.771 | 8621 | 106.785 | 0 | 131.798 | 88
32.207 | 0 | 33.989 | 211 | 41.975 | 97 | 46.99 | 21654 | 52.006 | 37 | 61.972 | 165 | 63.952 | 64 | 81.969 | 7909 | 106.983 | 46 | 131.996 | 39
32.405 | 0 | 34.187 | 208 | 42.173 | 85 | 47.188 | 20778 | 52.204 | 32 | 62.17 | 161 | 64.15 | 127 | 82.167 | 7157 | 107.181 | 43 | 132.194 | 59
32.603 | 0 | 34.385 | 226 | 42.371 | 46 | 47.386 | 18383 | 52.402 | 70 | 62.368 | 122 | 64.348 | 15 | 82.365 | 6525 | 107.379 | 57 | 132.392 | 70
32.801 | 0 | 34.583 | 193 | 42.569 | 47 | 47.584 | 15707 | 52.6 | 53 | 62.566 | 150 | 64.546 | 51 | 82.563 | 5857 | 107.577 | 62 | 132.59 | 58
32.999 | 0 | 34.781 | 212 | 42.767 | 108 | 47.782 | 12234 | 52.798 | 27 | 62.764 | 131 | 64.744 | 53 | 82.761 | 4953 | 107.775 | 49 | 132.788 | 50
33.197 | 0 | 34.979 | 206 | 42.965 | 105 | 47.98 | 9375 | 52.996 | 30 | 62.962 | 140 | 64.942 | 30 | 82.959 | 4231 | 107.973 | 96 | 132.986 | 111
33.395 | 0 | 35.177 | 204 | 43.163 | 31 | 48.178 | 6429 | 53.194 | 23 | 63.16 | 91 | 65.14 | 40 | 83.157 | 3478 | 108.171 | 62 | 133.184 | 70
33.593 | 0 | 35.375 | 218 | 43.361 | 47 | 48.376 | 4353 | 53.392 | 10 | 63.358 | 59 | 65.338 | 22 | 83.355 | 2703 | 108.369 | 30 | 133.382 | 88
33.791 | 0 | 35.573 | 228 | 43.559 | 21 | 48.574 | 2996 | 53.59 | 20 | 63.556 | 78 | 65.536 | 33 | 83.553 | 2470 | 108.567 | 78 | 133.58 | 71
33.989 | 0 | 35.771 | 322 | 43.757 | 32 | 48.772 | 1768 | 53.788 | 6 | 63.754 | 86 | 65.734 | 42 | 83.751 | 1855 | 108.765 | 54 | 133.778 | 63
34.187 | 0 | 35.969 | 328 | 43.955 | 53 | 48.97 | 1102 | 53.986 | 22 | 63.952 | 67 | 65.932 | 16 | 83.949 | 1594 | 108.963 | 134 | 133.976 | 134
34.385 | 0 | 36.167 | 360 | 44.153 | 81 | 49.168 | 764 | 54.184 | 6 | 64.15 | 43 | 66.13 | 32 | 84.147 | 1302 | 109.161 | 103 | 134.174 | 79
34.583 | 0 | 36.365 | 472 | 44.351 | 151 | 49.366 | 506 | 54.382 | 15 | 64.348 | 16 | 66.328 | 28 | 84.345 | 981 | 109.359 | 120 | 134.372 | 30
34.781 | 0 | 36.563 | 635 | 44.549 | 235 | 49.564 | 299 | 54.58 | 23 | 64.546 | 51 | 66.526 | 4 | 84.543 | 871 | 109.556 | 106 | 134.57 | 26
34.979 | 0 | 36.761 | 427 | 44.747 | 398 | 49.762 | 171 | 54.778 | 7 | 64.744 | 10 | 66.724 | 31 | 84.741 | 611 | 109.754 | 70 | 134.768 | 55
35.177 | 0 | 36.959 | 578 | 44.945 | 765 | 49.96 | 71 | 54.976 | 13 | 64.942 | 39 | 66.922 | 41 | 84.939 | 579 | 109.952 | 168 | 134.966 | 52
35.375 | 0 | 37.157 | 577 | 45.143 | 1436 | 50.158 | 78 | 55.174 | 0 | 65.14 | 33 | 67.12 | 36 | 85.137 | 423 | 110.15 | 171 | 135.164 | 83
35.573 | 0 | 37.355 | 345 | 45.341 | 2460 | 50.356 | 77 | 55.372 | 4 | 65.338 | 1 | 67.318 | 61 | 85.335 | 352 | 110.348 | 193 | 135.362 | 90
35.771 | 0 | 37.553 | 287 | 45.539 | 3620 | 50.554 | 24 | 55.57 | 13 | 65.536 | 6 | 67.516 | 16 | 85.533 | 305 | 110.546 | 230 | 135.56 | 57
35.969 | 0 | 37.751 | 421 | 45.737 | 6039 | 50.752 | 22 | 55.768 | 14 | 65.734 | 18 | 67.714 | 65 | 85.731 | 236 | 110.744 | 220 | 135.758 | 71
36.167 | 0 | 37.949 | 196 | 45.935 | 8732 | 50.95 | 23 | 55.966 | 0 | 65.932 | 7 | 67.912 | 35 | 85.929 | 310 | 110.942 | 289 | 135.956 | 92
36.365 | 0 | 38.147 | 189 | 46.133 | 12280 | 51.148 | 27 | 56.164 | 2 | 66.13 | 15 | 68.11 | 29 | 86.127 | 202 | 111.14 | 270 | 136.154 | 70
36.563 | 0 | 38.345 | 112 | 46.331 | 15926 | 51.346 | 39 | 56.362 | 0 | 66.328 | 71 | 68.308 | 47 | 86.325 | 200 | 111.338 | 334 | 136.352 | 42
36.761 | 0 | 38.543 | 126 | 46.529 | 19745 | 51.544 | 7 | 56.56 | 6 | 66.526 | 86 | 68.506 | 28 | 86.523 | 103 | 111.536 | 420 | 136.55 | 44
36.959 | 0 | 38.741 | 60 | 46.727 | 22635 | 51.742 | 24 | 56.758 | 8 | 66.724 | 49 | 68.704 | 32 | 86.721 | 79 | 111.734 | 558 | 136.748 | 81
37.157 | 0 | 38.939 | 123 | 46.924 | 24421 | 51.94 | 27 | 56.956 | 17 | 66.922 | 24 | 68.902 | 46 | 86.919 | 74 | 111.932 | 576 | 136.946 | 52
37.355 | 0 | 39.137 | 135 | 47.122 | 24424 | 52.138 | 17 | 57.154 | 4 | 67.12 | 42 | 69.1 | 103 | 87.117 | 46 | 112.13 | 812 | 137.144 | 7
37.553 | 0 | 39.335 | 116 | 47.32 | 22934 | 52.336 | 34 | 57.352 | 0 | 67.318 | 13 | 69.298 | 40 | 87.315 | 26 | 112.328 | 802 | 137.342 | 11
37.751 | 0 | 39.533 | 89 | 47.518 | 20968 | 52.534 | 32 | 57.55 | 0 | 67.516 | 39 | 69.496 | 15 | 87.513 | 52 | 112.526 | 982 | 137.54 | 6
37.949 | 0 | 39.731 | 135 | 47.716 | 17439 | 52.732 | 3 | 57.748 | 0 | 67.714 | 50 | 69.694 | 30 | 87.711 | 22 | 112.724 | 1057 | 137.738 | 16
38.147 | 0 | 39.929 | 68 | 47.914 | 13729 | 52.93 | 8 | 57.946 | 12 | 67.912 | 18 | 69.892 | 22 | 87.909 | 46 | 112.922 | 1362 | 137.936 | 24
38.345 | 0 | 40.127 | 58 | 48.112 | 10543 | 53.128 | 0 | 58.144 | 0 | 68.11 | 15 | 70.09 | 59 | 88.107 | 12 | 113.12 | 1559 | 138.134 | 10
38.543 | 0 | 40.325 | 145 | 48.31 | 7541 | 53.326 | 28 | 58.342 | 0 | 68.308 | 59 | 70.288 | 33 | 88.305 | 22 | 113.318 | 1735 | 138.332 | 25
38.741 | 0 | 40.523 | 155 | 48.508 | 5064 | 53.524 | 28 | 58.54 | 3 | 68.506 | 38 | 70.486 | 42 | 88.503 | 12 | 113.516 | 1900 | 138.53 | 10
38.939 | 0 | 40.721 | 123 | 48.706 | 3098 | 53.722 | 28 | 58.738 | 4 | 68.704 | 53 | 70.684 | 42 | 88.701 | 10 | 113.714 | 2292 | 138.728 | 7
39.137 | 0 | 40.919 | 240 | 48.904 | 2178 | 53.92 | 20 | 58.936 | 5 | 68.902 | 35 | 70.882 | 38 | 88.899 | 13 | 113.912 | 2746 | 138.926 | 4
39.335 | 0 | 41.117 | 117 | 49.102 | 1473 | 54.118 | 15 | 59.134 | 0 | 69.1 | 63 | 71.08 | 30 | 89.097 | 13 | 114.11 | 2766 | 139.124 | 9
39.533 | 0 | 41.315 | 187 | 49.3 | 891 | 54.316 | 8 | 59.332 | 0 | 69.298 | 40 | 71.278 | 67 | 89.295 | 11 | 114.308 | 3102 | 139.322 | 6
39.731 | 0 | 41.513 | 104 | 49.498 | 534 | 54.514 | 22 | 59.53 | 0 | 69.496 | 18 | 71.476 | 42 | 89.493 | 35 | 114.506 | 3658 | 139.52 | 8
39.929 | 0 | 41.711 | 232 | 49.696 | 369 | 54.712 | 0 | 59.728 | 0 | 69.694 | 6 | 71.674 | 33 | 89.691 | 11 | 114.704 | 4012 | 139.718 | 0
40.127 | 0 | 41.909 | 105 | 49.894 | 187 | 54.91 | 24 | 59.926 | 0 | 69.892 | 109 | 71.872 | 33 | 89.889 | 20 | 114.902 | 4078 | 139.916 | 19
40.325 | 0 | 42.107 | 74 | 50.092 | 85 | 55.108 | 25 | 60.124 | 0 | 70.09 | 51 | 72.07 | 31 | 90.087 | 0 | 115.1 | 4260 | 140.114 | 0
40.523 | 0 | 42.305 | 58 | 50.29 | 81 | 55.306 | 6 | 60.322 | 0 | 70.288 | 29 | 72.268 | 46 | 90.285 | 5 | 115.298 | 4595 | 140.312 | 5
40.721 | 0 | 42.503 | 59 | 50.488 | 86 | 55.504 | 17 | 60.52 | 12 | 70.486 | 16 | 72.466 | 0 | 90.483 | 0 | 115.496 | 4847 | 140.51 | 8
40.919 | 0 | 42.701 | 73 | 50.686 | 5 | 55.702 | 14 | 60.718 | 0 | 70.684 | 23 | 72.664 | 65 | 90.681 | 9 | 115.694 | 4970 | 140.708 | 8
41.117 | 0 | 42.899 | 94 | 50.884 | 54 | 55.9 | 9 | 60.916 | 0 | 70.882 | 0 | 72.862 | 9 | 90.879 | 10 | 115.892 | 4758 | 140.906 | 19
41.315 | 0 | 43.097 | 108 | 51.082 | 33 | 56.098 | 30 | 61.114 | 5 | 71.08 | 14 | 73.06 | 22 | 91.077 | 13 | 116.09 | 4772 | 141.103 | 15
41.513 | 0 | 43.295 | 101 | 51.28 | 27 | 56.296 | 13 | 61.312 | 0 | 71.278 | 0 | 73.258 | 53 | 91.275 | 16 | 116.288 | 4941 | 141.301 | 4
41.711 | 0 | 43.493 | 57 | 51.478 | 21 | 56.494 | 6 | 61.51 | 0 | 71.476 | 14 | 73.456 | 9 | 91.473 | 0 | 116.486 | 4934 | 141.499 | 16
41.909 | 0 | 43.691 | 54 | 51.676 | 44 | 56.692 | 0 | 61.708 | 0 | 71.674 | 12 | 73.654 | 26 | 91.671 | 0 | 116.684 | 4469 | 141.697 | 20
42.107 | 0 | 43.889 | 70 | 51.874 | 43 | 56.89 | 4 | 61.906 | 0 | 71.872 | 14 | 73.852 | 17 | 91.869 | 9 | 116.882 | 4464 | 141.895 | 22
42.305 | 0 | 44.087 | 33 | 52.072 | 9 | 57.088 | 0 | 62.104 | 2 | 72.07 | 16 | 74.05 | 17 | 92.067 | 4 | 117.08 | 4235 | 142.093 | 16
42.503 | 0 | 44.285 | 49 | 52.27 | 8 | 57.286 | 8 | 62.302 | 3 | 72.268 | 0 | 74.248 | 5 | 92.265 | 4 | 117.278 | 4120 | 142.291 | 32
42.701 | 0 | 44.483 | 115 | 52.468 | 23 | 57.484 | 0 | 62.5 | 0 | 72.466 | 6 | 74.446 | 15 | 92.463 | 0 | 117.476 | 3685 | 142.489 | 31
42.899 | 0 | 44.681 | 240 | 52.666 | 29 | 57.682 | 6 | 62.698 | 0 | 72.664 | 0 | 74.644 | 25 | 92.661 | 10 | 117.674 | 3298 | 142.687 | 17
43.097 | 0 | 44.879 | 380 | 52.864 | 26 | 57.88 | 3 | 62.896 | 0 | 72.862 | 6 | 74.842 | 20 | 92.859 | 6 | 117.872 | 3385 | 142.885 | 64
43.295 | 0 | 45.077 | 727 | 53.062 | 30 | 58.078 | 3 | 63.094 | 0 | 73.06 | 0 | 75.04 | 15 | 93.057 | 7 | 118.07 | 2678 | 143.083 | 81
43.493 | 0 | 45.275 | 1266 | 53.26 | 5 | 58.276 | 9 | 63.292 | 0 | 73.258 | 0 | 75.238 | 6 | 93.255 | 0 | 118.268 | 2502 | 143.281 | 26
43.691 | 0 | 45.473 | 2405 | 53.458 | 36 | 58.474 | 0 | 63.49 | 0 | 73.456 | 0 | 75.436 | 26 | 93.453 | 12 | 118.466 | 2237 | 143.479 | 69
43.889 | 0 | 45.671 | 3681 | 53.656 | 26 | 58.672 | 0 | 63.688 | 4 | 73.654 | 0 | 75.634 | 2 | 93.651 | 0 | 118.664 | 1904 | 143.677 | 80
44.087 | 0 | 45.869 | 6035 | 53.854 | 27 | 58.87 | 21 | 63.886 | 0 | 73.852 | 0 | 75.832 | 0 | 93.849 | 6 | 118.862 | 1748 | 143.875 | 113
44.285 | 0 | 46.067 | 9273 | 54.052 | 12 | 59.068 | 4 | 64.084 | 0 | 74.05 | 0 | 76.03 | 0 | 94.047 | 13 | 119.06 | 1499 | 144.073 | 155
44.483 | 0 | 46.265 | 13239 | 54.25 | 8 | 59.266 | 0 | 64.282 | 3 | 74.248 | 0 | 76.228 | 2 | 94.245 | 0 | 119.258 | 1375 | 144.271 | 160
44.681 | 0 | 46.463 | 17709 | 54.448 | 20 | 59.464 | 3 | 64.48 | 0 | 74.446 | 6 | 76.426 | 7 | 94.443 | 4 | 119.456 | 1082 | 144.469 | 217
44.879 | 0 | 46.661 | 22956 | 54.646 | 3 | 59.662 | 0 | 64.678 | 0 | 74.644 | 0 | 76.624 | 12 | 94.641 | 11 | 119.654 | 951 | 144.667 | 149
45.077 | 0 | 46.859 | 26435 | 54.844 | 4 | 59.86 | 2 | 64.876 | 0 | 74.842 | 0 | 76.822 | 0 | 94.839 | 11 | 119.852 | 820 | 144.865 | 142
45.275 | 0 | 47.056 | 29080 | 55.042 | 4 | 60.058 | 0 | 65.074 | 0 | 75.04 | 0 | 77.02 | 0 | 95.037 | 4 | 120.05 | 712 | 145.063 | 178
45.473 | 0 | 47.254 | 30749 | 55.24 | 17 | 60.256 | 3 | 65.272 | 0 | 75.238 | 0 | 77.218 | 4 | 95.235 | 17 | 120.248 | 576 | 145.261 | 196
45.671 | 0 | 47.452 | 28845 | 55.438 | 18 | 60.454 | 0 | 65.47 | 0 | 75.436 | 0 | 77.416 | 3 | 95.433 | 21 | 120.446 | 423 | 145.459 | 173
45.869 | 0 | 47.65 | 26344 | 55.636 | 15 | 60.652 | 4 | 65.668 | 0 | 75.634 | 0 | 77.614 | 0 | 95.631 | 0 | 120.644 | 436 | 145.657 | 158
46.067 | 0 | 47.848 | 22820 | 55.834 | 25 | 60.85 | 0 | 65.866 | 0 | 75.832 | 0 | 77.812 | 0 | 95.829 | 8 | 120.842 | 334 | 145.855 | 158
46.265 | 0 | 48.046 | 18382 | 56.032 | 7 | 61.048 | 0 | 66.064 | 0 | 76.03 | 6 | 78.01 | 0 | 96.027 | 13 | 121.04 | 379 | 146.053 | 70
46.463 | 0 | 48.244 | 14290 | 56.23 | 22 | 61.246 | 0 | 66.262 | 0 | 76.228 | 0 | 78.208 | 0 | 96.225 | 0 | 121.238 | 244 | 146.251 | 63
46.661 | 0 | 48.442 | 10188 | 56.428 | 6 | 61.444 | 0 | 66.46 | 4 | 76.426 | 0 | 78.405 | 0 | 96.423 | 4 | 121.436 | 163 | 146.449 | 51
46.859 | 0 | 48.64 | 6979 | 56.626 | 3 | 61.642 | 0 | 66.658 | 0 | 76.624 | 0 | 78.603 | 0 | 96.621 | 0 | 121.634 | 173 | 146.647 | 31
47.056 | 0 | 48.838 | 4635 | 56.824 | 0 | 61.84 | 0 | 66.856 | 0 | 76.822 | 0 | 78.801 | 0 | 96.819 | 0 | 121.832 | 138 | 146.845 | 24
47.254 | 0 | 49.036 | 2931 | 57.022 | 0 | 62.038 | 0 | 67.054 | 0 | 77.02 | 0 | 78.999 | 0 | 97.017 | 0 | 122.03 | 117 | 147.043 | 18
47.452 | 0 | 49.234 | 2074 | 57.22 | 7 | 62.236 | 0 | 67.252 | 0 | 77.218 | 0 | 79.197 | 0 | 97.215 | 7 | 122.228 | 120 | 147.241 | 6
47.65 | 0 | 49.432 | 1070 | 57.418 | 11 | 62.434 | 0 | 67.45 | 0 | 77.416 | 0 | 79.395 | 0 | 97.413 | 0 | 122.426 | 136 | 147.439 | 9
47.848 | 0 | 49.63 | 678 | 57.616 | 9 | 62.632 | 0 | 67.648 | 0 | 77.614 | 0 | 79.593 | 0 | 97.611 | 0 | 122.624 | 68 | 147.637 | 0
48.046 | 0 | 49.828 | 444 | 57.814 | 8 | 62.83 | 0 | 67.846 | 0 | 77.812 | 0 | 79.791 | 0 | 97.809 | 3 | 122.822 | 58 | 147.835 | 0
48.244 | 0 | 50.026 | 221 | 58.012 | 0 | 63.028 | 0 | 68.044 | 0 | 78.01 | 0 | 79.989 | 0 | 98.007 | 7 | 123.02 | 42 | 148.033 | 3
48.442 | 0 | 50.224 | 158 | 58.21 | 0 | 63.226 | 0 | 68.242 | 0 | 78.208 | 0 | 80.187 | 0 | 98.205 | 12 | 123.218 | 16 | 148.231 | 0
48.64 | 0 | 50.422 | 148 | 58.408 | 0 | 63.424 | 0 | 68.44 | 0 | 78.405 | 0 | 80.385 | 0 | 98.403 | 0 | 123.416 | 16 | 148.429 | 0
48.838 | 0 | 50.62 | 81 | 58.606 | 5 | 63.622 | 0 | 68.638 | 0 | 78.603 | 0 | 80.583 | 0 | 98.601 | 4 | 123.614 | 18 | 148.627 | 0
49.036 | 0 | 50.818 | 52 | 58.804 | 0 | 63.82 | 0 | 68.836 | 0 | 78.801 | 0 | 80.781 | 0 | 98.799 | 0 | 123.812 | 38 | 148.825 | 0
49.234 | 0 | 51.016 | 67 | 59.002 | 0 | 64.018 | 0 | 69.034 | 0 | 78.999 | 0 | 80.979 | 0 | 98.997 | 11 | 124.01 | 13 | 149.023 | 0
49.432 | 0 | 51.214 | 32 | 59.2 | 7 | 64.216 | 0 | 69.232 | 0 | 79.197 | 0 | 81.177 | 0 | 99.195 | 19 | 124.208 | 24 | 149.221 | 0
|  | 51.412 | 13 | 59.398 | 0 | 64.414 | 0 | 69.43 | 0 | 79.395 | 0 | 81.375 | 0 | 99.393 | 0 | 124.406 | 11 | 149.419 | 0

## Sheet1
Compound | Chemical Formula | Exact Mass | Single Pass t_p | t_pp | [M-H]- | Observed m/z | Error (ppm) | Accepted CCS | Single Pass t_p | t_pp | [M+H]+ | Observed m/z | Error (ppm) | Accepted CCS
Theophylline | C7H8N4O2 | 180.0647 |  |  | 179.0569 | 179.0583 | -7.8288 | 133.6 |  |  | 181.0725 |  |  | 140.9
Sulfadimethoxide | C12H14N4O4S | 310.0736 |  |  | 309.0657 | 309.0666 | -2.7573 | 170.1 |  |  | 311.0814 | 311.0808 | 1.9191 | 168.4
Perfluorooctanoic acid -CO2 | HC6F13 | 319.9871 |  |  | 318.9792 | 318.9803 | -3.3181 | 130.1
Perfluorooctanoic acid -CO2 | HC7F15 | 369.9839 |  |  | 368.976 | 368.9777 | -4.4773 | 137.2
Val-Tyr-Val | C19H29N3O5 | 379.2107 |  |  | 378.2029 | 378.2033 | -1.0952 | 192.5 |  |  | 380.2185 | 380.2182 | 0.88107 | 191.7
Terfenadine | C32H41NO2 | 471.3137 | [8 empty] | 472.3215 | 472.321 | 1.1374 | 228.7
Polyalanine n7 | C21H37N7O8 | 515.2703 |  |  | 514.2625 | 514.2644 | -3.6487 |  |  |  | 516.2782 | 516.2763 | 3.6275 | 211
Leucine Enkephalin | C28H37N5O7 | 555.2693 |  |  | 554.2615 | 554.2621 | -1.1536 | 225.3 |  |  | 556.2771 | 556.2786 | -2.6789 | 229.8
Polyalanine n8 | C24H42N8O9 | 586.3075 |  |  | 585.2996 | 585.2993 | 0.57338 | 227.7 |  |  | 587.3153 | 587.3156 | -0.53668 | 228
Reserpine | C33H40N2O9 | 608.2734 | [8 empty] | 609.2812 | 609.2818 | -0.99855 | 252.3
Polyalanine n9 | C27H47N9O10 | 657.3446 |  |  | 656.3367 | 656.3391 | -3.5841 | 242.1 |  |  | 658.3524 | 658.3498 | 3.9444 | 243
Polyalanine n10 | C30H52N10O11 | 728.3817 |  |  | 727.3739 | 727.3741 | -0.3305 | 255.9 |  |  | 729.3895 | 729.3893 | 0.28627 | 256
Polyalanine n11 | C33H57N11O12 | 799.4188 |  |  | 798.411 | 798.4109 | 0.089678 | 268.5 |  |  | 800.4266 | 800.4257 | 1.1504 | 271
Polyalanine n12 | C36H62N12O13 | 870.4559 |  |  | 869.4481 | 869.4493 | -1.399 | 280.2 |  |  | 871.4637 | 871.4611 | 3.0211 | 282
Polyalanine n13 | C39H67N13O14 | 941.493 |  |  | 940.4852 | 940.4849 | 0.31431 | 294.6 |  |  | 942.5008 | 942.5037 | -3.0294 | 294
Polyalanine n14 | C42H72N14O15 | 1012.5301 |  |  | 1011.5223 | 1011.5225 | -0.19021 | 308.8 |  |  | 1013.538 | 1013.5332 | 4.6933 | 306
Polyalanine n15 | C45H77N15O16 | 1083.5672 |  |  | 1082.5594 | 1082.5625 | -2.8455 | 322.4 |  |  | 1084.5751 | 1084.5692 | 5.4112 | 321.5
|  | 1154.6044 |  |  | 1153.5965 | 1153.5957 | 0.72088 |  |  |  | 1155.6122 | 1155.6073 | 4.2236 | 333.6
other possible candidates | succinic
salicylic acid
Accpeted CCS Values from | https://mcleanresearchgroup.shinyapps.io/CCS-Compendium/ | SPLASH LIPIDS
measured with an Aligent reference system
| https://pubs.acs.org/doi/suppl/10.1021/acs.jafc.2c00724/suppl_file/jf2c00724_si_001.pdf |  |  |  | Major Mix Standards
Acetaminophen | C8H9NO2 | 151.0633 |  |  | 150.0555 |  |  | 131.5 |  |  | 152.0711 |  |  | 130.4
Resperine Fragment | C8H10N4O2 | 194.0804 |  |  | 193.0725 | [5 empty] | 195.0882 |  |  | 138.2
Sulfaguanidine | C7H10N4O2S | 214.0524 |  |  | 213.0446 |  |  | 145.2 |  |  | 215.0603 |  |  | 146.8
Verapamil | C27H38N2O4 | 454.2831 |  |  | 453.2753 | [5 empty] | 455.291 |  |  | 208.8
Polyalanine n16 | C48H82N16O17 | 1154.6044 |  |  | 1153.5965 | [5 empty] | 1155.6122 |  |  | 333.6
Ultramark 1621 | C20H18O6N3P3F28 | 1020.9961 |  |  | 1019.9883
Ultramark 1621 | C22H18O6N3P3F32 | 1120.9897 |  |  | 1119.9819 | [5 empty] | 1121.9976 |  |  | 276.5